INSTRUCTIONS

PART 1 - PURPOSE

GENERAL

The type 862 Valve and Circuit tester is a self-contained instrument for making the following types  of  measurements rapidly and conveniently:

(a)	Resistors: Four Ranges.  Direct reading Ohm-meter for resistances between 0.1 Ohm and 10 Megohms

(b)	Condensers: Leakage and open circuit on Mica and Paper Condensers on   direct   current insulation tester. Leakage current of low voltage and high voltage Electrolytic Condensers.

(c)	Milliamperes: Four ranges  of  direct current milliamperes. Full scale  deflections:  ImA,  lOmA,  lOOmA, 250mA.  100 mV  voltage  drop  at  full scale deflection.

(d)	Volts D.C: Four ranges.   Full scale deflections of lOV; l00V, 250V,  1000V.   Meter  resistance: 1000 Ohms/Volt.

.(e)	Volts A.C: Four  Ranges Valve Rectification.  Full scale deflections: l0V, 100V, 250V, 1000v,

(f)	Output Volts:	Identical with A.C. ranges.

(g)	Valves: Merit test on  direct reading scale and short circuit or  leakage  between  elements of all common type  receiver  and  small transmitter valves.

PART 2 - OPERATING INSTRUCTIONS.

POWER SUPPLY:

The type 862 Valve and Circuit Tester can be operated from either A.C (200 to 260 volts) or from a 6 Volt Accumulator.

The 5 pin  plug  on  the  supply  line  of the instrument, coming through a hole inside  the  small compartment on the top of the case, is wired according to the circuit given below:-

0 240V 0

0	6v -  0

0

Earth

(a)	A.C. Operation.  The three core power cord supplied with the instrument and fitted with a 5 pin base is wired according to S.A.A. rules, green is  earth,  red and black 240 Volts A.C.  After connecting the  power  cord  to  the Mains and inserting the plug on the supply  line into the 5 pin base the instrument is ready for use.

(b)	Operation from  6  Volt  Battery. The  5  pin  plug on the instruments supply line has to  be inserted into the 5 pin base on the battery cord and the battery cord connected to a 6V accumulator.

The current drawn from a  6V battery is approximately 1.75 Amperes.  When  a  valve  is  inserted  for  emission  test the current drawn from the battery may raise to 6 Amperes according to the filament  current  and  the  voltage  of the valve under test

MULTITESTER RANGES:

NOTE:- It is necessary to  have the instrument connected to the power  supply  (A.C.   or   battery)   for   all  tests  except Milliamperes, D.C. Volts and the Low Ohms, Ohms and Ohms x 10.

MEGOHMS:- The voltage across  the  test  prods is 225 volts and care has to be taken not to  touch them when the Megohm test is carried out.  However since the maximum power present is only a fraction of  a  watt  a  shock  would  be  more unpleasant than dangerous -

Before taking a reading the switch in the left hand bottom corner is placed into  "MEGOHMS"  position  and  the test prods shorted  together.   Adjust  the   instrument   to  full  scale deflection by means  of  the  "LINE  ADJUST"  switch  (top left hand.)  When the  prods  are  placed  across  the  points to be measured the  meter  will  show  the  resistance.  Multiply the reading on the Ohms (top scale) by 500 to obtain the resistance value.  The Megohm range which extends from 0 to 10 megohms can be used to measure  high  resistors  and potentiometers and for testing insulation properties of  component  parts like sockets etc.

OHMS:- Two ranges are  available,  both  operated  from the dry cell placed in  the  compartment  on  the  top  of the cabinet, Place left hand bottom  switch  into  "OHMS" position and right hand bottom switch into either  "OHMS"  or "OHMS x 10" position as required.  Before taking  a  reading,  short  the test prods together and adjust the instrument  to full scale deflection by means of the control marked "OHMS ADJ."  Read the resistance on the Ohms scale.  Multiply  value  by  10  for  the  "Ohms x 10" range.

LOW OHMS:- This range, for which  the voltage is applied by the battery and  which  extends  from  0  to  30  ohms,  allows the measurement of resistance values as low  as  .1 of an ohm which gives a reading of half a  division.   It can therefore be used to check poor contacts,  short  circuits  in  coils, dry joints etc.

The left hand bottom switch is placed into "OHMS" position and the right  hand  bottom  switch  is  placed into "LOW OHMS" positon.  The pointer of  the  instrument  will go up scale and has to be adjusted to full scale deflection by means of control marked "OHMS ADJ."

When the test prods are  shorted together the pointer will not return exactly to zero owing  to the resistance of the test leads themselves.

NOTE:- Do  not  leave  the  right  hand  switch  in  "LOW OHMS" position when not using this  range,  because this will exhaust the battery in a  very  short  time  owing  to the high current drain.

ELECTROLYTIC CONDENSERS:-  The  Valve  and  Circuit  Tester  is provided with two ranges for the test of either the low voltage type electrolytic condenser up to  25 volts test voltage or the high voltage types which are encountered in the filter net-work of the power  supply  in  wireless  receivers  etc.  The rotary switch  in  the  left  hand   bottom   corner  has  to  be  set accordingly.  Preferably alligator  clip  should  be placed on the test  prods  and  the  connection  across  the electrolytic condenser under test made permanent.

NOTE: IT IS IMPORTANT TO OBSERVE THE CORRECT POLARITY OTHERWISE THE ELECTROLYTIC CONDENSER WILL BE DESTROYED.

The direct  reading  dial  on  the  instrument  will  show  the condition of the condenser.   If  an electrolytic condenser has not been used for sometime it  will be necessary to allow it to "form" for a minute or two  after it has been connected for the test.  The pointer of the  instrument will, if the condenser is good1 gradually go down the scale, finally becoming stationary,

NOTE:- If the condenser is wired  into  a piece of equipment it will be  necessary  to  disconnect  the  positive  lead  of the condenser  under test from the circuit before making the test.

PAPER CONDENSERS: Leakage or  open  circuit in paper condensers can be detected.  Turn left  hand rotary switch to "Paper Cond. Test."  Now apply the prods  across the condenser.  In case the condenser is in parallel to  a resistor or another component or an Electrolytic Condenser it  will  necessary to disconnect one end.  For sizes from  0.1  mfd.  or  over  clip  test prod onto condenser and observe the  neon.  On  a good condenser the neon will flash regularly, the period  of the flash depending on the size of the condenser,  For lower capacity values than 0.1 mfd. touch the terminals of the  condenser  with the prods and watch for a single flash  that  should  occur.  Very low values (1000 mmfd and lower) will only generate  a very minute flash and the neon should be watched closely.   A  second flash will not take place  unless  the  condenser  is  discharged  or  the polarity reversed.

A pronounced  continuous  glow  in  the  neon  will  indicate a leakage.  Sometimes a  faint  glow  will be present, especially when the flexible  cords  of  the  test  prods are placed close together or the operators hands  are  touching them,  This will indicate a leakage and is quite insignificant.

MILLAMPERES:	The left  hand  rotary  switch  is  placed into mA position and the correct range  selected  by means of the right hand rotary switch.  Always  take  the  initial readings on the highest range first  and  switch  to  lower  ranges if readings obtained are too far down the scale.

D.C. VOLTS: Turn  left  hand  rotary  switch  into "D.C. Volts" position and select the correct  range on the right hand bottom switch.  Always take a  reading  on  the highest range first to avoid overloading the instrument.

A.C. VOLTS: A valve rectifier is incorporated in the tester for A.C, volt ranges.   Turn  left  hand  rotary  switch into "A.C.Volts" position.  Right hand switch  serves as selector for the

range.

NOTE:	THE INSTRUMENT  MUST  BE  CONNECTED  TO  THE POWER SUPPLY (A.C. OR BATTERY) TO ALLOW THE RECTIFIER VALVE TO OPERATE.

Measurements on the A.C. line  can be taken notwithstanding the fact that the instrument is plugged into same.

It will be noted that on the lower A.C. ranges the pointer will advance a few degrees above the  zero.   This is due to the use of a valve as a rectifier;  this idle current is caused by some electrons reaching the  plate  owing  to their initial velocity despite no voltage on  the  plate.   No  inaccuracy however is caused by this fact.  The effect will not be so apparent on the higher ranges.

OUTPUT VOLTS: The "Output Volt" ranges use the valve rectifier. A built in condenser of 2  Microfarad  blocks the D.C. from the signal, Set left-hand switch into  0/Pt. V. position and select correct range with right hand switch.

VALVE TESTING

I.	Plug the instrument into power supply and switch on   Turn the bottom  left  hand  switch  to  position  marked "LINE CHECK."  The  meter needle  should  now  swing  over to full scale deflection.  Should the  needle  stop above or below this point, adjust by  means  of  the top left hand switch marked "LINE ADJ."  It is  important  in some districts to check the line voltage  constantly.   When testing a valve with the tester  operating  from  a  6  volt battery it is necessary to adjust  line  with  valve  under  test in its correct socket, after carrying out  the operations 2 and 3 (Below).

2.	Select correct filament voltage as shown on chart.

- Plug valve into correct socket as shown on chart.

4.	Test for inter-element  shorts  and leakages after waiting for valve to heat. A shorted element or leakage as high as 100,000 ohms will  cause  the  neon  electrodes  to show a definite glow. A very  faint  glow,  which may be apparent even with the valve out of a socket, can be disregarded,

NOTE:	UNDER NO CONDITION CARRY  OUT  ANY FURTHER TESTS IF

ANY SHORT IS  INDICATED,  AS  THE  METER  MAY BE SERIOUSLY

OVERLOADED; RESULTING  IN  THE  PERMANENT  DAMAGE  OF  THE

MOVEMENT. If the valve passes  the  above test, proceed as follows

5.	Move bottom left hand switch to position marked "MERIT."

6.	Turn selector switch (Middle  right  hand) to number shown on chart.

7.	Adjust range  control  (Middle  potentiometer  below meter) accurately to correct number shown on chart.

8.	Press  button  marked  "MERIT"  and  the  meter  will show directly the condition of  the valve.  When testing diodes the merit button must  be  only  depressed for the section marked TA and  PA  and  the  button  marked diodes for the selector switch numbers with the suffix DA.  On no account press either button before valve is heated sufficiently to give full emission or for  a  period longer than three (3) seconds  which is ample time  to read the condition of the valve.

The reading of diode plates will vary by as much as 20% to 50% between various valves of  the same type; but approximately equal readings will be obtained for the diode plates of any one particular valve.

9.	TESTING OF 1.4 VOLT VALVES: For merit test on these valve, press the diode button.  DO  NOT  use the MERIT button, as this might cause damage to the valves.

10.	This valve tester indicates "end-of-life" point, so far as emission is  concerned.   The  emission  readings  of  new valves have  no  direct  significance,  provided  they are above this point.

NOTE:	To identify the  shorting of elements, the following information can be utilized.

Selector Switch  Position  Number: (NOTE:  The  purpose  of this switch is to  select  that  electrode  in  the valve which collects enough electrons from the  cathode to serve as an emission  indicator,,  Generally,  this  is  the electrode nearest to the cathode),

1.	Control or screen grid.

2.	Plate

3.& 4.   Filaments (not shown on selector switch,)

5.	Cathode

6.	Supressor-screen or osc-grid.

7.	Osc.-grid or plate

8.	Top cap or control grid on screen grid valves.

Method of Calibrating Valves not Listed:

Take two or  three  valves  of  type  required  and follow this procedure:

1.	Check line voltage,

2.	Set filament switch to correct filament voltage of valve

3.	Turn  bottom  left  hand  switch to  "MERIT",  turn Range Control to about (5)  and	 plug	in valve, connecting grid cap,  if any.

4.	Depress Merit Button  and  rotate selector switch, setting it to the point giving the highest reading on meter.

5.	Now advance Range control until meter reads between 90 and

100 (on  the  voltage  scale).  Repeat  above  with valves available, then note readings such as Socket No. Filament, Range settings and Selector position for further use.

The Use of The Adapter Panel:

To provide test facilities for the valve which are only now being released on the market, and  to cover those which will be available shortly, the adapter  panel  has  been included as an integral part of model 862 Valve and Circuit Tester.

The adapter panel which  is  mounted  in  the front lid is connected to the main instrument by  means of an octal plug and also a banana plug.  These are connected to NO.7 socket and the TC socket respectively.  This  makes  it  possible  to select a suitable electrode by means of  the Selector and also to supply heater voltages.

If the valve to be tested is listed in the chart,  proceed as directed under Valve Testing, Sections 1-10.  In the case of valves which are not as yet listed, proceed in accordance with the directions  shown  in  the  proceeding  section  (Method of Calibrating Valves Not Listed).

The sockets  F,  G  and  H  are  for  Loctal-based valves. Sockets J and K are for  Octal-based valves, in which pins Nos. 1 and  8,  and  pins  7  and  8,  are  the  heater  connections respectively,  Socket E is for  type 832 Valves, whilst sockets N and D   cover  the  miniature  types,  including  9000 series. Socket C is for acorn valves and sockets A and B cover European types of 5 and 7  pin  valves.   Positions  L,  M,  0 and P are provided to  permit  the  inclusion  of  any  future  types  of sockets.  This precludes the  possibility of the Tester becoming obsolete.  A circuit diagram of  the  adapter panel is provided at the end of the book.

NOTE:	If a valve is not listed, care must be taken to apply the correct filament voltage and  to  choose  that socket which has the correct filament connections for the valve to be listed,

ooOoo