## Flatness worksheet
FM Flatness test (note FG Amplitude = 2.0Vp-p)
Locked External FM Flatness
FG Freq | FM sensitivity | Vmodoff (dBm) | Vmodon (dBm) | Mod Index | FMflat (dB) | Specification (dB) | Measurement Uncertainty (dB)
10kHz | 10kHz | -3.4 | -6.3 | 1.1074 | -0.60865 | +/-1.0 | 0.07 | 1 | 1.1115 | 1.1074 | 1.1074 | 1.1074 | 1.1074 | 1.1074 | 1.1074 | 1.1074 | 1.1074 | 1.1074 | 1.1074 | 1.1074 | 1.1074 | 1.1074
20kHz | 20kHz | -3.4 | -6.28 | 1.1039 | -0.58121 | +/-1.0 | 0.07 | 1 | 1.1077 | 1.1039 | 1.1039 | 1.1039 | 1.1039 | 1.1039 | 1.1039 | 1.1039 | 1.1039 | 1.1039 | 1.1039 | 1.1039 | 1.1039 | 1.1039
50kHz | 50kHz | -3.42 | -6.05 | 1.0588 | -0.21939 | +/-1.0 | 0.07 | 1 | 1.0601 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588
99.8kHz | 99.8kHz | -3.38 | -5.87 | 1.0324 | reference | N/A | N/A | 1 | 1.0328 | 1.0324 | 1.0324 | 1.0324 | 1.0324 | 1.0324 | 1.0324 | 1.0324 | 1.0324 | 1.0324 | 1.0324 | 1.0324 | 1.0324 | 1.0324
200kHz | 200kHz | -3.4 | -5.97 | 1.0476 | -0.12695 | +/-1.0 | 0.07 | 1 | 1.0485 | 1.0476 | 1.0476 | 1.0476 | 1.0476 | 1.0476 | 1.0476 | 1.0476 | 1.0476 | 1.0476 | 1.0476 | 1.0476 | 1.0476 | 1.0476
500kHz | 500kHz | -3.4 | -5.8 | 1.015 | 0.14821 | +/-1.0 | 0.07 | 1 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015
1MHz | 1MHz | -3.38 | -5.88 | 1.0343 | -0.016109 | +/-1.0 | 0.07 | 1 | 1.0348 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343
9.98MHz** | 9.98MHz | -3.4 | -5.7 | 0.99507 | 0.3201 | +/-3.0 | 0.07 | 1 | 0.99508 | 0.99507 | 0.99507 | 0.99507 | 0.99507 | 0.99507 | 0.99507 | 0.99507 | 0.99507 | 0.99507 | 0.99507 | 0.99507 | 0.99507 | 0.99507
** bandwidth test included.
Narrow External Phase Modulation Flatness
FG Freq | Sensitivity | Vmodoff (dBm) | Vmodon (dBm) | Mod Index | PMflat (dB) | Specification (dB) | Measurement Uncertainty (dB)
10kHz | 1 | -3.4 | -5.78 | 1.011 | 0.033752 | +/-1.0 | 0.07 | 1 | 1.0111 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011
20kHz | 1 | -3.38 | -5.75 | 1.009 | 0.050743 | +/-1.0 | 0.07 | 1 | 1.0091 | 1.009 | 1.009 | 1.009 | 1.009 | 1.009 | 1.009 | 1.009 | 1.009 | 1.009 | 1.009 | 1.009 | 1.009 | 1.009
50kHz | 1 | -3.38 | -5.53 | 0.96423 | 0.44539 | +/-1.0 | 0.07 | 1 | 0.96471 | 0.96423 | 0.96423 | 0.96423 | 0.96423 | 0.96423 | 0.96423 | 0.96423 | 0.96423 | 0.96423 | 0.96423 | 0.96423 | 0.96423 | 0.96423
99.8kHz | 1 | -3.4 | -5.8 | 1.015 | reference | N/A | N/A | 1 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015
200kHz | 1 | -3.42 | -5.82 | 1.015 | 0 | +/-1.0 | 0.07 | 1 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015
500kHz | 1 | -3.38 | -5.48 | 0.95366 | 0.54113 | +/-1.0 | 0.07 | 1 | 0.95446 | 0.95366 | 0.95366 | 0.95366 | 0.95366 | 0.95366 | 0.95366 | 0.95366 | 0.95366 | 0.95366 | 0.95366 | 0.95366 | 0.95366 | 0.95366
1MHz | 1 | -3.4 | -5.53 | 0.96002 | 0.63161 | +/-1.0 | 0.07 | 1 | 0.96062 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002
10MHz** | 1 | -3.4 | -5.23 | 0.89381 | 1.2523 | +/-3.0 | 0.07 | 1 | 0.89812 | 0.89382 | 0.89381 | 0.89381 | 0.89381 | 0.89381 | 0.89381 | 0.89381 | 0.89381 | 0.89381 | 0.89381 | 0.89381 | 0.89381 | 0.89381
** bandwidth test included.
Wide External Phase Modulation Flatness
FG Freq | Sensitivity | Vmodoff (dBm) | Vmodon (dBm) | Mod Index | PMflat (dB) | Specification (dB) | Measurement Uncertainty (dB)
10kHz | 1 | -3.42 | -5.92 | 1.0343 | -0.13087 | +/-1.0 | 0.07 | 1 | 1.0348 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343 | 1.0343
20kHz | 1 | -3.4 | -5.87 | 1.0286 | -0.082334 | +/-1.0 | 0.07 | 1 | 1.0289 | 1.0286 | 1.0286 | 1.0286 | 1.0286 | 1.0286 | 1.0286 | 1.0286 | 1.0286 | 1.0286 | 1.0286 | 1.0286 | 1.0286 | 1.0286
50kHz | 1 | -3.4 | -5.82 | 1.0189 | 0 | +/-1.0 | 0.07 | 1 | 1.019 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189
99.8kHz | 1 | -3.4 | -5.82 | 1.0189 | reference | N/A | N/A | 1 | 1.019 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189
200kHz | 1 | -3.4 | -5.82 | 1.0189 | 0 | +/-1.0 | 0.07 | 1 | 1.019 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189 | 1.0189
500kHz | 1 | -3.4 | -5.8 | 1.015 | 0.03345 | +/-1.0 | 0.07 | 1 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015

## FM Accuracy worksheet
FM Flatness test (note FG Amplitude = 2.0Vp-p)
| FG Freq | FM sensitivity/PM Sensitivity | Vmodoff (dBm) | Vmodon (dBm) | Mod Index | Measured Deviation | FM/PM error% | Specification (%) | Measurement Uncertainty (%)
Locked External FM Accuracy at 5GHz | 99800 | 99800 | -3.67 | -6 | 1.0011 | 99909.2691 | 0.1095% | 10% | 1.5% | 1 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011
Locked low-noise External FM Accuracy at 5GHz | 99800 | 99800 | -3.67 | -6 | 1.0011 | 99909.2691 | 0.1095% | 10% | 1.5% | 1 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011 | 1.0011
Locked External FM Accuracy at 20GHz | 99800 | 99800 | -5.37 | -8 | 1.0588 | 105672.046 | 5.884% | 10% | 1.5% | 1 | 1.0601 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588
Locked low-noise External FM Accuracy at 20GHz | 99800 | 99800 | -6.1 | -8.55 | 1.0247 | 102266.5473 | 2.471% | 10% | 1.5% | 1 | 1.0249 | 1.0247 | 1.0247 | 1.0247 | 1.0247 | 1.0247 | 1.0247 | 1.0247 | 1.0247 | 1.0247 | 1.0247 | 1.0247 | 1.0247 | 1.0247
Locked internal FM Accuracy at 5GHz | 99800 | 99800 | -3.78 | -6.17 | 1.013 | 101097.0069 | 1.3% | 10% | 1.5% | 1 | 1.0131 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013
Locked low-noise internal FM Accuracy at 5GHz | 99800 | 99800 | -3.78 | -6.17 | 1.013 | 101097.0069 | 1.3% | 10% | 1.5% | 1 | 1.0131 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013 | 1.013
Locked internal FM Accuracy at 20GHz | 99800 | 99800 | -5.68 | -8.32 | 1.0607 | 105856.9107 | 6.069% | 10% | 1.5% | 1 | 1.062 | 1.0607 | 1.0607 | 1.0607 | 1.0607 | 1.0607 | 1.0607 | 1.0607 | 1.0607 | 1.0607 | 1.0607 | 1.0607 | 1.0607 | 1.0607
Locked low-noise internal FM Accuracy at 20GHz | 99800 | 99800 | -5.6 | -7.98 | 1.011 | 100900.3317 | 1.103% | 10% | 1.5% | 1 | 1.0111 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011 | 1.011
Wide External Phase Mod Accuracy at 5GHz | 99800 | 1 | -3.97 | -6.4 | 1.0208 | 101878.6831 | 2.04% | 10% | 1.5% | 1 | 1.021 | 1.0208 | 1.0208 | 1.0208 | 1.0208 | 1.0208 | 1.0208 | 1.0208 | 1.0208 | 1.0208 | 1.0208 | 1.0208 | 1.0208 | 1.0208
Narrow External Phase Mod Accuracy at 5GHz | 99800 | 1 | -3.82 | -5.95 | 0.96002 | 95809.8575 | 4.165% | 10% | 1.5% | 1 | 0.96062 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002 | 0.96002
Wide External Phase Mod Accuracy at 20GHz | 99800 | 1 | -5.5 | -8.13 | 1.0588 | 105672.046 | 5.557% | 10% | 1.5% | 1 | 1.0601 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588 | 1.0588
Narrow External Phase Mod Accuracy at 20GHz | 99800 | 1 | -5.45 | -7.65 | 0.97465 | 97270.1855 | 2.601% | 10% | 1.5% | 1 | 0.97489 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465
Wide Internal Phase Mod Accuracy at 5GHz | 99800 | 1 | -3.77 | -6.17 | 1.015 | 101293.1767 | 1.474% | 10% | 1.5% | 1 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015 | 1.015
Narrow Internal Phase Mod Accuracy at 5GHz | 99800 | 1 | -3.7 | -5.77 | 0.94724 | 94535.0003 | 5.569% | 10% | 1.5% | 1 | 0.94829 | 0.94724 | 0.94724 | 0.94724 | 0.94724 | 0.94724 | 0.94724 | 0.94724 | 0.94724 | 0.94724 | 0.94724 | 0.94724 | 0.94724 | 0.94724
Wide Internal Phase Mod Accuracy at 20GHz | 99800 | 1 | -5.82 | -8.23 | 1.0169 | 101488.844 | 1.664% | 10% | 1.5% | 1 | 1.017 | 1.0169 | 1.0169 | 1.0169 | 1.0169 | 1.0169 | 1.0169 | 1.0169 | 1.0169 | 1.0169 | 1.0169 | 1.0169 | 1.0169 | 1.0169
Narrow Internal Phase Mod Accuracy at 20GHz | 99800 | 1 | -5.85 | -8.05 | 0.97465 | 97270.1855 | 2.601% | 10% | 1.5% | 1 | 0.97489 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465 | 0.97465

## mod-index-calculations.txt
Vmodon = dBm measurement of the CW with Modulation on
Vmodoff = dBm measurement of the CW with Modulation off
Constant_A= 1
RESULT_FINAL = Modulation Index
RESULT_1 = CONSTANT_A-(1-CONSTANT_A^2/4+CONSTANT_A^4/64-CONSTANT_A^6/2304+CONSTANT_A^8/147456-CONSTANT_A^10/14745600-10^((Vmodon-Vmodoff)/20))/(-CONSTANT_A/2+CONSTANT_A^3/16-CONSTANT_A^5/384+CONSTANT_A^7/18432-CONSTANT_A^9/1474560)
RESULT_2 = RESULT_1-(1-RESULT_1^2/4+RESULT_1^4/64-RESULT_1^6/2304+RESULT_1^8/147456-RESULT_1^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_1/2+RESULT_1^3/16-RESULT_1^5/384+RESULT_1^7/18432-RESULT_1^9/1474560)
RESULT_3 = RESULT_2-(1-RESULT_2^2/4+RESULT_2^4/64-RESULT_2^6/2304+RESULT_2^8/147456-RESULT_2^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_2/2+RESULT_2^3/16-RESULT_2^5/384+RESULT_2^7/18432-RESULT_2^9/1474560)
RESULT_4 = RESULT_3-(1-RESULT_3^2/4+RESULT_3^4/64-RESULT_3^6/2304+RESULT_3^8/147456-RESULT_3^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_3/2+RESULT_3^3/16-RESULT_3^5/384+RESULT_3^7/18432-RESULT_3^9/1474560)
RESULT_5 = RESULT_4-(1-RESULT_4^2/4+RESULT_4^4/64-RESULT_4^6/2304+RESULT_4^8/147456-RESULT_4^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_4/2+RESULT_4^3/16-RESULT_4^5/384+RESULT_4^7/18432-RESULT_4^9/1474560)
RESULT_6 = RESULT_5-(1-RESULT_5^2/4+RESULT_5^4/64-RESULT_5^6/2304+RESULT_5^8/147456-RESULT_5^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_5/2+RESULT_5^3/16-RESULT_5^5/384+RESULT_5^7/18432-RESULT_5^9/1474560)
RESULT_7 = RESULT_6-(1-RESULT_6^2/4+RESULT_6^4/64-RESULT_6^6/2304+RESULT_6^8/147456-RESULT_6^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_6/2+RESULT_6^3/16-RESULT_6^5/384+RESULT_6^7/18432-RESULT_6^9/1474560)
RESULT_8 = RESULT_7-(1-RESULT_7^2/4+RESULT_7^4/64-RESULT_7^6/2304+RESULT_7^8/147456-RESULT_7^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_7/2+RESULT_7^3/16-RESULT_7^5/384+RESULT_7^7/18432-RESULT_7^9/1474560)
RESULT_9 = RESULT_8-(1-RESULT_8^2/4+RESULT_8^4/64-RESULT_8^6/2304+RESULT_8^8/147456-RESULT_8^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_8/2+RESULT_8^3/16-RESULT_8^5/384+RESULT_8^7/18432-RESULT_8^9/1474560)
RESULT_10 = RESULT_9-(1-RESULT_9^2/4+RESULT_9^4/64-RESULT_9^6/2304+RESULT_9^8/147456-RESULT_9^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_9/2+RESULT_9^3/16-RESULT_9^5/384+RESULT_9^7/18432-RESULT_9^9/1474560)
RESULT_11 = RESULT_10-(1-RESULT_10^2/4+RESULT_10^4/64-RESULT_10^6/2304+RESULT_10^8/147456-RESULT_10^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_10/2+RESULT_10^3/16-RESULT_10^5/384+RESULT_10^7/18432-RESULT_10^9/1474560)
RESULT_12 = RESULT_11-(1-RESULT_11^2/4+RESULT_11^4/64-RESULT_11^6/2304+RESULT_11^8/147456-RESULT_11^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_11/2+RESULT_11^3/16-RESULT_11^5/384+RESULT_11^7/18432-RESULT_11^9/1474560)
RESULT_13 = RESULT_12-(1-RESULT_12^2/4+RESULT_12^4/64-RESULT_12^6/2304+RESULT_12^8/147456-RESULT_12^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_12/2+RESULT_12^3/16-RESULT_12^5/384+RESULT_12^7/18432-RESULT_12^9/1474560)
RESULT_FINAL = RESULT_13-(1-RESULT_13^2/4+RESULT_13^4/64-RESULT_13^6/2304+RESULT_13^8/147456-RESULT_13^10/14745600-10^((Vmodon-Vmodoff)/20))/(-RESULT_13/2+RESULT_13^3/16-RESULT_13^5/384+RESULT_13^7/18432-RESULT_13^9/1474560)