Burster Digistant Type 4462

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for Burster Digistant Type 4462.

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

For GPIB instrument drivers:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

For VXI instrument drivers:

-  the instrument is installed in the VXI mainframe and you are

using one of the following controller options:

Embedded controller

MXI

MXI2

GPIB-VXI

-  the logical address supplied to the initialize function must

match the logical address of the instrument.

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

________________________________________________________________________________

Error and Status Information:

Each function in this instrument driver returns a status code that

either indicates success or describes an error or warning condition.

Your program should examine the status code from each call to an

instrument driver function to determine if an error occurred.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

The description of each instrument driver function lists possible error

codes and their meanings

________________________________________________________________________________

How To Use This Document:

Use this document as a programming reference manual.

It describes each function in the

Burster Digistant Type 4462

instrument.  The functions appear in alphabetical order, with a description

of the function and its Visual Basic syntax, a description of each parameter,

and a list of possible error codes.

________________________________________________________________________________

Function Tree Layout:

Class/Panel Name:                              Function Name:

Initialize

Initialize GPIB                          bp4462_InitGpib

Initialize Serial                        bp4462_InitSerial

Configure

Ramp

Ramp 1

Set Config Ramp1                   bp4462_SetConfigRamp1

Get Config Ramp1                   bp4462_GetConfigRamp1

Ramp 2

Set Config List Ramp2              bp4462_SetConfigListRamp2

Get Config List Ramp2              bp4462_GetConfigListRamp2

Set Config Ramp2                   bp4462_SetConfigRamp2

Get Config Ramp2                   bp4462_GetConfigRamp2

Set Ramp Step                         bp4462_SetRampStep

Set Ramp Start/Stop                   bp4462_SetRampStart

Range

Get Config Range Mode                 bp4462_GetConfigRangeMode

Set Config Range Mode                 bp4462_SetConfigRangeMode

RJ

Get RJ Temperature Value Ext.         bp4462_GetRjTempValue

Set RJ Temperat. Value Manual         bp4462_SetRjTempValueMan

Get RJ Temperat. Value Manual         bp4462_GetRjTempValueMan

Set Config RJ Type                    bp4462_SetConfigRjType

Get Config RJ Type                    bp4462_GetConfigRjType

TC

Set Config TC Unit                    bp4462_SetConfigTcUnit

Set Config TC                         bp4462_SetConfigTc

Get Config TC                         bp4462_GetConfigTc

Get TC Voltage Extern                 bp4462_GetTcVoltageExtern

PT 100

Get Din Coefficient PT100             bp4462_GetConfigDinCoefficient

Get Config Coefficient PT100          bp4462_GetConfigCoefficient

Set Config Coefficient PT100          bp4462_SetConfigCoefficient

Output

Limit

Get Config Output Limit            bp4462_GetConfigOutputLimit

Set Config Output Limit Volt       bp4462_SetConfigOutputLimitV

Set Config Output Limit Current    bp4462_SetConfigOutputLimitC

Status / Values

Get Config Output Status           bp4462_GetConfigOutputStatus

Set Config Output Status           bp4462_SetConfigOutputStatus

Mode

Set Config Output Mode             bp4462_SetConfigOutputMode

Get Config Output Mode             bp4462_GetConfigOutputMode

Sense

Set Config Sense                      bp4462_SetConfigSense

Get Config Sense                      bp4462_GetConfigSense

Config Display

Set Config Display Mode               bp4462_SetConfigDisplayMode

Set Config Display Contrast           bp4462_SetConfigDisplayContrast

Get Config Display Contrast           bp4462_GetConfigDisplayContrast

Device Access

Get Config Device Access              bp4462_GetConfigDeviceAccess

Set Config Device Access              bp4462_SetConfigDeviceAccess

Voltage Divider

Get Config Voltage Div. State         bp4462_GetConfigVoltageDivState

Set Config Voltage Div. State         bp4462_SetConfigVoltageDivState

Set Config Voltage Divider            bp4462_SetConfigVoltageDiv

Delta Value

Get Config Delta Value                bp4462_GetConfigDeltaValue

Set Config Delta Value                bp4462_SetConfigDeltaValue

Set Delta Value Up/Down               bp4462_SetAddDeltaValue

System

Check TC Option                          bp4462_CheckTcOption

Set Lock System Keyboard                 bp4462_SetLockSystemKeyboard

Get Lock System Keyboard                 bp4462_GetLockSystemKeyboard

Device Error

Check Device Err. Status Output       bp4462_CheckErrorStatusOutput

Check Device Err. Status Temp.        bp4462_CheckErrorStatusTemp

Check Device Err. Status Source       bp4462_CheckErrorStatusSource

Get System Error                         bp4462_GetSystemError

SCPI

Receive a SCPI-Message                bp4462_receive_scpi

Send a SCPI-Message                   bp4462_send_scpi

Register

Read Standard Event Status Reg        bp4462_ReadStdEventStatusReg

Read Operation Status Event Reg       bp4462_ReadOperStatusEventReg

Read Operation Status Cond. Reg       bp4462_ReadOperStatusCondReg

Read Quest.Status Event Regist.       bp4462_ReadQuestStatusEventReg

Read Quest. Status Output Reg         bp4462_ReadQuestStatusOutputReg

Read Quest Status Temperat. Reg       bp4462_ReadQuestStatusTempReg

Reset Quest./Oper.Stat.En.Reg.        bp4462_ResetQuestOperStEnReg

Store/Load

Set Memory Label                         bp4462_SetMemoryLabel

Get Memory Label                         bp4462_GetMemoryLabel

Set Store Output Value                   bp4462_SetStoreOutputValue

Get Store Output Value                   bp4462_GetStoreOutputValue

Set Store ID Output Value                bp4462_SetStoreIdOutputValue

Close                                       bp4462_close

________________________________________________________________________________

Burster Digistant Type 4462

Instrument Name: burster DIGISTANT Type 4462

Description:   This instrument module provides

RS232 and IEEE488.2 support for

DIGISTANT Type 4462 from burster.

Functions and Classes:

(1) Initialize:

Initialize the instrument and sets to

default configuration

(2) Configure:

Configures the instrument by setting the

Voltage, Current, TC, RJ and Output a Value

(3) System:

Reads the system error

Reads the instrument ID-String

Controls the keyboard

Sends a SCPI-Message

Receives a SCPI-Message

(4) Store/Load:

Stores and loads values to and from device.

(5) Close:

Takes the instrument off-line.

(6) Calibration:

Special functions for the manufacturer to calibrate

the device.

*********

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

bp4462_CheckErrorStatusOutput

Function bp4462_CheckErrorStatusOutput& (ByVal instrumentID&)

Purpose

Panel Name:    Check Device Error Status Output

Description:   Check the questionable status register and

returns a error-ID if an error occurs.

Checked errors:

- Voltage output error

- Current output error

- Analog 15V supply error

- Temperature TC output error

- Calibrating error

- PT100 measuring error

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

310   Error Voltage Output

311   Error Current Output

312   Error Device Voltage 15 Volt Analog

313   Error Temperature Output

314   Error Calibration

315   Error PT 100 Measure

________________________________________________________________________________

bp4462_CheckErrorStatusSource

Function bp4462_CheckErrorStatusSource& (ByVal instrumentID&)

Purpose

Panel Name:    Check Device Error Status Source

Description:   Checks the status of the output

in the questionable status output register

and returns an error-ID if an error occurs.

Checked errors:

- Overrange of the end stage

- PT100 source error

- Overrange of the D/A converter

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

316   Error Output Overrange

317   Error PT 100

318   Error D/A-Converter Overrange

________________________________________________________________________________

bp4462_CheckErrorStatusTemp

Function bp4462_CheckErrorStatusTemp& (ByVal instrumentID&)

Purpose

Panel Name:    Check Device Error Status Temperature

Description:   Checks the status of a temperature measurement

in the questionable status temperature register

and returns an error-ID if an error occurs.

Checked errors:

- Break of a cable by current measuring

- Break of a cable by voltage measuring

- Resistor (PT100) out of range

- Invalid temperature

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

320   Current Cable Error

321   Voltage Cable Error

322   PT 100 Out of Range

323   Invalid Temperature

________________________________________________________________________________

bp4462_CheckTcOption

Function bp4462_CheckTcOption& (ByVal instrumentID&, optionStatus&)

Purpose

Panel Name:    Check TC Option

Description:   Check if a TC-option is integrated in device.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

optionStatus&

Control Name:  Option Status

Description:   Displays the result of TC-option

integrated in device.

Variable Type: Integer

Valid Range:   0 to 1

0 = No TC-option integrated

1 = TC-option integrated

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_close

Function bp4462_close& (ByVal instrumentID&)

Purpose

Panel Name:    Close Serial

Description:   Takes the instrument off-line.

Note: The  instrument must  be reinitialized

to use it again.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

________________________________________________________________________________

bp4462_GetConfigCoefficient

Function bp4462_GetConfigCoefficient& (ByVal instrumentID&, r0#, coefficientA#, coefficientB#, coefficientC#, t100#)

Purpose

Panel Name:    Get Config Coefficient PT100

Description:   Gets the values of the PT100 coefficients.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

r0#

Control Name:     R0

Description:      Displays the value of R0.

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

coefficientA#

Control Name:     Coefficient "A"

Description:      Displays the value of Coefficient "A".

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

coefficientB#

Control Name:     Coefficient "B"

Description:      Displays the value of Coefficient "B".

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

coefficientC#

Control Name:     Coefficient "C"

Description:      Displays the value of Coefficient "C".

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

t100#

Control Name:     t100

Description:      Displays the value of t100.

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigDeltaValue

Function bp4462_GetConfigDeltaValue& (ByVal instrumentID&, ByVal type&, unit&, value#)

Purpose

Panel Name:    Get Config Delta Value

Description:   Gets the delta value and unit

of the selected type.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal type&

Control Name:   Type

Description:    Specifies the type of delta value.

Variable Type:  Integer

Valid Range:    0 to 2

0 = Voltage

1 = Current

3 = Temperature

Default Value:  0 = Voltage

unit&

Control Name:   Unit

Description:    Displays the unit of the delta value.

Variable Type:  Integer

Valid Range:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

value#

Control Name:   Value

Description:    Displays the delta value

of the selected type.

Variable Type:  Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Type

-3   Error Unit return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigDeviceAccess

Function bp4462_GetConfigDeviceAccess& (ByVal instrumentID&, speechConfig&, outputFunction&, rampFunction&, ramp1&, ramp2&, deltaValue&, storeConfig&, recallConfig&, senseConfig&, outputLimit&, rangeMode&, divideConfig&,

TCConfig&, RJConfig&, PT100Scale&, deviceStore&, deviceRecall&, interface&, contrast&, statusDisplay&)

Purpose

Panel Name:    Get Config Device Access

Description:   Gets the access configuration of the device.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

speechConfig&

Control Name:  Speech Config

Description:   This parameter gets the state of

speech configuration access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

outputFunction&

Control Name:  Output Function

Description:   This parameter gets the state of

output function access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

rampFunction&

Control Name:  Ramp Function

Description:   This parameter gets the state of the

ramp function access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

ramp1&

Control Name:  Ramp1

Description:   This parameter gets the state of the

ramp1 menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

ramp2&

Control Name:  Ramp2

Description:   This parameter gets the state of the

ramp2 menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

deltaValue&

Control Name:  Delta Value

Description:   This parameter gets the state of the

delta value menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

storeConfig&

Control Name:  Store Config

Description:   This parameter gets the state of the

store config menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

recallConfig&

Control Name:  Recall Config

Description:   This parameter gets the state of the

recall configuration menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

senseConfig&

Control Name:  Sense Config

Description:   This parameter gets the state of the

sense config menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

outputLimit&

Control Name:  Output Limit

Description:   This parameter gets the state of the

output limit menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

rangeMode&

Control Name:  Range Mode

Description:   This parameter gets the state of the

range mode menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

divideConfig&

Control Name:  Divide Config

Description:   This parameter gets the state of the

divide configuration menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

TCConfig&

Control Name:  TC-Config

Description:   This parameter gets the state of the

TC-configuration menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

RJConfig&

Control Name:  RJ-Config

Description:   This parameter gets the state of the

RJ-configuration menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

PT100Scale&

Control Name:  PT100 Scale

Description:   This parameter gets the state of the

PT100 scale menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

deviceStore&

Control Name:  Device Store

Description:   This parameter gets the state of the

device store menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

deviceRecall&

Control Name:  Device Recall

Description:   This parameter gets the state of the

device recall menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

interface&

Control Name:  Interface

Description:   This parameter gets the state of the

interface menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

contrast&

Control Name:  Contrast

Description:   This parameter gets the state of the

contrast menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

statusDisplay&

Control Name:  Status Display

Description:   This parameter gets the state of the

status display menu access.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigDinCoefficient

Function bp4462_GetConfigDinCoefficient& (ByVal instrumentID&, r0#, coefficientA#, coefficientB#, coefficientC#, t100#)

Purpose

Panel Name:    Get Din Coefficient PT100

Description:   Gets the values of the DIN EN PT100

coefficients.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

r0#

Control Name:     R0

Description:      Displays the value of R0.

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

coefficientA#

Control Name:     Coefficient "A"

Description:      Displays the value of Coefficient "A".

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

coefficientB#

Control Name:     Coefficient "B"

Description:      Displays the value of Coefficient "B".

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

coefficientC#

Control Name:     Coefficient "C"

Description:      Displays the value of Coefficient "C".

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

t100#

Control Name:     t100

Description:      Displays the value of t100.

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigDisplayContrast

Function bp4462_GetConfigDisplayContrast& (ByVal instrumentID&, displayContrast#)

Purpose

Panel Name:    Get Config Display Contrast

Description:   Gets the contrast value of the display.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

displayContrast#

Control Name:  Display Contrast

Description:   Displays the display contrast value.

Variable Type: Double

Valid Range:   0.0 - 1.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigListRamp2

Function bp4462_GetConfigListRamp2& (ByVal instrumentID&, ByVal listNo&, unit&, value#, stepTimeh&, stepTimemin&, stepTimesec#)

Purpose

Panel Name:    Get Config List Ramp2

Description:   Gets the ramp function 2 configuration

of the instrument in following

ways :

- Output Unit

- List No.

- List Value

- List Step Time

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal listNo&

Control Name:   List No.

Description:    Specifies the List No. of the

ramp2 configuration

Variable Type:  Integer

Valid Range:    1 to 31

Default:        1

unit&

Control Name:   Unit

Description:    Specifies the output type and the

unit of start, stop and step

value

Variable Type:  Integer

Valid Units:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

value#

Control Name:   Value

Description:    Specifies the value of the

selected List No.

Variable Type:  Double

Valid Range:    by Voltage -30 to 30 V

by Current -52 to 52 mA

by TC    -3000 to 3000 K

stepTimeh&

Control Name:   Step Time[h]

Description:    Displays the hours of step time

in the selected List No.

Variable Type:  Integer

stepTimemin&

Control Name:   Step Time[min]

Description:    Displays the minutes of step time

in the selected List No.

Variable Type:  Integer

stepTimesec#

Control Name:   Step Time[sec]

Description:    Displays the seconds of step time

in the selected List No.

Variable Type:  Integer

Valid Range:    0.0 to 59.9

Default:        1.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Error Unit return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigOutputLimit

Function bp4462_GetConfigOutputLimit& (ByVal instrumentID&, voltageLimitValue&, currentLimitValue&)

Purpose

Panel Name:    Get Config Output Limit

Description:   Gets the specified voltage- and current-

limit values.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

voltageLimitValue&

Control Name:   Voltage Limit Value

Description:    Displays the voltage limit value

at the current output state.

Variable Type:  Integer

Valid Range:    Voltage 1 to 30 V

currentLimitValue&

Control Name:   Current Limit Value

Description:    Displays the current limit value

at the voltage output state.

Variable Type:  Integer

Valid Range:    Current 1 to 52 mA

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigOutputMode

Function bp4462_GetConfigOutputMode& (ByVal instrumentID&, mode&)

Purpose

Panel Name:    Get Config Output Mode

Description:   Gets the specified mode of the instrument in

following ways:

- Normal mode

- Ramp 1 mode

- Ramp 2 mode

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

mode&

Control Name:   Mode

Description:    Displays the Ramp Mode of the device.

Variable Type:  Integer

Valid Values:   0 = OFF / Normal

1 = Ramp1

2 = Ramp2

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Error Ramp Mode return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigOutputStatus

Function bp4462_GetConfigOutputStatus& (ByVal instrumentID&, unit&, value#)

Purpose

Panel Name:    Get Config Output Status

Description:   Gets the configuration of the

output value and unit.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

unit&

Control Name:   Unit

Description:    Displays the unit of the output value.

Variable Type:  Integer

Valid Range:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

value#

Control Name:   Value

Description:    Displays the output value.

Variable Type:  Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Error Unit return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigRamp1

Function bp4462_GetConfigRamp1& (ByVal instrumentID&, waveform&, numberOfCounts&, stepTimeh&, stepTimemin&, stepTimesec#, stepUnit&, stepValue#, startUnit&, startValue#, stopUnit&, stopValue#)

Purpose

Panel Name:    Get Config Ramp1

Description:   Gets the ramp function 1 configuration

of the instrument in following

ways :

- Selected waveform

- Number of counts

- Step time

- Output type and unit

- Start value

- End value

- Step value

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

waveform&

Control Name:   Waveform

Description:    Displays the waveform of ramp1

Variable Type:  Integer

Valid Values:   0 = Sawtooth

1 = Triangle

Default Value:  0 = Sawtooth

numberOfCounts&

Control Name:   Number of Counts

Description:    Displays the number of counts

of ramp1

Variable Type:  Integer

Valid Range:    0 to 99

0 = Countiue

Default:        1

stepTimeh&

Control Name:   Step Time (hours)

Description:    Displays the hours of step time

Variable Type:  Integer

stepTimemin&

Control Name:   Step Time (minutes)

Description:    Displays the minutes of step time

Variable Type:  Integer

stepTimesec#

Control Name:   Step Time (seconds)

Description:    Displays the seconds of step time

Variable Type:  Double

stepUnit&

Control Name:   Step Unit

Description:    Displays the step unit of the

ramp1 step value

Variable Type:  Integer

Valid Units:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

stepValue#

Control Name:   Step Value

Description:    Displays the step value of ramp1

Variable Type:  Double

startUnit&

Control Name:   Start Unit

Description:    Displays the start unit of the

ramp1 start value

Variable Type:  Integer

Valid Units:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

startValue#

Control Name:   Start Value

Description:    Displays the start value of

ramp1

Variable Type:  Double

stopUnit&

Control Name:   Stop Unit

Description:    Displays the stop unit of the

ramp1 stop value

Variable Type:  Integer

Valid Units:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

stopValue#

Control Name:   Stop Value

Description:    Displays the stop value of

ramp1

Variable Type:  Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Error Waveform return

-3   Error Start Unit return

-4   Error Stop Unit return

-5   Error Step Unit return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigRamp2

Function bp4462_GetConfigRamp2& (ByVal instrumentID&, ByVal listNo&, waveform&, numberOfCounts&, numberOfPoints&)

Purpose

Panel Name:    Get Config Ramp2

Description:   Gets the ramp function 2 configuration

of the instrument in following

ways :

- Waveform

- Number of counts

- Number of points

- List No.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal listNo&

Control Name:   List No

Description:    Specifies the List No. of the

ramp2 configuration

Variable Type:  Integer

Valid Range:    1 to 31

Default:        1

waveform&

Control Name:   Waveform

Description:    Displays the waveform of ramp2

Variable Type:  Integer

Values:         0 = Sawtooth

1 = Triangle

numberOfCounts&

Control Name:   Number of Counts

Description:    Displays the number of counts

of ramp2

Variable Type:  Integer

numberOfPoints&

Control Name:   Number of Points

Description:    Displays the number of points

of ramp2

Variable Type:  Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Error Waveform return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigRangeMode

Function bp4462_GetConfigRangeMode& (ByVal instrumentID&, rangeMode&)

Purpose

Panel Name:    Get Config Range Mode

Description:   Gets the configuration of the Range Mode

in following ways:

- Automatic

- Manual

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

rangeMode&

Control Name:  Range

Description:   Displays the range mode and the range value

if range mode is set to manual range.

Variable Type: Integer

Valid Range:   0 = Autorange

1 = 300 mVolt

2 = 3 Volt

3 = 30 Volt

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Range

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigRjType

Function bp4462_GetConfigRjType& (ByVal instrumentID&, RJType&)

Purpose

Panel Name:    Get Config RJ Type

Description:   Gets the configured RJ-Type in following

ways:

- RJ-External

- RJ-Manual

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

RJType&

Control Name:  RJ-Type

Description:   Displays the mode of reference junction type.

Variable Type: Integer

Valid Range:   0 to 1

0 = Manual Reference Junction Type

1 = External Reference Junction Type

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid RJ-Type return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigSense

Function bp4462_GetConfigSense& (ByVal instrumentID&, senseMode&)

Purpose

Panel Name:    Get Config Sense

Description:   Gets the configuration of the sense in

following ways:

- External

- Internal

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

senseMode&

Control Name:   Mode

Description:    Displays the Sense Mode of the device.

Variable Type:  Integer

Valid Values:   0 = Sense Internal

1 = Sense External

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Error Sense Mode return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigTc

Function bp4462_GetConfigTc& (ByVal instrumentID&, unit&, type&, scale&)

Purpose

Panel Name:    Get Config TC

Description:   Gets the configuration of TC.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

unit&

Control Name:  Unit

Description:   Displays the tc-temperature unit.

Variable Type: Integer

Valid Range:   5 to 7

5 = C (Celsius)

6 = F (Fahrenheit)

7 =  K (Kelvin)

type&

Control Name:   Type

Description:    Displays the TC-Type.

Variable Type:  Integer

Valid Values:   0 = R

1 = S

2 = B

3 = J

4 = T

5 = E

6 = K

7 = U

8 = L

9 = N

10 = C

11 = D

12 = G2

13 = M

scale&

Control Name:  Scale

Description:   Displays the tc-temperature scale.

Variable Type: Integer

Valid Range:   0 to 1

0 = IPTS 68

1 = ITS 90

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Error TC-Type return

-3   Error TC-Scale return

-4   Error TC-Unit return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetConfigVoltageDivState

Function bp4462_GetConfigVoltageDivState& (ByVal instrumentID&, dividerMode&, dividerFactor#)

Purpose

Panel Name:    Get Config Voltage Divider State

Description:   Gets the mode of the voltage divider.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

dividerMode&

Control Name:   Divider Mode

Description:    Displays the divider mode of the device.

Variable Type:  Integer

Valid Range:    0 to 1

0 = Divider OFF

1 = Divider ON

dividerFactor#

Control Name:   Factor

Description:    Displays the factor value

of the voltage dividor.

Variable Type:  Double

Valid Range:    1.0 to 1010.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetLockSystemKeyboard

Function bp4462_GetLockSystemKeyboard& (ByVal instrumentID&, lockState&)

Purpose

Panel Name:    Get Lock System Keyboard

Description:   Gets the lock state of the System Keyboard.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

lockState&

Control Name:  Lock-State

Description:  Displays the lock-state for

the front panel of the instrument.

Variable Type: Integer

Valid Range:   0  = Unlocked

1  = Locked

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetMemoryLabel

Function bp4462_GetMemoryLabel& (ByVal instrumentID&, ByVal memoryNo&, ??memoryLabel??)

Purpose

Panel Name:   Get Memory Label

Description:  Gets the label text of a

selected device memory ID.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal memoryNo&

Control Name:  Memory No.

Description:   Specificies the memory No. to the

recalled label text in the device.

Variable Type: Integer

Valid Range:   1 -  33

??memoryLabel??

Control Name:  Memory Label

Description:   Displays the memory label text.

Variable Type: String (10 Byte + NULL)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Memory ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetRjTempValue

Function bp4462_GetRjTempValue& (ByVal instrumentID&, unit&, value#)

Purpose

Panel Name:    Get RJ Temperature Value Extern

Description:   Gets the temperature value and unit of the

external reference junction.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

unit&

Control Name:  Unit

Description:   Displays the temperature unit of the

rj-temperature value.

Variable Type: Integer

Valid Range:   5 to 7

5 = C (Celsius)

6 = F (Farenheit)

7 =  K (Kelvin)

value#

Control Name:   Value

Description:    Displays the rj-temperature value.

Variable Type:  Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Unit return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetRjTempValueMan

Function bp4462_GetRjTempValueMan& (ByVal instrumentID&, unit&, value#)

Purpose

Panel Name:    Get RJ Temperature Value Manual

Description:   Gets the temperature value and unit of the

manual adjusted reference junction.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

unit&

Control Name:  Temperature Unit

Description:   Displays the temperature unit of the

rj-temperature value.

Variable Type: Integer

Valid Range:   5 to 7

5 = C (Celsius)

6 = F (Farenheit)

7 =  K (Kelvin)

value#

Control Name:   Value

Description:    Displays the rj-temperature value.

Variable Type:  Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetStoreOutputValue

Function bp4462_GetStoreOutputValue& (ByVal instrumentID&, ByVal memoryNo&, outputUnit&, outputValue#, deltaUnit&, deltaValue#)

Purpose

Panel Name:    Get Store Output Value

Description:   Displays the stored output and delta value

under the selected memory No. with value unit.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal memoryNo&

Control Name:   Memory No.

Description:    Specifies the Memory No. of the stored

output and delta value.

Variable Type:  integer

Valid Range:    0 to 99

Default:        0

outputUnit&

Control Name:   Output Unit

Description:    Displays the unit of the output value.

Variable Type:  Integer

Valid Range:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

outputValue#

Control Name:   Output Value

Description:    Displays the output value

of the selected memory No.

Variable Type:  Double

deltaUnit&

Control Name:   Delta Unit

Description:    Displays the unit of the delta value.

Variable Type:  Integer

Valid Range:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

deltaValue#

Control Name:   Delta Value

Description:    Displays the delta value

of the selected memory No.

Variable Type:  Double

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Memory No.

-3   Error Output Unit return

-4   Error Delta Unit return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_GetSystemError

Function bp4462_GetSystemError& (ByVal instrumentID&)

Purpose

Panel Name:    Get System Error

Description:   Gets an actually error ID from device

if an error occours.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success, No Error

-1   Invalid Instrument ID

-100   Command error

-101   Invalid character in the SCPI-command

-105   Get-command not allowed

-109   Missing Parameter in the SCPI-syntax

-110   Command header error

-120   Invalid numeric data error

-200   Execution error

-204   Illegal device state

-213   No init command allowed

-220   SCPI-command contained invalid parameter

-221   Setting conflict

-222   Data out of range

-224   Illegal parameter value

-231   Data value is questionable

-350   Query overflow

-400   Query error

-410   Query interrupted

-420   Query unterminated

110   Voltage overrange

210   Current overrange

300   15 Volt analogue power failure

510   Temperature overrange

520   PT 100 error

602   Out over

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

________________________________________________________________________________

bp4462_GetTcVoltageExtern

Function bp4462_GetTcVoltageExtern& (ByVal instrumentID&, valueU_T#, unitU_T&, valueU_0C#, unitU_0C&)

Purpose

Panel Name:    Get TC Voltage Extern

Description:   Gets the external TC-voltage values U(T)

and U(0C).

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

valueU_T#

Control Name:  Value U(T)

Description:   Displays the voltage-value of the

external tc-temperature U(T).

Variable Type: Double

unitU_T&

Control Name:  Unit U(T)

Description:   Displays the voltage-unit of the

external tc-temperature value U(T).

Variable Type: Integer

Valid Range:   0 to 2

0 = Volt

1 = m Volt

2 = Volt

valueU_0C#

Control Name:  Value U(0C)

Description:   Displays the voltage-value of the

external tc-temperature U(0C).

Variable Type: Double

unitU_0C&

Control Name:  Unit U(0C)

Description:   Displays the voltage-unit of the

external tc-temperature value U(0C).

Variable Type: Integer

Valid Range:   0 to 2

0 = Volt

1 = m Volt

2 = Volt

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Unit U(T) return

-3   Invalid Unit U(0C) return

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_InitGpib

Function bp4462_InitGpib& (ByVal adress&, ByVal reset&, instrumentID&)

Purpose

Panel Name:    Initialize Gpib

Description:   Initializes the instrument by:

1) finding an unused device in the device table

using the supplied GPIB address

2) device clear and RESET  (if desired)

3) returns the ID-String of the instrument

4) returns an instrument ID

This is a multi-instrument driver, allowing

multiple Digistant 4462 instruments to be

controlled.

During the initialization, the returned ID is

based on the GPIB address.  This ID is used in

all of the other functions to distinguish

between multiple Digistant 4462 instruments.

Parameter List

ByVal adress&

ByVal reset&

Control Name:  Reset

Description:   If reset is on, a device clear and

power on reset is done during initialization.

Variable Type: Integer

Valid Range:   0 to 1

0  off

1  on

instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid GPIB Address

-2     Invalid Reset

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

________________________________________________________________________________

bp4462_InitSerial

Function bp4462_InitSerial& (ByVal COMPort&, ByVal baudRate&, ByVal dataConfig&, ByVal reset&, instrumentID&)

Purpose

Panel Name:    Initialize Serial

Description:   Initializes the instrument by:

1) select the used COM-Port

2) device clear and RESET  (if desired)

3) configurate the interface

Parameter List

ByVal COMPort&

Control Name:   COM-Port

Description:    Specifies the COM-Port to be

used to communicate with the

instrument.

Variable Type:  Integer

Valid Values:   1 - COM1   (=> Adress 0x3f8 IRQ4)

2 - COM2   (=> Adress 0x2f8 IRQ3)

3 - COM3   (=> Adress 0x3e8 IRQ4)

4 - COM4   (=> Adress 0x2e8 IRQ3)

Default:        1

ByVal baudRate&

Control Name:   Baud Rate

Description:    Specifies the baud rate of the

RS232-Port to communicate with

the instrument. This value must

correspond to the current

internal setting of the

instrument.

Variable Type:  Integer

Valid Values:   300,600,1200,2400,4800,9600,19200 or 38400 baud

Default:        9600

ByVal dataConfig&

Control Name:   DataConfig

Description:   The protocoll of the data transmittion can be

configurated.

Variable Type: Integer

Valid Range:   0 to 6

0 = 8 Databits, 1 Stopbit,  No Parity

1 = 8 Databits, 1 Stopbit,  Even Parity

2 = 8 Databits, 1 Stopbit,  Odd Parity

3 = 8 Databits, 2 Stopbits, No Parity

4 = 7 Databits, 1 Stopbit,  Even Parity

5 = 7 Databits, 1 Stopbit,  Odd Parity

6 = 7 Databits, 2 Stopbits, No Parity

Default Value: 0

ByVal reset&

Control Name:  Reset

Description:   If reset is on, a device clear and

power on reset is done during initialization.

Variable Type: Integer

Valid Range:   0 to 1

0 = off

1 = on

instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   COM-Port is out of range

-2   Baud Rate is out of range

-3   Data-Configuration is out of range

-4   Delay-Time is out of range

-5   Reset Device is out of range

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_ReadOperStatusCondReg

Function bp4462_ReadOperStatusCondReg& (ByVal instrumentID&, registerValue&, bit0&, bit8&)

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

registerValue&

bit0&

bit8&

________________________________________________________________________________

bp4462_ReadOperStatusEventReg

Function bp4462_ReadOperStatusEventReg& (ByVal instrumentID&, registerValue&, bit0&, bit8&)

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

registerValue&

bit0&

bit8&

________________________________________________________________________________

bp4462_ReadQuestStatusEventReg

Function bp4462_ReadQuestStatusEventReg& (ByVal instrumentID&, registerValue&, bit0&, bit1&, bit3&, bit4&, bit8&, bit9&)

Purpose

Panel Name:   Read Questionable Status Event Register

Description:  Reads the individual error bits of

the Questionable Status Event Register.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

registerValue&

Control Name:  Register Value

Description:   Return the 16 bit integer value of

the register.

Variable Type: 16-bit Integer

bit0&

Control Name:  Bit 0

Description:   Return the value of Bit 0 from

the register.

'Voltage Error'

Variable Type: Integer

bit1&

Control Name:  Bit 1

Description:   Return the value of Bit 1 from

the register.

'Current Error'

Variable Type: Integer

bit3&

Control Name:  Bit 3

Description:   Return the value of Bit 3 from

the register.

'Power Error'

Variable Type: Integer

bit4&

Control Name:  Bit 4

Description:   Return the value of Bit 4 from

the register.

'Temperature Error'

Variable Type: Integer

bit8&

Control Name:  Bit 8

Description:   Return the value of Bit 8 from

the register.

'Calibration Error'

Variable Type: Integer

bit9&

Control Name:  Bit 9

Description:   Return the value of Bit 9 from

the register.

'PT 100 Error'

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_ReadQuestStatusOutputReg

Function bp4462_ReadQuestStatusOutputReg& (ByVal instrumentID&, registerValue&, bit0&, bit3&, bit4&)

Purpose

Panel Name:   Read Questionable Status Output Register

Description:  Reads the individual error bits of

the Questionable Status Output Register.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

registerValue&

Control Name:  Register Value

Description:   Return the 8 bit hex value of

the register.

Variable Type: 8-bit Hexadecimal

bit0&

Control Name:  Bit 0

Description:   Return the value of Bit 0 from

the register.

'Output Stage Overrange'

Variable Type: Integer

bit3&

Control Name:  Bit 3

Description:   Return the value of Bit 3 from

the register.

'PT 100 Output Error'

Variable Type: Integer

bit4&

Control Name:  Bit 4

Description:   Return the value of Bit 4 from

the register.

'D/A Device Error'

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_ReadQuestStatusTempReg

Function bp4462_ReadQuestStatusTempReg& (ByVal instrumentID&, registerValue&, bit0&, bit1&, bit2&, bit3&)

Purpose

Panel Name:   Read Questionable Status Temperature Register

Description:  Reads the individual error bits of

the Questionable Status Temperature Register.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

registerValue&

Control Name:  Register Value

Description:   Return the 8 bit hex value of

the register.

Variable Type: 8-bit Hexadecimal

bit0&

Control Name:  Bit 0

Description:   Return the value of Bit 0 from

the register.

'Current Wire Error'

Variable Type: Integer

bit1&

Control Name:  Bit 1

Description:   Return the value of Bit 1 from

the register.

'Voltage Wire Error'

Variable Type: Integer

bit2&

Control Name:  Bit 2

Description:   Return the value of Bit 2 from

the register.

'PT 100 Overrange'

Variable Type: Integer

bit3&

Control Name:  Bit 3

Description:   Return the value of Bit 3 from

the register.

'Invalid Temperature'

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_ReadStdEventStatusReg

Function bp4462_ReadStdEventStatusReg& (ByVal instrumentID&, registerValue&, bit0&, bit2&, bit3&, bit4&, bit5&)

Purpose

Panel Name:   Read Standard Event Status Register

Description:  Reads the individual error bits of

the Standard Event Status Register.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

registerValue&

Control Name:  Register Value

Description:   Return the 8 bit integer value of

the register.

Variable Type: 8-bit Integer

bit0&

Control Name:  Bit 0

Description:   Return the value of Bit 0 from

the register.

'Operation Complete'

Variable Type: Integer

bit2&

Control Name:  Bit 2

Description:   Return the value of Bit 2 from

the register.

'Query Error'

Variable Type: Integer

bit3&

Control Name:  Bit 3

Description:   Return the value of Bit 3 from

the register.

'Device Dependent Error'

Variable Type: Integer

bit4&

Control Name:  Bit 4

Description:   Return the value of Bit 4 from

the register.

'Execution Error'

Variable Type: Integer

bit5&

Control Name:  Bit 5

Description:   Return the value of Bit 5 from

the register.

'Command Error'

Variable Type: Integer

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_receive_scpi

Function bp4462_receive_scpi& (ByVal instrumentID&, ByVal sendSCPIMessage$, ??answer_fromSlave??)

Purpose

Panel Name:   Receive a SCPI-Message

Description:  Sends "Query form" from SCPI-

Message and receive a answers

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal sendSCPIMessage$

Control Name:  Send SCPI-Message

Description:   This control is a string parameter

You must input Query-Form from

the SCPI-message without the

LineFeed termination.

Example      : "Sour:volt?"

Variable Type: String

??answer_fromSlave??

Control Name:  Answer from Slave

Description:   This control return a string from

the slave.

Variable Type: String (Lenght => see the 4462

Manual SCPI-Commands)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_ResetQuestOperStEnReg

Function bp4462_ResetQuestOperStEnReg& (ByVal instrumentID&)

Purpose

Panel Name:   Reset Questionable/Operation Status Enable

Register

Description:  Resets the Questionable/Operation Status

Enable Register.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_send_scpi

Function bp4462_send_scpi& (ByVal instrumentID&, ByVal sendSCPIMessage$)

Purpose

Panel Name:   Send a SCPI-Message

Description:  Sends a SCPI-Message to a

slave

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal sendSCPIMessage$

Control Name:  Send SCPI-Message

Description:   This control is a string parameter

You must input the SCPI-message

without the LineFeed termination.

Example:       "SOUR:VOLT 10V"

Variable Type: String

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetAddDeltaValue

Function bp4462_SetAddDeltaValue& (ByVal instrumentID&, ByVal stepMode&)

Purpose

Panel Name:    Set Delta Value Up/Down

Description:   Sets the direction for the delta value steps.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal stepMode&

Control Name:   Step Mode

Description:    Specifies the step mode

of the delta value.

Variable Type:  Integer

Valid Range:    0 to 1

0 = Step Down

1 = Step Up

Default Value:  0 = Step Down

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Step Mode

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigCoefficient

Function bp4462_SetConfigCoefficient& (ByVal instrumentID&, ByVal r0#, ByVal coefficientA#, ByVal coefficientB#, ByVal coefficientC#, ByVal t100#)

Purpose

Panel Name:    Set Config Coefficient PT100

Description:   Sets the values of the PT100 coefficients.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal r0#

Control Name:     R0

Description:      Specifies the R0.

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

Default:          0.0

ByVal coefficientA#

Control Name:     Coefficient "A"

Description:      Specifies the Coefficient "A".

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

Default:          0.0

ByVal coefficientB#

Control Name:     Coefficient "B"

Description:      Specifies the Coefficient "B".

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

Default:          0.0

ByVal coefficientC#

Control Name:     Coefficient "C"

Description:      Specifies the Coefficient "A".

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

Default:          0.0

ByVal t100#

Control Name:     t100

Description:      Specifies the t100.

Formula:

at -200C to 0C: Rt = R0 [1 + At + Bt^2 + C [t - t100] t^3]

at 0C to 850C:  Rt = R0 [1 + At + Bt^2 ]

Variable Type:    Double

Default:          0.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigDeltaValue

Function bp4462_SetConfigDeltaValue& (ByVal instrumentID&, ByVal unit&, ByVal value#, ByVal deviceVersion&)

Purpose

Panel Name:    Set Config Delta Value

Description:   Sets the delta value for the selected unit.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal unit&

Control Name:   Unit

Description:    Specifies the unit of the delta value.

Variable Type:  Integer

Valid Range:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

Default Value:  0 = Voltage (V)

ByVal value#

Control Name:   Value

Description:    Specifies the delta value

of the selected unit.

Variable Type:  Double

Valid Range:

Voltage (V):     0.0 to 60 000 000.0 V

Voltage (mV):     0.0 to 60 000.0 mV

Voltage (V):      0.0 to 60.0 V

Current (A):     0.0 to 104 000.0 A

Current (mA):     0.0 to 104.0 mA

Temperature (C):

Temperature (F):

Temperature ( K): 0.0 to 3 000.0 K

Default Value:  0.0

ByVal deviceVersion&

Control Name:   Device Version

Description:    Switch between the 30 Volts and 60 Volts

Device Version of Digistant 4462.

The 60 Volts Version is comming since

Device Software Version 200201.

Variable Type:  Integer

Valid Values:   0 = 30 Volts Version

1 = 60 Volts Version

Default Value:  0 = 30 Volts Version

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Unit

-3   Invalid Voltage Value

-4   Invalid Current Value

-5   Invalid Temperature Value

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigDeviceAccess

Function bp4462_SetConfigDeviceAccess& (ByVal instrumentID&, ByVal speechConfig&, ByVal outputFunction&, ByVal rampFunction&, ByVal ramp1&, ByVal ramp2&, ByVal deltaValue&, ByVal storeConfig&, ByVal recallConfig&,

ByVal senseConfig&, ByVal outputLimit&, ByVal rangeMode&, ByVal divideConfig&, ByVal TCConfig&, ByVal RJConfig&, ByVal PT100Scale&, ByVal deviceStore&, ByVal deviceRecall&,

ByVal interface&, ByVal contrast&, ByVal statusDisplay&)

Purpose

Panel Name:    Set Config Device Access

Description:   Sets the access configuration for the device.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal speechConfig&

Control Name:  Speech Config

Description:   This parameter activate or

deactivate the access to the

speech configuration menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal outputFunction&

Control Name:  Output Function

Description:   This parameter activate or

deactivate the access to the

output function.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal rampFunction&

Control Name:  Ramp Function

Description:   This parameter activate or

deactivate the access to the

ramp function menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal ramp1&

Control Name:  Ramp1

Description:   This parameter activate or

deactivate the access to the

ramp1 menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal ramp2&

Control Name:  Ramp2

Description:   This parameter activate or

deactivate the access to the

ramp2 menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal deltaValue&

Control Name:  Delta Value

Description:   This parameter activate or

deactivate the access to the

delta value menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal storeConfig&

Control Name:  Store Config

Description:   This parameter activate or

deactivate the access to the

store configuration menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal recallConfig&

Control Name:  Recall Config

Description:   This parameter activate or

deactivate the access to the

recall configuration menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal senseConfig&

Control Name:  Sense Config

Description:   This parameter activate or

deactivate the access to the

sense configuration menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal outputLimit&

Control Name:  Output Limit

Description:   This parameter activate or

deactivate the access to the

output limit menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal rangeMode&

Control Name:  Range Mode

Description:   This parameter activate or

deactivate the access to the

range mode menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal divideConfig&

Control Name:  Divide Config

Description:   This parameter activate or

deactivate the access to the

divide configuration menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal TCConfig&

Control Name:  TC-Config

Description:   This parameter activate or

deactivate the access to the

TC-configuration menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal RJConfig&

Control Name:  RJ-Config

Description:   This parameter activate or

deactivate the access to the

RJ-configuration menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal PT100Scale&

Control Name:  PT100 Scale

Description:   This parameter activate or

deactivate the access to the

PT100 scale menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal deviceStore&

Control Name:  Device Store

Description:   This parameter activate or

deactivate the access to the

device store menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal deviceRecall&

Control Name:  Device Recall

Description:   This parameter activate or

deactivate the access to the

device recall menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal interface&

Control Name:  Interface

Description:   This parameter activate or

deactivate the access to the

interface menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal contrast&

Control Name:  Contrast

Description:   This parameter activates or

deactivates the access to the

contrast menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

ByVal statusDisplay&

Control Name:  Status Display

Description:   This parameter activate or

deactivate the access to the

status display menu.

Variable Type: Integer

Valid Range:   0 = OFF

1 = ON

Default Value: 1 = ON

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0     Success

-1     Invalid Instrument ID

-2     Invalid Speech menu

-3     Invalid Output Mode menu

-4     Invalid Ramp Mode menu

-5     Invalid Ramp1 menu

-6     Invalid Ramp2 menu

-7     Invalid Delta Value menu

-8     Invalid Store Configuration menu

-9     Invalid Recall Configuration menu

-10     Invalid Sense menu

-11     Invalid Output Limit menu

-12     Invalid Range Mode menu

-13     Invalid Divide Configuration menu

-14     Invalid TC-Configuration menu

-15     Invalid RJ-Configuration menu

-16     Invalid PT100 Scale menu

-17     Invalid Device Store menu

-18     Invalid Device Recall menu

-19     Invalid Interface menu

-20     Invalid Contrast menu

-21     Invalid Status Display menu

220    Unable to open instrument

221    Unable to close instrument

224    Error clearing instrument

230    Error writing to instrument

231    Error reading from instrument

232    Instrument not initialized

233    Error configuring GPIB address

236    Error interpreting instrument response

237    Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigDisplayContrast

Function bp4462_SetConfigDisplayContrast& (ByVal instrumentID&, ByVal displayContrast#)

Purpose

Panel Name:    Set Config Display Contrast

Description:   Sets the contrast value for the display.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal displayContrast#

Control Name:   Display Contrast

Description:    Specifies the contrast value

for the Display.

Variable Type:  Double

Valid Values:   0.0 to 1.0

Default Value:  0.5

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Display Contrast

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigDisplayMode

Function bp4462_SetConfigDisplayMode& (ByVal instrumentID&, ByVal mode&)

Purpose

Panel Name:    Set Config Display Mode

Description:   Sets the status of the display.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal mode&

Control Name:   Mode

Description:    Specifies the Status of the

Display Mode of the device.

Variable Type:  Integer

Valid Values:   0 = Delta X  /  Limit Value

1 = TC-Typ   /  RJ-Temperature

2 = TC-Typ   /  U(t)

3 = TC-Typ   /  U(0)

4 = Output Value

5 = PT100 Temperature

Default Value:  0 = Delta X  /  Limit Value

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Display Mode

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigListRamp2

Function bp4462_SetConfigListRamp2& (ByVal instrumentID&, ByVal listNo&, ByVal value#, ByVal stepTimeh&, ByVal stepTimemin&, ByVal stepTimesec#, ByVal unit&, ByVal deviceVersion&)

Purpose

Panel Name:    Set Config List Ramp2

Description:   Sets the ramp function 2 configuration

of the instrument in following

ways :

- Select the output type and unit

- Input the List No.

- Input the List value

- Input the List step time

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal listNo&

Control Name:   List No

Description:    Specifies the List No. of the

ramp2 configuration

Variable Type:  Integer

Valid Range:    1 to 31

Default:        1

ByVal value#

Control Name:   Value

Description:    Specifies the value of the

selected List No. of ramp2

Variable Type:  Double

Valid Range:    by Voltage -30 to 30 V

by Current -52 to 52 mA

by TC    -3000 to 3000 K

Default:        0.0

ByVal stepTimeh&

Control Name:   Step Time[h]

Description:    Specifies the hours of step time

in the selected List No.

Variable Type:  Integer

Valid Range:    0 to 99

Default:        0

ByVal stepTimemin&

Control Name:   Step Time[min]

Description:    Specifies the minutes of the step time

in the selected List No.

Variable Type:  Integer

Valid Range:    0 to 59

Default:        0

ByVal stepTimesec#

Control Name:   Step Time[sec]

Description:    Specifies the seconds of step time

in the selected List No.

Variable Type:  Double

Valid Range:    0.0 to 59.9 seconds

Default:        1.0

ByVal unit&

Control Name:   Unit

Description:    Specifies the unit of the ramp2 value

Variable Type:  Integer

Valid Range:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

Default Value:  0 = Voltage (V)

ByVal deviceVersion&

Control Name:   Device Version

Description:    Switch between the 30 Volts and 60 Volts

Device Version of Digistant 4462.

The 60 Volts Version is comming since

Device Software Version 200201.

Variable Type:  Integer

Valid Values:   0 = 30 Volts Version

1 = 60 Volts Version

Default Value:  0 = 30 Volts Version

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Unit

-3   Invalid List No.

-4   Invalid Step Time hour

-5   Invalid Step Time minute

-6   Invalid Step Time seconds

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigOutputLimitC

Function bp4462_SetConfigOutputLimitC& (ByVal instrumentID&, ByVal currentLimitValue&, ByVal deviceVersion&)

Purpose

Panel Name:    Set Config Output Limit Current

Description:   Sets the current limit value.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal currentLimitValue&

Control Name:   Current Limit Value

Description:    Specifies the current limit value

at the voltage output state.

Variable Type:  Integer

Valid Range:    Current 1 to 55 mA

Default:        55 mA

ByVal deviceVersion&

Control Name:   Device Version

Description:    Switch between the 30 Volts and 60 Volts

Device Version of Digistant 4462.

The 60 Volts Version is comming since

Device Software Version 200201.

Variable Type:  Integer

Valid Values:   0 = 30 Volts Version

1 = 60 Volts Version

Default Value:  0 = 30 Volts Version

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Current Limit Value

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigOutputLimitV

Function bp4462_SetConfigOutputLimitV& (ByVal instrumentID&, ByVal voltageLimitValue&, ByVal deviceVersion&)

Purpose

Panel Name:    Set Config Output Limit Volt

Description:   Sets the voltage limit value.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal voltageLimitValue&

Control Name:   Voltage Limit Value

Description:    Specifies the voltage limit value

at the current output state.

Variable Type:  Integer

Valid Range:    Voltage 1 to 32 V

Default:        32 V

ByVal deviceVersion&

Control Name:   Device Version

Description:    Switch between the 30 Volts and 60 Volts

Device Version of Digistant 4462.

The 60 Volts Version is comming since

Device Software Version 200201.

Variable Type:  Integer

Valid Values:   0 = 30 Volts Version

1 = 60 Volts Version

Default Value:  0 = 30 Volts Version

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Voltage Limit Value

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigOutputMode

Function bp4462_SetConfigOutputMode& (ByVal instrumentID&, ByVal mode&)

Purpose

Panel Name:    Set Config Output Mode

Description:   Switch to the normal mode or the selected

ramp mode of the instrument.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal mode&

Control Name:   Mode

Description:    This parameter switches the

Ramp Mode of the device.

Variable Type:  Integer

Valid Values:   0 = OFF / Normal

1 = Ramp1

2 = Ramp2

Default:        0 = OFF / Normal

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Ramp Mode

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigOutputStatus

Function bp4462_SetConfigOutputStatus& (ByVal instrumentID&, ByVal unit&, ByVal value#, ByVal deviceVersion&)

Purpose

Panel Name:    Set Config Output Status

Description:   Sets the output value and unit.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal unit&

Control Name:   Unit

Description:    Specifies the unit of the output value.

Variable Type:  Integer

Valid Range:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

Default Value:  0 = Voltage (V)

ByVal value#

Control Name:   Value

Description:    Specifies the output value.

Variable Type:  Double

Valid Range:    Voltage -30 to 30 V

Current -52 to 52 mA

TC    -3000 to 3000 K

Default:        0.0

ByVal deviceVersion&

Control Name:   Device Version

Description:    Switch between the 30 Volts and 60 Volts

Device Version of Digistant 4462.

The 60 Volts Version is comming since

Device Software Version 200201.

Variable Type:  Integer

Valid Values:   0 = 30 Volts Version

1 = 60 Volts Version

Default Value:  0 = 30 Volts Version

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Unit

-3   Invalid Voltage Value

-4   Invalid Current Value

-5   Invalid Temperature Value

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigRamp1

Function bp4462_SetConfigRamp1& (ByVal instrumentID&, ByVal waveform&, ByVal numberOfCounts&, ByVal stepTimeh&, ByVal stepTimemin&, ByVal stepTimesec#, ByVal unit&, ByVal stepValue#, ByVal startValue#, ByVal endValue#,

ByVal deviceVersion&)

Purpose

Panel Name:    Set Config Ramp1

Description:   Configure the ramp function 1

of the instrument in following

ways :

- Select the instrument

- Select the waveform

- Input the step time (hour, minute, seconds)

- Select the unit

- Input the step value

- Input the start value

- Input the stop value

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal waveform&

Control Name:   Waveform

Description:    Specifies the waveform  of ramp1

Variable Type:  Integer

Valid Values:   0 = Sawtooth

1 = Triangle

Default Value:  0 = Sawtooth

ByVal numberOfCounts&

Control Name:   Number of Counts

Description:    Specifies the number of counts

of ramp1

Variable Type:  Integer

Valid Range:    0 to 99

0 = Countiue

Default:        1

ByVal stepTimeh&

Control Name:   Step Time (hours)

Description:    Specifies the hours of step time

Variable Type:  Integer

Valid Range:    0 to 99 hours

Default:        0

ByVal stepTimemin&

Control Name:   Step Time (minutes)

Description:    Specifies the minutes of step time

Variable Type:  Integer

Valid Range:    0 to 59 minutes

Default:        0

ByVal stepTimesec#

Control Name:   Step Time (seconds)

Description:    Specifies the seconds of step time

Variable Type:  Double

Valid Range:    0.0 to 59.9 seconds

Default:        1.0

ByVal unit&

Control Name:   Unit

Description:    Specifies the Unit of the Ramp1 Values.

Variable Type:  Integer

Valid Values:   0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

Default Value:  0 = Voltage (V)

ByVal stepValue#

Control Name:   Step Value

Description:    Specifies the step value of

ramp1

Variable Type:  Double

Valid Range:    by Voltage 0 to 60 V

by Current 0 to 104 mA

by TC      0 to 3000 K

Default:        0.0

ByVal startValue#

Control Name:   Start Value

Description:    Specifies the start value of

ramp1

Variable Type:  Double

Valid Range:    by Voltage -30 to 30 V

by Current -52 to 52 mA

by TC      0   to 3000 K

Default:        0.0

ByVal endValue#

Control Name:   End Value

Description:    Specifies the end value of

ramp1

Variable Type:  Double

Valid Range:    by Voltage -30 to 30 V

by Current -52 to 52 mA

by TC      0   to 3000 K

Default:        0.0

ByVal deviceVersion&

Control Name:   Device Version

Description:    Switch between the 30 Volts and 60 Volts

Device Version of Digistant 4462.

The 60 Volts Version is comming since

Device Software Version 200201.

Variable Type:  Integer

Valid Values:   0 = 30 Volts Version

1 = 60 Volts Version

Default Value:  0 = 30 Volts Version

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Waveform

-3   Invalid Unit

-4   Invalid Number of Counts

-5   Invalid Step Value

-6   Invalid Step Time (hour)

-7   Invalid Step Time (minute)

-8   Invalid Step Time (seconds)

-9   Invalid Start Value

-10  Invalid Stop Value

-11  Invalid Step Value

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigRamp2

Function bp4462_SetConfigRamp2& (ByVal instrumentID&, ByVal waveform&, ByVal numberOfCounts&, ByVal numberOfPoints&, ByVal listNo&)

Purpose

Panel Name:    Set Config Ramp2

Description:   Sets the ramp function 2 configuration

of the instrument in following

ways :

- Select the waveform

- Input the number of counts

- Input the number of points

- Input the List No.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal waveform&

Control Name:   Waveform

Description:    Specifies the waveform of ramp2

Variable Type:  Integer

Valid Values:   0 = Sawtooth

1 = Triangle

Default:        0 = Sawtooth

ByVal numberOfCounts&

Control Name:   Number of Counts

Description:    Specifies the number of counts

of ramp2

Variable Type:  Integer

Valid Range:    0 to 99

0 = Countiue

Default:        1

ByVal numberOfPoints&

Control Name:   Number of Points

Description:    Specifies the number of points

of ramp2

Variable Type:  Integer

Valid Range:    2 to 31

Default:        2

ByVal listNo&

Control Name:   List No

Description:    Specifies the List No. of the

input values

Variable Type:  Integer

Valid Range:    1 to 31

Default:        1

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Waveform

-3   Invalid Number of Counts

-4   Invalid Number of Points

-5   Invalid List Number

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigRangeMode

Function bp4462_SetConfigRangeMode& (ByVal instrumentID&, ByVal rangeMode&, ByVal deviceVersion&)

Purpose

Panel Name:    Set Config Range Mode

Description:   Switch between Automatic and Manual Range.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal rangeMode&

Control Name:  Range

Description:   Specifies the range mode and the range value

if range mode is set to manual range.

Variable Type: Integer

Valid Range:   0 = Autorange

1 = 300 mVolt

2 = 3 Volt

3 = 30 Volt

ByVal deviceVersion&

Control Name:   Device Version

Description:    Switch between the 30 Volts and 60 Volts

Device Version of Digistant 4462.

The 60 Volts Version is comming since

Device Software Version 200201.

Variable Type:  Integer

Valid Values:   0 = 30 Volts Version

1 = 60 Volts Version

Default Value:  0 = 30 Volts Version

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Range Mode

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigRjType

Function bp4462_SetConfigRjType& (ByVal instrumentID&, ByVal RJType&)

Purpose

Panel Name:    Set Config RJ Type

Description:   Switch between RJ-External and RJ-Manual.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal RJType&

Control Name:  RJ-Type

Description:   Specifies the mode of reference junction type.

Variable Type: Integer

Valid Range:   0 to 1

0 = Manual Reference Junction Type

1 = External Reference Junction Type

Default Value: 0 = Manual Reference Junction Type

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid RJ-Type

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigSense

Function bp4462_SetConfigSense& (ByVal instrumentID&, ByVal mode&)

Purpose

Panel Name:    Set Config Sense

Description:   Sets the configuration for the sense in

following ways:

- External

- Internal

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal mode&

Control Name:   Mode

Description:    Specifies the Sense Mode of the device.

Variable Type:  Integer

Valid Values:   0 = Sense Internal

1 = Sense External

Default Value:  0 = Sense Internal

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Sense Mode

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigTc

Function bp4462_SetConfigTc& (ByVal instrumentID&, ByVal unit&, ByVal type&, ByVal scale&)

Purpose

Panel Name:    Set Config TC

Description:   Sets the configuration for the TC-mode.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal unit&

Control Name:  Unit

Description:   Specifies the tc-temperature unit.

Variable Type: Integer

Valid Range:   5 to 7

5 = C (Celsius)

6 = F (Fahrenheit)

7 =  K (Kelvin)

Default Value: 5 = C (Celsius)

ByVal type&

Control Name:   Type

Description:    Specifies the TC-Type.

Variable Type:  Integer

Valid Values:   0 = R

1 = S

2 = B

3 = J

4 = T

5 = E

6 = K

7 = U

8 = L

9 = N

10 = C

11 = D

12 = G2

13 = M

Default Value:  0 = R

ByVal scale&

Control Name:  Scale

Description:   Specifies the tc-temperature scale.

Variable Type: Integer

Valid Range:   0 to 1

0 = IPTS 68

1 = ITS 90

Default Value: 1 = ITS 90

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid TC-Type

-3   Invalid TC-Scale

-4   Invalid TC-Unit

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigTcUnit

Function bp4462_SetConfigTcUnit& (ByVal instrumentID&, ByVal unit&)

Purpose

Panel Name:    Set Config TC Unit

Description:   Sets the unit for the TC-mode.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal unit&

Control Name:  Unit

Description:   Specifies the tc-temperature unit.

Variable Type: Integer

Valid Range:   5 to 7

5 = C (Celsius)

6 = F (Fahrenheit)

7 =  K (Kelvin)

Default Value: 5 = C (Celsius)

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid TC-Unit

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigVoltageDiv

Function bp4462_SetConfigVoltageDiv& (ByVal instrumentID&, ByVal dividerFactor#)

Purpose

Panel Name:    Set Config Voltage Divider

Description:   Sets the factor value for the voltage divider.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal dividerFactor#

Control Name:   Divider Factor

Description:    Specifies the factor value

of the voltage dividor.

Variable Type:  Double

Valid Range:    1.0 to 1010.0

Default Value:  1.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Divider Factor Value

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetConfigVoltageDivState

Function bp4462_SetConfigVoltageDivState& (ByVal instrumentID&, ByVal dividerMode&)

Purpose

Panel Name:    Set Config Voltage Divider State

Description:   Sets the mode of the voltage divider.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal dividerMode&

Control Name:   Divider Mode

Description:    Specifies the divider mode of the device.

Variable Type:  Integer

Valid Range:    0 to 1

0 = Divider OFF

1 = Divider ON

Default Value:  0 = Divider OFF

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Divider Mode

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetLockSystemKeyboard

Function bp4462_SetLockSystemKeyboard& (ByVal instrumentID&, ByVal keyboard&)

Purpose

Panel Name:    Set Lock System Keyboard

Description:   Sets the System Keyboard of the device

to the selected state.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal keyboard&

Control Name:  Keyboard

Description:   This parameter locks the keyboard

on the front panel of the

instrument.

Variable Type: Integer

Valid Range:   0  = Unlocked

1  = Locked

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Value

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetMemoryLabel

Function bp4462_SetMemoryLabel& (ByVal instrumentID&, ByVal memoryNo&, ByVal memoryLabel$)

Purpose

Panel Name:   Set Memory Label

Description:  Sets the label text of a

selected device memory ID.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal memoryNo&

Control Name:  Memory No.

Description:   Specificies the memory No. to

store the label text

Variable Type: Integer

Valid Range:   1 -  33

ByVal memoryLabel$

Control Name:  Memory Label

Description:   Specifices the memory label text.

Variable Type: String

Valid Range:   max. 10 Byte

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid State

-3   Memory label to long

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetRampStart

Function bp4462_SetRampStart& (ByVal instrumentID&, ByVal startStop&)

Purpose

Panel Name:    Set Ramp Start/Stop

Description:   Switch between ramp Start or

Stop

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal startStop&

Control Name:   Start/Stop

Description:    Specifies the status of the ramp mode.

Variable Type:  Integer

Valid Values:   0 = Stop, 1 = Start

Default:        0 = Stop

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Start/Stop

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetRampStep

Function bp4462_SetRampStep& (ByVal instrumentID&, ByVal stepDirection&)

Purpose

Panel Name:    Set Ramp Step

Description:   Sets the direction for the ramp steps

in following ways:

- Ramp steps up

- Ramp steps down

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal stepDirection&

Control Name:   Step Direction

Description:    Specifies the step direction

of the ramp value.

Variable Type:  Integer

Valid Values:   0 = Down, 1 = Up

Default:        0 = Down

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Step Direction

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetRjTempValueMan

Function bp4462_SetRjTempValueMan& (ByVal instrumentID&, ByVal unit&, ByVal value#)

Purpose

Panel Name:    Set RJ Temperature Value Manual

Description:   Sets the temperature value and unit for the

manual reference junction.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal unit&

Control Name:  Unit

Description:   Specifies the temperature unit of the

rj-temperature value.

Variable Type: Integer

Valid Range:   5 to 7

5 = C (Celsius)

6 = F (Farenheit)

7 =  K (Kelvin)

Default Value: 5 = C (Celsius)

ByVal value#

Control Name:   Value

Description:    Specifies the rj-temperature value.

Variable Type:  Double

Valid Range:    -50.00 to 150.00 C

-58.00 to 302.00 F

223.15 to 423.15  K

Default Value:  0.0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrument ID

-2   Invalid Range Mode

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetStoreIdOutputValue

Function bp4462_SetStoreIdOutputValue& (ByVal instrumentID&, ByVal memoryNo&, ByVal outputUnit&, ByVal outputValue#, ByVal deltaValue#, ByVal deltaUnit&, ByVal deviceVersion&)

Purpose

Panel Name:    Set Store ID Output Value

Description:   Stores the specified output and delta value

under the selected memory No. with value unit.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal memoryNo&

Control Name:   Memory No.

Description:    Specifies the Memory No. of the stored

output and delta value.

Variable Type:  integer

Valid Range:    0 to 99

Default:        0

ByVal outputUnit&

Control Name:   Output Unit

Description:    Specifies the unit of the output value.

Variable Type:  Integer

Valid Range:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

Default Value:  0 = Voltage (V)

ByVal outputValue#

Control Name:   Output Value

Description:    Specifies the output value of the

selected memory No.

Variable Type:  Double

Valid Range:    by Voltage -30 to 30 V

by Current -52 to 52 mA

by TC        0 to 3000 K

Default:        0.0

ByVal deltaValue#

Control Name:   Delta Value

Description:    Specifies the delta value of the

selected memory No.

Variable Type:  Double

Valid Range:    by Voltage -30 to 30 V

by Current -52 to 52 mA

by TC        0 to 3000 K

Default:        0.0

ByVal deltaUnit&

Control Name:   Delta Unit

Description:    Specifies the unit of the delta value.

Variable Type:  Integer

Valid Range:    0 to 7

0 = Voltage (V)

1 = Voltage (mV)

2 = Voltage (V)

3 = Current (A)

4 = Current (mA)

5 = Temperature (C)

6 = Temperature (F)

7 = Temperature ( K)

Default Value:  0 = Voltage (V)

ByVal deviceVersion&

Control Name:   Device Version

Description:    Switch between the 30 Volts and 60 Volts

Device Version of Digistant 4462.

The 60 Volts Version is comming since

Device Software Version 200201.

Variable Type:  Integer

Valid Values:   0 = 30 Volts Version

1 = 60 Volts Version

Default Value:  0 = 30 Volts Version

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Memory No.

-3   Error Output Unit return

-4   Error Delta Unit return

-5   Invalid Voltage Output Value

-6   Invalid Current Output Value

-7   Invalid Temperature Output Value

-8   Invalid Voltage Delta Value

-9   Invalid Current Delta Value

-10  Invalid Temperature Delta Value

-11  Unequal Output and Delta Types

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

________________________________________________________________________________

bp4462_SetStoreOutputValue

Function bp4462_SetStoreOutputValue& (ByVal instrumentID&, ByVal memoryNo&)

Purpose

Panel Name:    Set Store Output Value

Description:   Stores the actually output and delta value

under a specified memory No. in the device.

Parameter List

ByVal instrumentID&

Control Name:  Instrument ID

Description:   The initialize function call

returns an instrument ID, which is used as an

input here.  (The ID is used to select among

like-instruments).

Variable Type: Integer

Valid Range:   1 to 10 (typical)

The value 10 is set internally in the driver.

ByVal memoryNo&

Control Name:   Memory No.

Description:    Specifies the Memory No. of the stored

output and delta value.

Variable Type:  integer

Valid Range:    0 to 99

Default:        0

Return Value

Control Name:  Error

Description:   Displays the results of the

function call.

Variable Type: Integer

Errors:

0   Success

-1   Invalid Instrumet ID

-2   Invalid Memory No.

1 - 124   RS-232 error (see RS-232 Library

error codes)

220   Unable to open instrument

221   Unable to close instrument

224   Error clearing instrument

230   Error writing to instrument

231   Error reading from instrument

232   Instrument not initialized

233   Error configuring GPIB address

236   Error interpreting instrument response

237   Error configuring EOS terminator

300   The instrument sends NAK

301   The instrument sends EOT

320   Datalogger is empty