## calculator
Agilent E932xA Power Sensor Uncertainty Calculator
Input | Value
Device Under Test (SWR) | 1.1
Sensor Model | E9321A
Power Meter Model | EPM-P
Frequency | 3 | GHz
Power | 0 | dBm |  | 0.001 | W
Number of Average Reading | 8
Measurement mode (Avg only / Normal) | Average Only
Speed (normal, x2, fast) | Normal
Below settings are not required if Measurement Mode = Avg Only:
Acquisition type (Trigger/ Free Run) | Trigger
VBW setting (Low, Med, High, Off) | Off
Gate length (not required for Free Run Acquisition) | 0.0005 | s
Source of Uncertainty | Symbol | Value ±% | Probability Distribution | Divisor | Result
Mismatch Gain Between Generator and Sensor | Mu | ¦Гg¦ = 0.048 | U-shape | 1.4142 | 0.4988%
¦Гs¦ = 0.074
Mismatch Gain Between Calibration Source and Sensor | Muc | ¦Гc¦ = 0.029 | U-shape | 1.4142 | 0.2332%
¦Гsc¦ = 0.057
Power Meter Instrumentation Error | Pm | 0.5% | Gaussian | 2 | 0.25%
Power Meter Instrumentation Error During Calibration | Pmc | 0.5% | Gaussian | 2 | 0.25%
Power Meter Calibrator Output Power | Pcal | 0.5% | Gaussian | 2 | 0.25%
Zero drift | D | 6e-11 | Gaussian | 2 | 3e-06%
Power Sensor Calibration Factor Uncertainties | Kb | 2.1% | Gaussian | 2 | 1.05%
Power Sensor Linearity | Pl | 3.7% | Gaussian | 2 | 1.85%
Zero set | Zs | 1.7e-10 | Gaussian | 2 | 8.5e-06%
Sensor noise | N | 3.201e-10 | Gaussian | 2 | 1.601e-05%
|  |  | Combined Uncertainty-RSSed | = | 2.24%
|  |  | K | = | 2
|  |  | Expanded Uncertainty | = | 4.479%
|  |  | Upper Limit Uncertainty | = | + 0.190dB
|  |  | Lower Limit Uncertainty | = | − 0.199dB
Revision 3
Date updated: 25 Nov 2010
© Agilent Technologies, Inc. 2008
The data used in this calculation is base on the specifications as in the EPM-P and P-Series datasheet, literature number 5980-1469E and 5989-2471EN.
This calculation is base on ISO Guide to the Expression of Uncertainty in Measurement, often referred to as the GUM.
For more info about GUM, please refer to Agilent Application Note 1449-3 "Fundamental of RF and Microwave Power Measurements (Part 3), literature number 5988-9215EN.

## Data (hidden)
Power Meter instrumentation Errors & Calibrator Accuracy
Max SWR | [7 empty] | PowerMeter | Power Meter Instrumentation Error - Average Only | Power Meter Instrumentation Error - Normal | Power Meter Instrumentation Error During Calibration - Average Only | Power Meter Instrumentation Error During Calibration - Normal | Power Meter Calibrator Output Power
[8 empty] | EPM-P | 0.5% | 0.8% | 0.5% | 0.8% | 0.5%
Frequency | E9321A | E9322A | E9323A | E9325A | E9326A | E9327A |  | P-Series | 0.5% | 0.8% | 0.5% | 0.8% | 0.4%
0.05 | 1.12 | 1.12 | 1.14 | 1.12 | 1.12 | 1.14
0.1 | 1.12 | 1.12 | 1.14 | 1.12 | 1.12 | 1.14 |  | Power Meter Instrumentation Error - Average Only
2 | 1.16 | 1.18 | 1.22 | 1.16 | 1.18 | 1.22 |  | Power Meter Instrumentation Error During Calibration - Average Only
4 | 1.16 | 1.18 | 1.22 | 1.16 | 1.18 | 1.22
4.2 | 1.16 | 1.18 | 1.22 | 1.16 | 1.18 | 1.22
6 | 1.16 | 1.18 | 1.22 | 1.16 | 1.18 | 1.22
8 | # | # | # | 1.16 | 1.18 | 1.22
10 | # | # | # | 1.23 | 1.18 | 1.22
11 | # | # | # | 1.23 | 1.18 | 1.22
12 | # | # | # | 1.23 | 1.21 | 1.22
12.4 | # | # | # | 1.23 | 1.21 | 1.22
14 | # | # | # | 1.23 | 1.21 | 1.22
15 | # | # | # | 1.23 | 1.21 | 1.22
16 | # | # | # | 1.28 | 1.27 | 1.26
18 | # | # | # | 1.28 | 1.27 | 1.26
Power Linearity (25+/-10C)
Power | E9321A: Average Only; E9322A: Normal
-65 | E9321A: 3.7%; E9322A: 4.2%
-60 | E9321A: 3.7%; E9322A: 4.2%
-50 | E9321A: 3.7%; E9322A: 4.2%
-45 | E9321A: 3.7%; E9322A: 4.2%
-40 | E9321A: 3.7%; E9322A: 4.2%
-35 | E9321A: 3.7%; E9322A: 4.2%
-30 | E9321A: 3.7%; E9322A: 4.2%
-18.5 | E9321A: 3.7%; E9322A: 4.2%
-17.5 | E9321A: 3.7%; E9322A: 4.2%
-15 | E9321A: 3.7%; E9322A: 4.2%
-14.5 | E9321A: 3.7%; E9322A: 4.2%
-13.5 | E9321A: 3.7%; E9322A: 4.2%
-11.5 | E9321A: 3.7%; E9322A: 4.2%
-10.5 | E9321A: 3.7%; E9322A: 4.2%
-9.5 | E9321A: 3.7%; E9322A: 4.2%
-5 | E9321A: 3.7%; E9322A: 4.2%
0.5 | E9321A: 3.7%; E9322A: 4.2%
20 | E9321A: 3.7%; E9322A: 4.2%
Zero Set
Power Range | E9321A: Model; E9322A: Normal; E9323A: Average only
-65 to +20dBm | E9321A: E9321A; E9322A: 5e-09; E9323A: 1.7e-10
-60 to +20dBm | E9321A: E9322A; E9322A: 1.9e-08; E9323A: 5e-10
-60 to +20dBm | E9321A: E9323A; E9322A: 6e-08; E9323A: 6e-10
-65 to +20dBm | E9321A: E9325A; E9322A: 5e-09; E9323A: 1.7e-10
-60 to +20dBm | E9321A: E9326A; E9322A: 1.9e-08; E9323A: 5e-10
-60 to +20dBm | E9321A: E9327A; E9322A: 6e-08; E9323A: 6e-10
Zero Drift
Model | E9321A: Normal; E9322A: Average Only
E9321A | E9321A: 5e-09; E9322A: 6e-11
E9322A | E9321A: 5e-09; E9322A: 1e-10
E9323A | E9321A: 4e-08; E9322A: 1e-10
E9325A | E9321A: 5e-09; E9322A: 6e-11
E9326A | E9321A: 5e-09; E9322A: 1e-10
E9327A | E9321A: 4e-08; E9322A: 1e-10
Sensor Noise (noise per sample) | E9321A: Normal; E9326A: Normal | Average Only
Model | Trigger |  |  | Sensor Noise | Trigger | Free Run
E9321A | 7.5e-08 |  |  | Acq | Error | Error | 1.65e-10
E9322A | E9321A: 1.8e-07; E9325A: Multiplier; E9327A: 0.82 | 1.94
E9323A | E9321A: 5.5e-07; E9325A: Video Bandwidth; E9326A: 1
E9325A | E9321A: 7.5e-08; E9325A: Gate Length; E9326A: 0.01
E9326A | 1.8e-07 |  |  | Total Noise | #VALUE! | #VALUE! | 3.201e-10
E9327A | E9321A: 5.5e-07; E9325A: Overall Sensor Noise; E9326A: 3.201e-10
Sensor Noise (Normal, free run mode)
Model | E9321A: Free Run
E9321A | E9321A: 6e-09
E9322A | E9321A: 1.2e-08
E9323A | E9321A: 2.5e-08
E9325A | E9321A: 6e-09
E9326A | E9321A: 1.2e-08
E9327A | E9321A: 2.5e-08
Sensor Noise (Average only mode)
Model | E9321A: Average Only
E9321A | E9321A: 1.65e-10
E9322A | E9321A: 3.3e-10
E9323A | E9321A: 4e-10
E9325A | E9321A: 1.65e-10
E9326A | E9321A: 3.3e-10
E9327A | E9321A: 4e-10
Measurement Type
Average Only | E9321A: Normal
Speed
Normal | E9321A: x2; E9322A: Fast
Acquisition type (triggered/ free run)
Free Run | E9321A: Trigger
Video Bandwidth
Low | E9321A: Medium; E9322A: High; E9323A: Off
Noise multiplier
Average Only mode (EPM-P)
Number of avg | Normal | x2 |  | Multiplier | Average Only | Normal
1 | 5.5 | 6.5 |  | EPM-P | 1.94 | 0.82
2 | E9321A: 3.89; E9322A: 4.6
4 | E9321A: 2.75; E9322A: 3.25
8 | E9321A: 1.94; E9322A: 2.3
16 | E9321A: 1; E9322A: 1.63
32 | E9321A: 0.85; E9322A: 1
64 | E9321A: 0.61; E9322A: 0.72
128 | E9321A: 0.49; E9322A: 0.57
256 | E9321A: 0.34; E9322A: 0.41
512 | E9321A: 0.24; E9322A: 0.29
1024 | E9321A: 0.17; E9322A: 0.2
Normal mode, Free Run(EPM-P)
Number of avg | E9321A: Normal; E9322A: x2; E9323A: Fast
1 | E9321A: 1; E9322A: 1.2; E9323A: 3.4
2 | E9321A: 0.94; E9322A: 1.128; E9323A: 3.4
4 | E9321A: 0.88; E9322A: 1.056; E9323A: 3.4
8 | E9321A: 0.82; E9322A: 0.984; E9323A: 3.4
16 | E9321A: 0.76; E9322A: 0.912; E9323A: 3.4
32 | E9321A: 0.7; E9322A: 0.84; E9323A: 3.4
64 | E9321A: 0.64; E9322A: 0.768; E9323A: 3.4
128 | E9321A: 0.58; E9322A: 0.696; E9323A: 3.4
256 | E9321A: 0.52; E9322A: 0.624; E9323A: 3.4
512 | E9321A: 0.46; E9322A: 0.552; E9323A: 3.4
1024 | E9321A: 0.4; E9322A: 0.48; E9323A: 3.4
Cal factor uncertainties (-65 to +20dBm, 25+/-10degC)
Greater Than 0dBm | 0.1% | 0.15% | 0.2% | 0.1% | 0.15% | 0.2%
Frequency | E9321A | E9322A | E9323A | E9325A | E9326A | E9327A | Reference
0.05 | 0.1% | 0.15% | 0.2% | 0.1% | 0.15% | 0.2% | 0%
0.1 | 1.9% | 1.95% | 2% | 1.9% | 1.95% | 2% | 1.8%
0.3 | 1.9% | 1.95% | 2% | 1.9% | 1.95% | 2% | 1.8%
0.5 | 1.9% | 1.95% | 2% | 1.9% | 1.95% | 2% | 1.8%
0.8 | 1.9% | 1.95% | 2% | 1.9% | 1.95% | 2% | 1.8%
1 | 2.2% | 2.25% | 2.3% | 2.2% | 2.25% | 2.3% | 2.1%
1.2 | 2.2% | 2.25% | 2.3% | 2.2% | 2.25% | 2.3% | 2.1%
1.5 | 2.2% | 2.25% | 2.3% | 2.2% | 2.25% | 2.3% | 2.1%
2 | 2.2% | 2.25% | 2.3% | 2.2% | 2.25% | 2.3% | 2.1%
3 | 2.2% | 2.25% | 2.3% | 2.2% | 2.25% | 2.3% | 2.1%
4 | 2.2% | 2.25% | 2.3% | 2.2% | 2.25% | 2.3% | 2.1%
5 | 2.2% | 2.25% | 2.3% | 2.2% | 2.25% | 2.3% | 2.1%
6 | 2.2% | 2.25% | 2.3% | 2.2% | 2.25% | 2.3% | 2.1%
7 | NA | NA | NA | 2.4% | 2.45% | 2.5% | 2.3%
8 | NA | NA | NA | 2.4% | 2.45% | 2.5% | 2.3%
9 | NA | NA | NA | 2.4% | 2.45% | 2.5% | 2.3%
10 | NA | NA | NA | 2.4% | 2.45% | 2.5% | 2.3%
11 | NA | NA | NA | 2.4% | 2.45% | 2.5% | 2.3%
12 | NA | NA | NA | 2.4% | 2.45% | 2.5% | 2.3%
12.4 | NA | NA | NA | 2.4% | 2.45% | 2.5% | 2.3%
13 | NA | NA | NA | 2.4% | 2.45% | 2.5% | 2.3%
14 | NA | NA | NA | 2.6% | 2.65% | 2.7% | 2.5%
15 | NA | NA | NA | 2.6% | 2.65% | 2.7% | 2.5%
16 | NA | NA | NA | 2.6% | 2.65% | 2.7% | 2.5%
17 | NA | NA | NA | 2.6% | 2.65% | 2.7% | 2.5%
18 | NA | NA | NA | 2.6% | 2.65% | 2.7% | 2.5%
Less than & equal to 0dBm
Frequency | E9321A | E9322A | E9323A | E9325A | E9326A | E9327A
0.05 | [8 empty] | > 0dBm | <= 0dBm | bar: E9321A to E9327A
0.1 | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% |  | Uncertainty | 2.2% | 2.1%
0.3 | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8%
0.5 | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8%
0.8 | 1.8% | 1.8% | 1.8% | 1.8% | 1.8% | 1.8%
1 | 2.1% | 2.1% | 2.1% | 2.1% | 2.1% | 2.1%
1.2 | 2.1% | 2.1% | 2.1% | 2.1% | 2.1% | 2.1%
1.5 | 2.1% | 2.1% | 2.1% | 2.1% | 2.1% | 2.1%
2 | 2.1% | 2.1% | 2.1% | 2.1% | 2.1% | 2.1%
3 | 2.1% | 2.1% | 2.1% | 2.1% | 2.1% | 2.1%
4 | 2.1% | 2.1% | 2.1% | 2.1% | 2.1% | 2.1%
5 | 2.1% | 2.1% | 2.1% | 2.1% | 2.1% | 2.1%
6 | 2.1% | 2.1% | 2.1% | 2.1% | 2.1% | 2.1%
7 | NA | NA | NA | 2.3% | 2.3% | 2.3%
8 | NA | NA | NA | 2.3% | 2.3% | 2.3%
9 | NA | NA | NA | 2.3% | 2.3% | 2.3%
10 | NA | NA | NA | 2.3% | 2.3% | 2.3%
11 | NA | NA | NA | 2.3% | 2.3% | 2.3%
12 | NA | NA | NA | 2.3% | 2.3% | 2.3%
12.4 | NA | NA | NA | 2.3% | 2.3% | 2.3%
13 | NA | NA | NA | 2.3% | 2.3% | 2.3%
14 | NA | NA | NA | 2.5% | 2.5% | 2.5%
15 | NA | NA | NA | 2.5% | 2.5% | 2.5%
16 | NA | NA | NA | 2.5% | 2.5% | 2.5%
17 | NA | NA | NA | 2.5% | 2.5% | 2.5%
18 | NA | NA | NA | 2.5% | 2.5% | 2.5%
Video Bandwidth (noise multiplier)
| E9321A | E9322A | E9323A | E9325A | E9326A | E9327A
Low | 0.32 | 0.5 | 0.4 | 0.32 | 0.5 | 0.4
Medium | 0.5 | 0.63 | 0.63 | 0.5 | 0.63 | 0.63
High | 0.63 | 0.8 | 1 | 0.63 | 0.8 | 1
Off | bar: E9321A to E9327A
rho sensor @ calilbration | E9321A: 0.056604; E9322A: 1.12
rho DUT | E9321A: 0.047619
rho sensor | E9321A: 0.074074; E9322A: 1.16
rho calibrator | E9321A: 0.029126
Calibrator Mismatch | E9321A: EPM-P; E9322A: P-Series
SWR calibrator | E9321A: 1.06; E9322A: 1.08
rho calibrator | E9321A: 0.029126; E9322A: 0.038462
Rectangular | E9321A: 1.7321
Gaussian | E9321A: 2
U-shape | E9321A: 1.4142
Frequency Check
E9321A
E9322A
E9323A
E9325A
E9326A
E9327A
Power check
E9321A
E9322A
E9323A
E9325A
E9326A
E9327A
Note on N5: 25+/-10degC
Note on J5: 0.5% for CW measurements. 0.8% for peak measurements
Note on N4: 23+/-3degC
Note on A201: 0 to 55degC
Note on K5: 0.5% for CW measurements. 0.8% for peak measurements
Note on E67: For Triggered mode: Noise = 1/√(gate length/ X ) * Noise per sample * Noise per sample multiplier base on VBW (Table 11) X = 10ns for Fluffy sensors X = 50ns for Frodo sensors

## Revision
Revision 1 (23 Dec 2008)
New creation
Revision 2 (28 May 2009)
Corrected error on mismatch uncertainty calculation.
Revision 3 | (25 Nov 2010)
Corrected error on mismatch uncertainty calculation.