EPM: E4418/19B Power Meter Firmware Revision Information

Firmware Version A1.03.00 and A2.03.00

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| Zero/Cal Lockout | Feature can be activated to prevent measurements being made until connected sensor is zeroed and calibrated. | User Guide: Chapter 2 “Zeroing and Calibrating the Power Meter; Zero / Cal Lockout |
| Sensor Information over GPIB | Ability to read sensor model, serial number and last calibration date (for E-series sensors only) over GPIB. | Programming Guide: Chapter 14 “Service subsystem” |
| V8486A sensor linearity correction | Applies linearity correction for V8486A V-band sensor | User Guide: Chapter 2 “Zeroing and Calibrating the Power Meter; Calibration Procedure using Agilent 8480 series sensors”. |
| Filter length setting | The digital filter can be set to any number between 1 and 1024, to set the optimal averaging for a particular measurement. | User Guide: Chapter 2 “Setting Averaging” |
| Step detect on/off | Step detection (which re-initializes the filter upon detection of a step increase/decrease in power) can be turned on or off. Step detect may be turned off to prevent pulsed signals re-initializing the filter. | User Guide: Chapter 2 “Setting Averaging; Step Detection” |

Firmware Versions A1.04.00 and A2.04.00

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| E-series E9300 power sensor compatibility | Enables EPM series power meters to use the E-series E9300 power sensors. EPM “B” (E4418B/19B) power meters with serial prefix US3847 or GB3841 will require firmware revision A1.04.00/A2.04.00 for Agilent E9300 sensor operation. Please note that EPM “A” power meters (E4418A/19A) also require a simple hardware upgrade. Please contact your local Agilent Technologies Service Center or sales representative for information on this HW upgrade. | Not applicable. E-series E9300 power sensor information is contained in the E9300 Operating and Service Guide (p/n E9300-90001: English language version). |

Firmware Version A1.05.00

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| 436A code compatibility | Only provided in single channel power meter. | 436A compatibility information will be added to E4418B shipments. Will be included in E4418B Users Guide as soon as possible. |

436A compatibility information is available in a User’s Guide Flyer (E4418-90047).

Firmware Version A1.06.00 and A2.06.00. (Note: Firmware revision A2.05.00 was not used).

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| E-series E9300 “B” and “H” power sensors compatibility | Enables EPM series power meter to use “B” and “H” versions (high power) in the E9300 sensor family. Please note that the EPM ”A” power meters (E4418A/19A) also require a simple hardware upgrade. Please contact your local Agilent Technologies Service Center or sales representative for information on the HW upgrade. | Not applicable. E-series E9300 power sensor information is contained in the E9300 Operating and Service Guide (p/n E9300-90016: English language version). |
| Longer life battery with option 001 | Enables EPM series power meters to use the new battery supplied with option 001 | Not applicable. Refer to the flyer “Battery Information”, p/n E4418-90054 |

Firmware Version A1.07.00 and A2.07.00.

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| GPIB problem fixed for 436A, 437B, and 438A emulation modes. | Corrected a problem that made it possible for remote zero or calibration tasks to never be reported as complete. | User Guide: Chapter 3 “System/Inputs Menu” Compatibility information for HP436A emulation mode is located in the User’s Guide Flyer E4418-90047 |
| Service query command improved to remove ambiguity. | Query “SERV:SENS:<etc> now returns “UNKNOWN” whilst data is being read from the sensor. Prior to this update, the command would return details of the last sensor connected. | Programming Guide: Chapter 14 “Service Subsystem” |
| Calibrator self test improved. | Added a 1.5 second delay when switching on the power ref for the calibrator self test. Without the delay, an unsettled reading could cause the self test to fail. | User Guide: Chapter 2 “Self Test”. |
| Setup information added for self test confidence check | This test would frequently fail because the cal factor or measurement frequency had not been set. | User Guide: Chapter 2 “self test” |
| E4419A recorder output fixed for instances when a sensor is not connected | A channel that does not have a sensor attached to it cannot be routed to the recorder output. Prior to this update, routing a channel to the recorder output without connecting a sensor to it would cause the power meter to hang-up. | User Guide: chapter 2 “Recorder outputs” |
| E4419A recorder output “B” maximum power setting fixed | Corrected a problem that exists in Firmware Version A2.03.00, A2.04.00 and A2.06.00. Prior to this update, recorder output “B” was permanently set for a maximum power of 100mW. | User Guide: Chapter 2 “Recorder Outputs” |
| Triggering problem fixed for P438A emulation mode. | For sensor “A” measurement, the TR1/TR2 command will trigger channel ‘A’ only. For sensor “B” measurement, the TR1/TR2 command will trigger channel ‘B’ only. Prior to this update, the TR1/TR2 command could trigger both channels if a math function was selected. | User guide: Chapter 3 “System/Inputs Menu” |

Firmware Versions A1.08.00 and A2.08.00.

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| RS422 Serial Interface. | RS422 serial communications interface can be enabled or disabled to reflect the hardware fitted. | Not applicable. Full details are available in Service Notes E4418B-09 and E4419B-09 |

Firmware Versions A1.08.01 and A2.08.01.

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| Zero/Cal Information Retained Following A Power Cycle | Corrected a problem that causes the power meter to lose user calibration data for E930x sensors following a power cycle. This calibration data is generated by the power meter during the zero and cal process, and is not related to the calibration data stored in the sensor EEPROM. Also note that this problem did not affect any sensors other than the E930x family. | User Guide: Chapter 2 “Zeroing and Calibrating the Power Meter” |

Firmware Versions A1.08.03 and A2.08.03.

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| Duty cycle and Self Test Problem | To resolve narrow pulse duty cycle problem and to skip display self test for K06 option | Not applicable |

Firmware Versions A1.08.04 and A2.08.04.

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| Calibrator fault on self test | To resolve self test failure on the calibrator test. | Not applicable |

Firmware Versions A1.09.00 and A2.09.00 release on  1 September 2008.

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| Compatibility with N8481A, N8482A & N8485A power sensors | To enable the power meter to support N8480 thermocouple power sensors | Not applicable |

Firmware Versions A1.09.01 and A2.09.01 release on 9 June 2009.

| Feature | Notes | Refer to E4418/19B Power Meter Manual |
| --- | --- | --- |
| Compatibility with N8481/H, N8482B/H, N8487A & N8486AR/Q power sensors | To enable the power meter to support N8480 thermocouple power sensors | Not applicable |
| To fix “reading freezing” | Power measurement reading will not freeze with the recorder output turn on | Not applicable |