## calculator
U2000 Series USB Power Sensor Uncertainty Calculator
Input | Value
Device Under Test (SWR) | 1.25
Sensor Model | U2000A
Frequency | 1 | GHz
Power | 0 | dBm |  | 0.001 | W
Number of Average Reading | 1024
Average mode | Normal
Internal or External Zeroing | External zero
Source of Uncertainty | Symbol | Value ± | Probability Distribution | Divisor | Result
Mismatch Gain Between Generator and Sensor | Mu | ¦Гg¦ = 0.111 | U-shape | 1.4142 | 0.4795%
¦Гs¦ = 0.061
Zero drift | D | 5e-07 | Rectangular | 1.7321 | 0.02887%
Power Sensor Calibration Factor Uncertainties | Kb | 1.6% | Gaussian | 2 | 0.8%
Power Sensor Linearity | Pl | 3% | Rectangular | 1.7321 | 1.732%
Zero set | Zs | 1e-06 | Rectangular | 1.7321 | 0.05774%
Sensor noise | N | 1e-06 | Rectangular | 1.7321 | 0.01212%
|  |  | Combined Uncertainty-RSSed | = | 1.968%
|  |  | K | = | 2
|  |  | Expanded Uncertainty | = | 3.937%
|  |  | Upper Limit Uncertainty | = | + 0.168dB
|  |  | Lower Limit Uncertainty | = | − 0.174dB
Revision 2
Date updated: 11 Sept 2008
© Agilent Technologies, Inc. 2008
The data used in this calculation is base on the specifications as in the U2000 series USB power sensors datasheet, literature number 5989-6278EN.
This calculation is base on ISO Guide to the Expression of Uncertainty in Measurement, often referred to as the GUM.
For more info about GUM, please refer to Agilent Application Note 1449-3 "Fundamental of RF and Microwave Power Measurements (Part 3), literature number 5988-9215EN.

## Data (hidden)
Max SWR (25 +/-10C)
Frequency | U2000A | U2001A | U2002A | U2004A | U2000B | U2000H | U2001B | U2001H | U2002H
9e-06 | # | # | # | 1.13 | # | # | # | # | # | 9kHz
0.01 | 1.15 | 1.15 | # | 1.13 | 1.12 | 1.15 | 1.12 | 1.15 | # | 10MHz
0.03 | 1.13 | 1.13 | # | 1.13 | 1.12 | 1.15 | 1.12 | 1.15 | # | 30MHz
0.05 | 1.13 | 1.13 | 1.13 | 1.13 | 1.12 | 1.15 | 1.12 | 1.15 | 1.15 | 50MHz
2 | 1.19 | 1.19 | 1.19 | 1.19 | 1.17 | 1.15 | 1.17 | 1.15 | 1.15 | 2GHz
6 | 1.19 | 1.19 | 1.19 | 1.19 | 1.17 | 1.15 | 1.17 | 1.15 | 1.15 | 6GHz
8 | 1.19 | 1.19 | 1.19 | 1.19 | 1.17 | 1.25 | # | # | 1.25 | 8GHz
12.4 | 1.19 | 1.19 | 1.19 | 1.19 | 1.24 | 1.28 | # | # | 1.28 | 12.4GHz
14 | 1.22 | # | 1.22 | # | 1.24 | 1.28 | # | # | 1.28 | 14GHz
16 | 1.26 | # | 1.26 | # | 1.24 | 1.28 | # | # | 1.28 | 16GHz
18 | 1.26 | # | 1.26 | # | 1.24 | 1.28 | # | # | 1.3 | 18GHz
19 | # | # | 1.3 | # | # | # | # | # | 1.3 | 19GHz
24 | # | # | 1.3 | # | # | # | # | # | 1.3 | 24GHz
Power Linearity (25+/-10C)
| U2000A | U2001A | U2002A | U2004A | U2000B | U2000H | U2001B | U2001H | U2002H
| 3% | 3% | 3% | 3% | 3.5% | 4% | 3.5% | 4% | 4%
Internal zero
External zero
Zero Set (External) | (Note: For these parameters, there are overlapping power range. In this case, worst case spec is chosen for MU calculation).
| U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | 6e-10 | 6e-10 | 6e-10 | 6e-10 |  | -30 | 8e-07 | 8e-07 |  | -50 | 8e-09 | 8e-09 | 8e-09
-38 | 1.5e-09 | 1.5e-09 | 1.5e-09 | 1.5e-09 |  | -8 | 2e-06 | 2e-06 |  | -28 | 2e-08 | 2e-08 | 2e-08
-20 | 1e-08 | 1e-08 | 1e-08 | 1e-08 |  | 10 | 1e-05 | 1e-05 |  | -10 | 1e-07 | 1e-07 | 1e-07
-11 | 5e-07 | 5e-07 | 5e-07 | 5e-07 |  | 19 | 0.001 | 0.001 |  | -1 | 2e-05 | 2e-05 | 2e-05
-7 | 1e-06 | 1e-06 | 1e-06 | 1e-06 |  | 23 | 0.002 | 0.002 |  | 3 | 3e-05 | 3e-05 | 3e-05
10 | 5e-06 | 5e-06 | 5e-06 | 5e-06 |  | 44 | 0.002 | 0.002 |  | 20 | 6e-05 | 6e-05 | 6e-05
20 | 5e-06 | 5e-06 | 5e-06 | 5e-06 | [5 empty] | 30 | 6e-05 | 6e-05 | 6e-05
Zero Set (Internal) | U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | 1.5e-09 | 1.5e-09 | 1.5e-09 | 2.8e-09 |  | -30 | 1.8e-06 | 1.8e-06 |  | -50 | 1.5e-09 | 1.5e-09 | 1.5e-09
-38 | 2e-09 | 2e-09 | 2e-09 | 3e-09 |  | -8 | 2e-06 | 2e-06 |  | -28 | 2e-08 | 2e-08 | 2e-08
-20 | 1.2e-08 | 1.2e-08 | 1.2e-08 | 1.2e-08 |  | 10 | 1.2e-05 | 1.2e-05 |  | -10 | 1.2e-07 | 1.2e-07 | 1.2e-07
-11 | 2e-06 | 2e-06 | 2e-06 | 2e-06 |  | 19 | 0.002 | 0.002 |  | -1 | 2e-05 | 2e-05 | 2e-05
-7 | 4e-06 | 4e-06 | 4e-06 | 4e-06 |  | 23 | 0.004 | 0.004 |  | 3 | 4e-05 | 4e-05 | 4e-05
10 | 6e-06 | 6e-06 | 6e-06 | 6e-06 |  | 44 | 0.004 | 0.004 |  | 20 | 6e-05 | 6e-05 | 6e-05
20 | 6e-06 | 6e-06 | 6e-06 | 6e-06 | [5 empty] | 30 | 6e-05 | 6e-05 | 6e-05
Zero Drift | U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | 2e-10 | 2e-10 | 2e-10 | 2e-10 |  | -30 | 2e-07 | 2e-07 |  | -50 | 2e-09 | 2e-09 | 2e-09
-38 | 4e-10 | 4e-10 | 4e-10 | 4e-10 |  | -8 | 4e-07 | 4e-07 |  | -28 | 4e-09 | 4e-09 | 4e-09
-20 | 1.5e-09 | 1.5e-09 | 1.5e-09 | 1.5e-09 |  | 10 | 1.5e-06 | 1.5e-06 |  | -10 | 1.5e-08 | 1.5e-08 | 1.5e-08
-11 | 5e-08 | 5e-08 | 5e-08 | 5e-08 |  | 19 | 5e-08 | 5e-08 |  | -1 | 5e-07 | 5e-07 | 5e-07
-7 | 5e-07 | 5e-07 | 5e-07 | 5e-07 |  | 23 | 0.0005 | 0.0005 |  | 3 | 5e-06 | 5e-06 | 5e-06
10 | 2e-06 | 2e-06 | 2e-06 | 2e-06 |  | 44 | 0.0005 | 0.0005 |  | 20 | 2e-05 | 2e-05 | 2e-05
20 | 2e-06 | 2e-06 | 2e-06 | 2e-06 | [5 empty] | 30 | 2e-05 | 2e-05 | 2e-05
Sensor noise | U2000A | U2001A | U2002A | U2004A |  |  | U2000B | U2001B |  |  | U2000H | U2001H | U2002H
-60 | 1e-09 | 1e-09 | 1e-09 | 1e-09 |  | -30 | 1e-06 | 1e-06 |  | -50 | 1e-08 | 1e-08 | 1e-08
-38 | 1.5e-09 | 1.5e-09 | 1.5e-09 | 1.5e-09 |  | -8 | 1.5e-06 | 1.5e-06 |  | -28 | 1.5e-08 | 1.5e-08 | 1.5e-08
-20 | 1.5e-08 | 1.5e-08 | 1.5e-08 | 1.5e-08 |  | 10 | 1.5e-05 | 1.5e-05 |  | -10 | 1.5e-07 | 1.5e-07 | 1.5e-07
-11 | 6.5e-07 | 6.5e-07 | 6.5e-07 | 6.5e-07 |  | 19 | 0.00065 | 0.00065 |  | -1 | 6.5e-06 | 6.5e-06 | 6.5e-06
-7 | 1e-06 | 1e-06 | 1e-06 | 1e-06 |  | 23 | 0.001 | 0.001 |  | 3 | 1e-05 | 1e-05 | 1e-05
10 | 1e-05 | 1e-05 | 1e-05 | 1e-05 |  | 44 | 0.001 | 0.001 |  | 20 | 0.0001 | 0.0001 | 0.0001
20 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | [5 empty] | 30 | 0.0001 | 0.0001 | 0.0001
Normal | x2
Noise multiplier
Number of avg | Normal |  | Number of avg | x2
1 | 2 |  | 1 | 2.7
2 | 1.8 |  | 2 | 2.4
4 | 1.7 |  | 4 | 2
8 | 1.5 |  | 8 | 1.6
16 | 1 |  | 16 | 1
32 | 0.95 |  | 32 | 0.91
64 | 0.74 |  | 64 | 0.78
128 | 0.55 |  | 128 | 0.53
256 | 0.39 |  | 256 | 0.34
512 | 0.29 |  | 512 | 0.29
1024 | 0.21 |  | 1024 | 0.2
Cal factor uncertainties
Frequency | U2000A | U2001A | U2002A | U2004A | U2000B | U2000H | U2001B | U2001H | U2002H
9e-06 | # | # | # | 1.8% | # | # | # | # | # | 9kHz
0.01 | 1.8% | 1.8% | # | 1.8% | 1.8% | 2% | 1.8% | 2% | # | 10Mhz
0.03 | 1.6% | 1.6% | # | 1.8% | 1.8% | 2% | 1.8% | 2% | # | 30MHz
0.05 | 1.6% | 1.6% | 2% | 1.8% | 1.8% | 2% | 1.8% | 2% | 2.5% | 50MHz
0.5 | 1.6% | 1.6% | 2% | 1.8% | 1.8% | 2% | 1.8% | 2% | 2.5% | 500MHz
1.2 | 1.6% | 1.6% | 2% | 1.8% | 1.8% | 2% | 1.8% | 2% | 2.5% | 1.2GHz
2 | 2% | 2% | 2.5% | 1.8% | 2% | 2% | 2% | 2% | 2.5% | 2GHz
6 | 2% | 2% | 2.5% | 1.8% | 2% | 2% | 2% | 2% | 2.5% | 6GHz
8 | 2% | 2% | 2.5% | 1.8% | 2% | 2% | # | # | 2.5% | 8GHz
12.4 | 2% | 2% | 2.5% | 1.8% | 2.2% | 2.2% | # | # | 2.7% | 12.4GHz
14 | 2.2% | # | 2.7% | # | 2.2% | 2.2% | # | # | 2.7% | 14GHZ
16 | 2.2% | # | 2.7% | # | 2.2% | 2.2% | # | # | 2.7% | 16GHz
18 | 2.2% | # | 3% | # | 2.2% | 2.2% | # | # | 3% | 18GHz
24 | # | # | 3% | # | # | # | # | # | 3% | 24GHz
rho DUT | 0.11111
rho sensor | 0.061033 | 1.13
Multiplier | 0.21
Rectangular | 1.7321
Gaussian | 2
U-shape | 1.4142
U2000A |  |  | U2000A
U2001A |  |  | U2001A
U2002A |  |  | U2002A
U2004A |  |  | U2004A
U2000B |  |  | U2000B
U2000H |  |  | U2000H
U2001B |  |  | U2001B
U2001H |  |  | U2001H
U2002H |  |  | U2002H
Zero Set
External zero | 1e-06
Internal zero | 4e-06