Understanding OmniBER 718 Option 210

- The Problem:  Your design must meet the 100 mUI peak-peak Jitter Generation requirement of GR-253.

Jitter Generation refers to the intrinsic jitter generated within a Network Element (NE).  It is measured at the output of the NE, with no jitter or wander applied at the input.

The 100 mUI peak-peak jitter requirement applies to Category II interfacesError: Reference source not found, which include all OC-N optical and STS-N electrical interfaces to SONET NEs.

Jitter is a noise process, so it is statistical in nature.  Maximum peak-peak jitter is accumulated and measured over a 60 second interval.  Low- and high-pass filters are also used, to limit the measurement to the jitter frequency range of interest.

For OC-48, GR-253 specifies

12 kHz high-pass filter

20 MHz low-pass filter

100 mUI

100 mUI

(GR-253)

# GR-253

Cat II Jit Gen

Requirement

- A possible solution?

What about using a tester with jitter measurement accuracy per O.172?

Most SONET jitter testers available today specify the accuracy of their jitter measurement according to ITU-T recommendation O.172.

O.172 specifies jitter measurement accuracy in a number of ways, all in terms of maximum peak-peak jitter using a concatenated SDH (SONET) signal with a 223 – 1 PRBS payload.  The total measurement error allowable comprises fixed and variable components:

±R% of reading ±W

- R is the variable component, expressed as a % of the reading.  R varies according to the frequency of jitter present in the signal.  At STM-16 (OC-48), R ranges from ±7% for low jitter frequencies to ±20% for high jitter frequencies.

- W is the fixed error, expressed in UI.  This is the intrinsic jitter of the tester’s receiver.

R

W

At OC-48, O.172 f3-f4 specifies max 50 mUI fixed error over 1 MHz – 20 MHz frequency range.

GR-253 tester requirements:

Frequency does not meet the GR-253 12 kHz – 20 MHz requirement

Fixed error contributes up to half of the GR-253 100 mUI budget

At OC-48, O.172 f1-f4 specifies max 100 mUI fixed error over 5 kHz – 20 MHz frequency range.

GR-253 tester requirements:

Frequency range exceeds the GR-253 12 kHz – 20 MHz requirement

Fixed error swamps GR-253 100 mUI budget

GR-253 12 kHz high-pass filter shown in red

- O.172 jitter measurement accuracy cannot guarantee compliance with GR-253!

100 mUI

100 mUI

(GR-253)

# GR-253

Cat II Jit Gen

Requirement

100 mUI

tester

intrinsic

jitter

20%

7%

50 mUI

tester

intrinsic

jitter

20%

10%

# O.172

f1–f4

# O.172

f3–f4

(narrow band)

- The best solution!

Select from two OmniBER 718 models with superior jitter intrinsic performance!

For some time now, OmniBER has delivered the industry’s best jitter intrinsic performance in the shape of option 200.  Typically, this allows designers to work to a 30 mUI safe design margin for peak-peak jitter.

Now, OmniBER also offers a new version – option 210 – that guarantees receiver jitter intrinsics, 30% lower than its current best-in-class levels.  You can now work, with even more confidence, to a 45 mUI safe design margin for peak-peak jitter.

With this significant increase in safe design margin, you can save R&D time, and so benefit from reduced costs and faster time-to-market!

R

W

At OC-48, OmniBER option 210‘s fixed error contributes only up to a third of the GR-253 100 mUI budget

At OC-48, both OmniBER jitter options are specified over the full O.172 f1-f4 bandwidth (5 kHz – 20 MHz), exceeding the GR-253 12 kHz – 20 MHz requirement

Choose OmniBER to prove your design’s output jitter performance meets GR-253!

100 mUI

100 mUI

(GR-253)

# GR-253

Cat II Jit Gen

Requirement

100 mUI

tester

intrinsic

jitter

20%

7%

50 mUI

tester

intrinsic

jitter

20%

10%

# O.172

f1–f4

# O.172

f3–f4

(narrow band)

50 mUI

tester

intrinsic

jitter

20%

7%

35 mUI

tester

intrinsic

jitter

20%

7%

# OmniBER

Option 210

f1–f4

30 mUI safe

design margin

45 mUI safe

design margin

# OmniBER

Option 200

(typical)

f1–f4

## The Full Story…

Bandwidth Comparison

|  | GR-253 |  | O.172 f1-f4 |  |  | O.172 f3-f4Error: Reference source not found |  |  | OmniBER 718 options 200 & 210 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | HP | LP | f1 HP | f4 LP | Meets GR-253? | f3 HP | f4 LP | Meets GR-253? | HP | LP | Meets GR-253? |
| OC-48 | 12 kHz | 20 MHz | 5 kHz | 20 MHz |  | 1 MHz | 20 MHz |  | 5 kHz | 20 MHz |  |
| OC-12 | 12 kHz | 5 MHz | 1 kHz | 5 MHz |  | 250 kHz | 5 MHz |  | 1 kHz | 5 MHz |  |
| OC-3 | 12 kHz | 1.3 MHz | 500 Hz | 1.3 MHz |  | 65 kHz | 1.3 MHz |  | 500 Hz | 1.3 MHz |  |
| STS-3 | 12 kHz | 1.3 MHz | 500 Hz | 1.3 MHz |  | 65 kHz | 1.3 MHz |  | 500 Hz | 1.3 MHz |  |
| OC-1 | 12 kHz | 400 kHz | 100 Hz | 400 kHz |  | 20 kHz | 400 kHz |  | 100 Hz | 400 kHz |  |
| STS-1 | 12 kHz | 400 kHz | 100 Hz | 400 kHz |  | 20 kHz | 400 kHz |  | 100 Hz | 400 kHz |  |

In addition to the high-pass filters listed above, OmniBER also offers a 12 kHz high-pass filter at all rates, convenient for GR-253 peak-peak and rms jitter measurements.

Peak-Peak Jitter Comparison

|  | GR-253 | O.172 f1-f4 |  |  | OmniBER 718 opt 200 |  |  | OmniBER 718 opt 210 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Max pk-pk | W | Max R | Meets GR-253? | W (typical) | Max R | Meets GR-253? | W | Max R | Meets GR-253? |
| OC-48 | 100 mUI | 100 mUI | ±20% |  | 50 mUI | ±20% |  | 35 mUI | ±20% |  |
| OC-12 | 100 mUI | 100 mUI | ±15% |  | 50 mUI | ±15% |  | 35 mUI | ±15% |  |
| OC-3 | 100 mUI | 70 mUI | ±10% |  | 35 mUI | ±10% |  | 35 mUI | ±10% |  |
| STS-3 | 100 mUI | 70 mUI | ±10% |  | 35 mUI | ±10% |  | 35 mUI | ±10% |  |
| OC-1 | 100 mUI | 70 mUI | n/a | ? | 35 mUI | ±8% |  | 35 mUI | ±8% |  |
| STS-1 | 100 mUI | n/a | n/a | ? | 35 mUI | ±8% |  | 35 mUI | ±8% |  |

RMS Jitter Comparison

|  | GR-253 | O.172 f1-f4 | OmniBER 718 options 200 & 210 |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Max rmsError: Reference source not found | W | W (typical) | Max R | Meets GR-253? |
| OC-48 | 10 mUI | n/a | 4 mUI | ±20% |  |
| OC-12 | 10 mUI | n/a | 4 mUI | ±15% |  |
| OC-3 | 10 mUI | n/a | 4 mUI | ±10% |  |
| STS-3 | 10 mUI | n/a | 4 mUI | ±10% |  |
| OC-1 | 10 mUI | n/a | 4 mUI | ±8% |  |
| STS-1 | 10 mUI | n/a | 4 mUI | ±8% |  |

O.172 treats rms jitter measurement as optional, so it contains no accuracy specification for rms jitter measurement.

Telcordia Technologies GR-253-CORE Issue 3 September 2000, section 5.6.2.3.6 Category II Jitter Generation, requirement R5-258.

See GR-253-CORE section 5.6.1 Network Interface Jitter Criteria.

See GR-253-CORE section 5.6.2.3 Jitter Generation.

ITU-T O.172 (03/2001), section 9.4 Measurement accuracy.

30% reduction in receiver jitter intrinsics applies at OC-48 and OC-12 data rates.

Figures from GR-253-CORE section 5.6.2.3 and Table 5-9.

Figures from ITU-T O.172 Table 7/O.172.

Figures from GR-253-CORE section 5.6.2.3.6, requirement R5-258.

Figures from ITU-T O.172 Table 8/O.172.

Figures from ITU-T O.172 Table 10/O.172.

See ITU-T O.172 section 9.2.3 Measurement of RMS jitter.