# IFR 4000 Test Set Option Features

## Section 1: Overview

The IFR 4000 allows for optional ELT software components to be activated to the base software install.

## Section 2: Software options management

To allow end-users to purchase and add ELT option to the 4000, the unit supports modification of the options list via an encrypted file.  This file is created using Aeroflex Wichita intranet applications (IFR Software Options).  It contains the lists of options values to be changed, and the serial number for which the options file is valid.  It is encoded using RSA asymmetric key encryption, using a private key known to IFR and decrypted using a public key stored within the instrument.

## Section 3: RCI commands

OPTION:APPLY		// to download option file to IFR 4000 through RS232 using comms. software

OPTION:APPLY?	// to query if ELT option is installed

#### Section 3: Examples

To download option file to IFR 4000:

Match 4000 RS-232 setting to settings on communication program Tera Term, MTTTy, or other communication program.  If using MTTY program, click ‘Connect’ under ‘File’.

Type 'option:apply' from IFR 4000 remote control interface (RCI) via terminal program.

Then 4000 will prompt 'Begin download of options now....'

Use comm. program menu "send file" to send option file (10300xxxx.new) to IFR 4000.

Make sure the file transfer is in binary mode and transmit with CR only (not CR LF).