Using Two Model 2400 SourceMeters to

Output up to Two Amps

Discussion

Each Model 2400 SourceMeter  can source up to one ampere of current and read the developed voltage. In applications that require more than one ampere the following method can be used to source up to two amperes. To accomplish this, the SOURCE OUTPUTS of two Model 2400 SourceMeters are connected in parallel as shown in Figure 1. Sourcing more than two amperes could be done with more SourceMeters.

Method

1. Configure the circuit as shown in Figure 1.

2. Select the SOURCE I mode on each Model 2400 SourceMeter. Select one amp of current on unit #1 and vary the current output of unit #2 to the desired total current.

3. Set the compliance voltage on each SourceMeter to the maximum compliance expected in the circuit. In the case of  Figure 1 this would be two volts. (1 ohm x 2A = 2V max)

4. Use CONFIG V and select 4-WIRE voltage sensing.

5. Make sure the leads and the terminal switch match. (e.g. Front for Front).

6. Press OPERATE on each unit.

7. Monitor the voltage levels. Notice that the voltage reads approximately 2 volts.

8. Press OPERATE on each unit to turn off  current.

Note: The second set of SENSE lines are not actually necessary. One set is sufficient to monitor the compliance voltage level. The sense lines on unit #2 can be removed and 2-WIRE sensing can be used.  The voltage reading Unit #2 would include the test lead resistance voltage drop and therefore would be higher than the reading on Unit #1. The voltage level on Unit #1 would be the most accurate because of the 4-WIRE sensing.

Note: Don’t  use the resistance mode on either Model 2400. The calculation of resistance is accomplished by using the forced current and the read voltage. Each unit is forcing a different current. In this example the resistance would be calculated with 1 Amp, but the voltage drop would be 2V. Therefore the resistance would be calculated as 2 ohms.

Settings:

SOURCE 		= I = 1.00A

MEASURE 		= V = Measured value

COMPLIANCE	= 2.1V

SENSE MODE	= 4-WIRE

DEVICE		= 1 Ohm resistor

UNIT #1

UNIT #2

Model 2400

Model 2400

OUTPUT HI

SENSE HI

OUTPUT LO

SENSE LO

SENSE LO

SENSE HI

Device Under Test

Figure 1. Connecting two Model 2400s in parallel