KEITHLEY

Safely Using the Interlock on the Keithley Model 4200-SCS

Safety interlock connections

WARNING 	It is strongly recommended that you use the safety interlock

circuit to avoid personal injury or death caused by hazardous

voltages.

The safety interlock feature on the Model 4200-SCS should be used to avoid possible shock hazards. It provides a means by which the outputs of the Model 42XX-SMUs can be automatically placed in a safe state regardless of the state of the Model 4200-SCS operating software. When the safety interlock signal is asserted (connected to +12V) all of the voltage ranges of the SMUs will be functional. However, when the safety interlock signal is not asserted, the 200V range on the SMUs will be disabled, limiting the nominal output to ± 20V. Under these conditions, all SMU and PreAmp signal terminals will be non-hazardous. Component test fixtures and probe station dark boxes typically have a safety switch that can be interfaced to the Model 4200-SCS safety circuit as outlined in the following paragraphs.

Interlock connector

Figure 1-1 shows the location of the safety interlock connector.

NOTE	The 4200-TM IN and OUT terminals shown in Figure 1-1 are not currently supported. Connecting cables to these terminals can cause unexpected system operation.

Figure 1-1

Interlock connector location

Model 4200-SCS Rear Panel

Interlock cables

Use the supplied interlock cable (236-ILC-3) or equivalent to make interlock connections whenever possible.

Test fixtures

There are two types of test fixtures for the Model 4200-SCS: low voltage fixture (less than ±20 volts) and high voltage (greater than ±20 volts). High voltage fixtures require extra precaution to ensure there are no shock hazards. Whenever the interlock of the Model 4200-SCS is asserted, the FORCE and GUARD terminals of the SMUs and PreAmps should be considered high voltage, even if they are programmed to a non-hazardous voltage and/or current.

Testing with less than ±20 volts

For testing discrete devices, a test fixture equipped with 3-lug triax connectors is necessary to allow the Model 4200-SCS to be connected to the discrete device. Figure 1-2 shows a basic test fixture to test a 2-terminal device. The test fixture’s exterior enclosure should be constructed of metal and the metal should be connected to COMMON. The DUT should be mounted on test terminals that are insulated using a material that has high resistivity such as Teflon. Guarding will improve the quality of the measurement by reducing leakage and parasitic capacitance. The Keithley “Low Level Measurement” handbook provides an in-depth discussion on guarding and other techniques that are useful for building quality test fixtures. Contact a Keithley sale or service office to obtain a copy.

NOTE The Model 4200-SCS will function on all current ranges and up to ±20 volts without the interlock being asserted. The maximum voltage on the SMU and PreAmp terminals is not hazardous when the interlock is not asserted.

4200-SCS User’s Manual Installation 2-11

Figure 1-2

Typical Test Fixture Connections.

Testing with more than ±20 volts

If voltages greater than ±20 volts are needed for testing, then the above step should be followed and the following additional steps are required. An interlock switch must be added to the fixture to ensure that hazardous voltages are not present when the fixture’s exterior enclosure is open and to enable the Model 4200-SCS to output higher voltages when the fixture’s exterior enclosure is closed. In addition, the exterior enclosure must be connected to COMMON and/or safety ground using #18 AWG wire or greater. Care must be taken to ensure that the wiring (FORCE, GUARD, and SENSE) within the fixture do not electrically contact the exterior enclosure. For more details on the Model 4200-SCS interlock system, see Reference Manual “Control and data connections,” page 4-23.

WARNING 	Asserting the interlock will allow the SMU and PreAmp terminals to become hazardous. SMU and PreAmp terminals should be considered hazardous even if the outputs are programmed to be low volt. Precautions must be taken to prevent a shock hazard by surrounding the test device and any unprotected leads (wiring) with double insulation for 250 volts, Category I. See EN61010-1 safety standards for details.

Typical interlock connections

Figure 1-3 shows typical interlock connections. In this example, the Model 4200-SCS is connected to a component test fixture. The test fixture has a safety interlock switch connected to its lid.  When the lid is closed, the interlock circuit is closed (asserted), and SMU ± 200V ranges are enabled. Conversely, the interlock circuit is open (de-asserted) when the lid is open, and the SMU ± 200V ranges are disabled. A safety interlock cable is supplied with the 4200-SCS, allowing it to directly interface the interlock circuits of the Keithley Models 8006 and 8007 test fixtures.

Figure 1-3

Typical interlock connections

Interlock connector wiring

Figure 1-4 shows typical interlock connector wiring. Note that a normally open switch should be used. And open interlock condition occurs when the switch is open.

WARNING 	Be sure the interlock switch is operating correctly to assure

proper, safe interlock operation.

Figure 1-4

Interlock connector wiring

NOTE	Older versions of the Model 4200-SCS Reference Manual may indicate Pin 3 = NC.  Please, be aware that Pin 3 is actually connected to Ground.  Inadvertently shorting Pin 1 or Pin 3 to ground may result in damage to the interlock circuit.

Making custom connections to the interlock

From Figure 1-4, notice that:

Pin 1 = Interlock

Pin 2 = +12V

Pin 3 = Ground

WARNING	Avoid making electrical contact between Pin 3 and the other two pins.  Connecting either Pin 1 or Pin 2 to Ground (Pin 3) will result in damage to the interlock circuit.

WARNING	Do not create a permanent link between Pins 1 and 2 to effectively disable the interlock feature of the Model 4200-SCS.

Users may wish based on their own needs and against commonly accepted safety procedures to permanently jumper across Pins 1 and 2.  A permanent jumper across these pins will permanently enable the ±200V ranges on the SMUs without the need or hassle of an external switch.  While this does not meet safety regulations, Keithley recognizes that permanently disabling the interlock is a possible use condition for the Model 4200-SCS.

When making any connections to the interlock connector, care must be taken not to make electrical contact between either Pin 1 or Pin 2 and Pin 3. If the user accidentally connects Pin 2 to Pin 3, they are connecting 12V directly to Ground, which may damage the interlock circuit.

WARNING	When making modifications to the interlock connector or cabling ensure that the 4200-SCS is powered off and all cables removed from the rear panel.  Accidentally shorting the interlock pins while making the modifications may result in damage to the interlock circuit.