MODEL A10-B

###### Rubidium Time & Frequency Standard

USER’S HANDBOOK

Contents

Safety Considerations	3

Voltage, Frequency and Power Characteristics	4

Environmental Conditions	4

Replaceable Fusing Characteristics	4

Cleaning Instructions	4

Rubidium Frequency Standard A10-B	5

Getting Started	5

Rubidium Unit Connections	6

Rubidium Frequency Standards	7

Appendix	8

# Safety Considerations

#### General

This product and related documentation must be reviewed for familiarisation before operation. If the equipment is used in a manner not specified by the manufacturer, the protection provided by the instrument may be impaired.

#### Before Applying Power

Verify that the product is set to match the available line voltage and the correct fuse is installed.

#### Before Cleaning

Disconnect the product from operating power before cleaning.

#### WARNING

Bodily injury or death may result from failure to heed a warning. Do not proceed beyond a warning until the indicated conditions are fully understood and met.

#### CAUTION

Damage to equipment, or incorrect measurement data, may result from failure to heed a caution. Do not proceed beyond a caution until the indicated conditions are fully understood and met.

#### This equipment must be earthed

An uninterruptible safety earth ground must be maintained from the mains power source to the product’s ground circuitry.

#### WARNING

When measuring power line signals, be extremely careful and use a step down isolation transformer whose output is compatible with the input measurement capabilities of this product. The product’s front and rear panels are typically at earth ground. Thus, never try to measure AC power line signals without an isolation transformer.

#### WARNING

Instructions for adjustments when covers are removed and for servicing are for use by service-trained personnel only. To avoid dangerous electrical shock, do not perform such adjustments or servicing unless qualified to do so.

#### WARNING

Any interruption of the protective grounding conductor (inside or outside the instrument) or disconnecting of the protective earth terminal will cause a potential shock hazard that could result in personal injury. Grounding one conductor of a two conductor out-let is not sufficient protection.

Whenever it is likely that the protection has been impaired, the instrument must be made inoperative and be secured against any unintended operation.

If the instrument is to be energised via an autotransformer (for voltage reduction) make sure the common terminal is connected to the earthed pole terminal (neutral) of the power source.

Instructions for adjustments while the covers are removed and for servicing are for use by service-trained personnel only. To avoid dangerous electrical shock, do not perform such adjustments or servicing unless qualified to do so.

For continued protections against fire, replace the line fuse(s) with fuses of the same current rating and type (for example, normal blow time delay). Do not use repaired fuses of short-circuited fuse holders.

# Voltage, Frequency and Power Characteristics

Voltage 220-240V AC

Frequency 40-50Hz

Power characteristics 500mA Max

# Environmental Conditions

#### Temperature

| Operating (ambient) | -10°C to +55°C (-65 to +65 op) |
| --- | --- |
| Storage | -40°C to +85°C |

#### Magnetic Field

| Sensitivity | ≤2x10-11/ Gauss |
| --- | --- |
| Atmospheric Pressure | -60m to 4000m |
|  | <1x10-13/ mbar |

# Replaceable Fusing Characteristics

800mA time-lag HBC

# Cleaning Instructions

To ensure long and trouble operation, keep the unit free from dust and use care with liquids around the unit.

Be careful not to spill liquids onto the unit. If the unit does get wet, turn the power off immediately and let the unit dry completely before turning it on again.

Clean with a damp (with water) cloth.

Never spray cleaner directly onto the unit or let liquid run into any part of it. Never use harsh or caustic products to clean the unit.

# Rubidium Frequency Standard A10-B

### Getting Started

Unpack the unit with its associated power supply, and check that there is no obvious damage.

Connect the lead from the power supply terminated in the XLR connector, to the A10-B’s rear socket, marked “24V DC ext. PSU.” Also connect the Euro connector, at the input to the PSU., to the mains supply.

Switch on the mains supply. The green “ON” LED at the lower right-hand side of the front panel should light, confirming that DC power is being fed to the unit. Initially the two red lights marked “BIT” and “Lamp” will also be illuminated.

Wait until both of the red LED’s go off (three to five minutes if the unit is being run from cold) when the internal circuitry will be correctly locked and the unit will be ready for use. For optimum accuracy and stability run the unit for approximately one (1) hour.

1 MHz, 5 MHz and 10 MHz, in both sine and square forms, are available at the front six BNC connectors - as marked. The 1 PPS output (at the lower, left-hand corner of the front panel) may be synchronised with an external, 1 PPS source by feeding synchronising pulses to the “SYNC.” Input socket, located at the top, left hand corner of the front panel.

The six BNC sockets, mounted on the back of the unit each normally supply 10 MHz sine outputs for distribution purposes - 50 output impedance’s. They may, however, be configured, at the customer’s request, to give identical outputs at any one of 10 MHz, 5 MHz or 1 MHz.

“SYNC” and “LOCK” rear panel, BNC. Connections may be supplied for use in conjunction with GPS systems and are not required for normal frequency standard operation.

As it is often important that the unit should remain operative when the mains supply fails, a battery back up unit (BBU) has been incorporated to maintain operational continuity.

If this facility is required, connect a 24V battery supply to the “24V DC BBU”, XLR connector on the back panel of the unit - the pin connections are indicated on the panel. When the mains fails and the power supply’s output voltage falls below approximately 22V, the 24V battery supply will be directly switched in to maintain the supply to the rubidium unit and its peripherals. When the mains supply is restored, the battery will be disconnected and the unit will continue to run from the mains-driven PSU.

### Rubidium Unit Connections

There are two forms of rubidium unit, which may be accommodated within the A10-B. The chassis base diagram(s) will be amended when the fixing details are finalized:

9-Pin, “D” - Connector. R.F. Output - S.M.A. Socket.

| Pin 1 | Lock Monitor |
| --- | --- |
| Pin 2 | External Frequency Adjustment |
| Pin 3 | No connection |
| Pin 4 | Rubidium Lamp Monitor |
| Pin 5 | No connection |
| Pin 6 | +22 to +30V DC Input |
| Pin 7 | VCXO Monitor |
| Pin 8 | Common Negative |
| Pin 9 | +22 to +30V DC Input |

Mixed “D” - Connector.

| A1 | RF Output (coaxial) |
| --- | --- |
| Pin 1 | Lock Monitor |
| Pin 2 | External Frequency Adjustment |
| Pin 3 | External Frequency Adjustment - Return |
| Pin 4 | Rubidium Lamp Monitor |
| Pin 5 | No Connection |
| Pin 6 | +22 to +24V DC Input |
| Pin 7 | VCXO Monitor |
| Pin 8 | Regulated output (for Ext. Freq. Adj.) |
| Pin 9 | +22 to +24V DC Input |
| Pin 10 | Common Negative |

MONITOR BOARD - IF FITTED.

WIRING FOR JP1:

| Pins 1, 2 and 3 | No Connection. |
| --- | --- |
| Pins 4, 7 and 9 | Common Return. |
| Pin 5 | +24V DC |
| Pin 6 | To VCXO Monitor on Rb. |
| Pin 8 | To BIT/LOCK Monitor on Rb |
| Pin 10 | To Rb Lamp Monitor on Rb. |

The wiring detail for JP4 “Remote Fault Detection” is not yet finalized.

# Rubidium Frequency Standards

A Rubidium frequency standard owes its outstanding accuracy and superb stability to a unique frequency control mechanism. The resonant transition frequency of the Rb 87 atom (6,834,682,614 Hz) is used as a reference against which an OCXO output is compared. The OCXO output is multiplied to the resonance frequency and is used to drive the microwave cavity where the atomic transition is detected by Electro-optical means. The detector is used to lock the OCXO output ensuring its medium and long-term stability.

The first realised Rubidium frequency standard arose out of the work of Carpenter (Carpenter et al 1960) and Arditi (Arditi 1960). It was a few years until the first commercial devices came onto the market and this was primarily due to the work of Packard and Schwartz who had been strongly influenced by the work of Arditi a few years before on Alkali atoms (of which Rb 87 is one). Unlike much of the research done into frequency standards at that time, practical realization of a Rubidium maser was high on the researchers agenda. This was mainly due to an understanding that such a device would have extremely good short-term stability relative to size and price. In 1964, Davidovits brought such research to fruition, with the first operational Rubidium frequency standard.

The Rubidium frequency standard, like its more expensive cousin, the Hydrogen maser, may be operated either as a passive or as an active device. The passive Rubidium frequency standard has proved the most useful, as it may be reduced to the smallest size whilst retaining excellent frequency stability. The applications for such a device abound in the communication, space and navigation fields.

The Rubidium frequency standard may be thought of as consisting of a cell containing the Rubidium in its vapour state, placed into a microwave cavity resonant at the hyperfine frequency of the ground state. Optical pumping ensures state selection. The cell contains a buffer gas primarily to inhibit wall relaxation and Doppler broadening. The Rubidium frequency standard essentially consists of a voltage controlled crystal oscillator, which is locked to a highly stable atomic transition in the ground state of the Rb 87 atom.

There are several reasons why Rubidium has an important role to play as a frequency standard. Perhaps most importantly is its accuracy and stability. Accuracy is comparable with that of the standard Caesium with an operating life approximately 5 times that of Caesium. Moreover the stability of a Rubidium frequency standard over short time-scales -100s of seconds- betters that of Caesium (Caesium is more stable over longer time periods, in the regions of hours to years).

There are, however, a few drawbacks to the use of Rubidium as a frequency standard. In the past, these included the limited life of the Rubidium lamp (since improved to >10 years), The Caesium is affected to a greater degree than this, whilst the Hydrogen Maser operates differently and is not affected. The thermal stability of Rubidium is inferior to that of Caesium or Hydrogen Masers, and the Rubidium previously required frequency access to a primary reference signal or synchronization source to maintain long-term Caesium level accuracy.

The cost of a Rubidium frequency standard is significantly cheaper than a Caesium, with a much reduced size and weight. Due to its small size, low weight and environmental tolerance the Rubidium frequency standard is ideal for mobile applications. Indeed, Rubidium atomic clocks are beginning to be implemented into the new generation of GPS satellites. This is in part due to the extended life of the Rubidium physics package compared to that of Caesium. The Rubidium is also extremely quick to reach operational performance, within 10 minutes reaching 5 parts in 10.

# Appendix

# Text boxes

Quartzlock UK Ltd

Gothic,

Plymouth Road,

Totnes,

Devon

TQ9 5LH,

England.

Tel:	+44 (0) 1803 862062

Fax:	+44 (0) 1803 867962

Web: 	www.quartzlock.com

E-mail:	quartzlock@quartzlock.com