The P450  Rubidium Reference Oscillator  provides a high stable signal for time-frequency measurements and reference signal for various measuring systems.

## Serial Number _______________________

# Date of Manufacture __________________

| Technical Data | Value allowed | Value measured |
| --- | --- | --- |
| Output: Frequency, MHz Voltage (into 50 ohms) | 10 0.5-1.0Vrms |  |
| Adjustment Mechanical Range Electrical Range Control Voltage | 2E-9 (min) 2E-9 (min) 0 - +5V |  |
| Accuracy at Shipment | +/- 5E-11 |  |
| Aging: 1 month 1 year | +/- 1E-10 +/- 5E-10 |  |
| Frequency Stability: 1 s 10 s 100 s | 3E-11 1E-11 3E-12 |  |
| Phase Noise, dBc/√Hz: 10 Hz 100 Hz 1 kHz 10 kHz | 100 125 135 145 |  |
| 7. Distortion, dBc: Harmonic Non-harmonic | -40 -80 |  |
| 8. Warm-up Time to 1E-9 | 5 min |  |
| Power Supply: Input voltage Power at warm up (16 mins) Power at steady state at +25°C | 22-30 45W 10W |  |
| 10. Temperature Operating (Ambient) Base plate, not more than Storage Frequency offset over operating temperature range | -10°C - +55°C +65°C -40°C - +85°C +/- 5E-10 |  |
| 11. Warranty | 1 year |  |

Quartzlock Limited Warranty

Quartzlock (UK) Ltd warrants each instrument it manufactures to be free from defects in material and workmanship.  The obligation under this warranty is limited to repairing or replacing any instrument or part thereof that shall be returned to us by the original purchaser, transportation charges prepaid, when upon examination it is disclosed to our satisfaction to be defective.  This warranty is effective for one year after delivery to the original purchaser.  If the fault has been caused by misuse or abnormal conditions of operation, repair will be made at cost.  In this case, an estimate will be submitted before work is started.

If any fault develops, notify Quartzlock UK Ltd and give full details including model and serial numbers.  You will then be notified as to the disposition of the defective instrument

This warranty is in lieu of all other warranties expressed or implied, including any warranty of fitness for a particular purpose.  The rights and remedies provided herein are exclusive and in lieu of any other rights or remedies.  In no event shall the seller be liable for consequential damage.

# Quartzlock UK Ltd

Gothic – Plymouth Road – Totnes – Devon –TQ9 5LH – UK

Tel: +44 (0) 1803 862062 Fax: +44 (0) 1803 867962

E-mail: quartzlock@quartzlock.com

Authorised Signatory……………………………