R&S Spectrum Analyzer FSU (VISA)

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for R&S Spectrum Analyzer FSU (VISA).

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

To successfully use this module, the following conditions must be met:

For GPIB instrument drivers:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

For VXI-11 instrument drivers:

-  the instrument is connected to the ethernet interface.

-  the IP address supplied to the initialize function must match the

settings of the instrument.

The RSIB interface is a R&S defined protocol for remote control of R&S

spectrum and network analyzers. Refer to the R&S application notes.

________________________________________________________________________________

Error and Status Information:

Each function in this instrument driver returns a status code that

either indicates success or describes an error or warning condition.

Your program should examine the status code from each call to an

instrument driver function to determine if an error occurred.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

The description of each instrument driver function lists possible error

codes and their meanings

________________________________________________________________________________

How To Use This Document:

Use this document as a programming reference manual.

It describes each function in the

R&S Spectrum Analyzer FSU (VISA)

instrument.  The functions appear in alphabetical order, with a description

of the function and its C syntax, a description of each parameter, and a list

of possible error codes.

________________________________________________________________________________

Function Tree Layout:

Class/Panel Name:                                           Function Name:

Initialize                                               rsfsu_init

Application Functions

Read Spectrum                                         rsfsu_appReadSpectrum

Configuration Functions

General Device Settings

Frequency Group

Center Freq and Span                            rsfsu_confSpanCenterFrq

Start And Stop Freq                             rsfsu_confStartStopFrq

Frequency Offset                                rsfsu_confFreqOffset

Signal Track                                    rsfsu_confSignalTrack

Frequency Axis Mode                             rsfsu_confFreqAxisMode

Amplitude Group

Reference Level                                 rsfsu_confRefLevel

Reference Level Opt                             rsfsu_confRefLevelOpt

Level Range                                     rsfsu_confLevelRange

Grid Mode                                       rsfsu_confGridMode

Input Group

RF Input Attenuator                             rsfsu_confRFInAtt

RF Input Attenuator Auto                        rsfsu_confRFInAttAuto

RF Input Protection Reset                       rsfsu_confRFInputProtectRst

RF Input Impedance                              rsfsu_confRFInImp

RF Input Coupling                               rsfsu_confRFInputCoupling

RF Input Electronic Attn                        rsfsu_confRFInEAtt

RF Input Electronic Attn Auto                   rsfsu_confRFInEAttAuto

RF Input Electronic Attn State                  rsfsu_confRFInEAttState

RF Input Gain                                   rsfsu_confRFInGain

RF Input YIG Filter State                       rsfsu_setRFInYIGFilterState

RF Input YIG Filter Temp Corr                   rsfsu_setRFInYIGFilterTempCorr

External Mixer

External Mixer                               rsfsu_confExtMix

External Mixer LO Level                      rsfsu_confExtMixerLOLevel

External Mixer Signal                        rsfsu_confExtMixSignal

External Mixer Parameters                    rsfsu_confExtMixParameters

Default Conversion Loss                      rsfsu_confDefConvLoss

Conversion Loss Table                        rsfsu_confExtMixLossTab

Conversion Loss Table Delete                 rsfsu_confExtMixLossTabDelete

Marker Group

Marker State                                    rsfsu_confMarkState

Marker Position                                 rsfsu_confMarkPosition

Marker Opt                                      rsfsu_confMarkOpt

Delta Marker State                              rsfsu_confDeltaMarkState

Delta Marker Position                           rsfsu_confDeltaMarkPosition

Delta Marker Opt                                rsfsu_confDeltaMarkOpt

Measure Group

Summary Marker (Abs)                            rsfsu_confSumMark

Summary Marker (Rel)                            rsfsu_confSumMarkRel

Lines Group

Display Lines                                   rsfsu_confDLines

Limit Lines State                               rsfsu_confLimitLineState

Set Limit Line Spacing                          rsfsu_actSetLimitLineSpacing

Set Param Limit Line Spacing                    rsfsu_actSetParamLimitLineSpacing

Limit Lines Data                                rsfsu_confLimitLineData

Limit Lines Switch                              rsfsu_confLimitLineSwitch

Limit Lines Shift                               rsfsu_confLimitLineShift

Limit Lines Trace                               rsfsu_confLimitLineTrace

Limit Lines Mode                                rsfsu_confLimitLineMode

Limit Lines Units                               rsfsu_confLimitLineUnits

Limit Lines Domain                              rsfsu_confLimitLineDomain

Limit Lines Offset                              rsfsu_confLimitLineOffset

Limit Lines Margin                              rsfsu_confLimitLineMargin

Limit Lines Threshold                           rsfsu_confLimitLineThreshold

Limit Lines Measurement Type                    rsfsu_confLimitLineMeasType

Trace Group

Trace Setup                                     rsfsu_confTrace

Detector                                        rsfsu_confSAMDetectorExt

Trace Math                                      rsfsu_confSAMTraceMathExt

Trace IQ Group

Trace IQ State                               rsfsu_confTraceIQState

Trace IQ Set                                 rsfsu_confTraceIQ

Trace IQ Sampling Rate                       rsfsu_confTraceIQSrate

Trace IQ Averaging                           rsfsu_configureTraceIQAver

Sweep Group

Sweep                                           rsfsu_confSweep

Get Sweep Count                                 rsfsu_getSweepCount

Sweep Time                                      rsfsu_confSweepTime

Sweep Time Auto                                 rsfsu_confSweepTimeAuto

Sweep Points                                    rsfsu_confSweepPoints

Trigger Group

Trigger                                         rsfsu_confTrg

Trigger Delay                                   rsfsu_confTrgDelay

Trigger Delay Compensation                      rsfsu_confTrgDelayComp

Gate Settings                                   rsfsu_confSAMGateSet

Bandwidth Group

Resolution BW                                   rsfsu_confResbw

Video BW                                        rsfsu_confSAMVibw

Set Video BW Type                               rsfsu_setVideoBWType

Analog Demodulation BW                          rsfsu_setADEMBandwidth

Coupling Settings                               rsfsu_confCoupExt

Display Group

Set Active Window                               rsfsu_confSetActiveWindow

Display Mode                                    rsfsu_confDisplayMode

Screen Coupling                                 rsfsu_confScreenCpl

Display Configuration

Display Control                              rsfsu_confDisplayControl

Display Comment                              rsfsu_confDisplayComment

Config Display                               rsfsu_confDisplayParams

Display Size                                 rsfsu_confDisplaySize

System Setup Group

Reference Oscillator                            rsfsu_confReferenceOsc

Time And Date                                   rsfsu_confSystemTime

Generate Transducer Factor                      rsfsu_confTransducerFactor

Setup Transducer                                rsfsu_confTransducer

Setup Transducer Ref Level Adj                  rsfsu_confTransducerRefLevAdj

Setup Transducer Def                            rsfsu_confTransducerDef

Setup Extern Gain Correct                       rsfsu_confExternGainCorrect

Emulation                                       rsfsu_confEmulation

Hardcopy Group

Hardcopy Options                                rsfsu_confHCOpt

Hardcopy Comment                                rsfsu_confHCComment

Hardcopy Device                                 rsfsu_confNTHCDeviceExt

Config Hardcopy Colors                          rsfsu_configHardcopyColors

Get Printer List                                rsfsu_confGetPrinterList

Set Printer                                     rsfsu_confSetPrinter

List Power Group

List Power Sequence                             rsfsu_confListPwrSeq

Meas List Power Seq                             rsfsu_actMeasListPwrSeq

List Power Set                                  rsfsu_confListPwrSet

List Power State                                rsfsu_confListPwrState

Burst Power Group

Burst Power Sequence                            rsfsu_confBurstPwrSeq

Meas Burst Power Seq                            rsfsu_actMeasBurstPwrSeq

Signal Analysis Mode

Marker Group

Noise Measurement                               rsfsu_confSAMMarkNoise

Marker Search N dB                              rsfsu_confSAMMarkSearchNdB

Frequency Counter                               rsfsu_confSAMMarkFreqCnt

Delta Marker Fixed Ref                          rsfsu_confSAMDeltaMarkFixRef

Delta Marker Fixed Ref Point                    rsfsu_confSAMDeltaMarkFixRefPoint

Delta Marker Fixed Ref Max                      rsfsu_confSAMDeltaMarkFixRefMax

Phase Noise Measurement                         rsfsu_confSAMMarkPhaseNoise

Marker Search Parameter                         rsfsu_confSAMMarkSearchParamExt

Marker Search Limits                            rsfsu_confSAMMarkSearchLimits

Marker Demodulation                             rsfsu_confSAMMarkDemod

Measure Group

Channel Power / ACP

Channel Power Channels                       rsfsu_confSAMMarkChPowChanChannels

Channel Power Standard                       rsfsu_confSAMMarkChPowChanStandard

Channel Power High Speed                     rsfsu_confSAMMarkChPowHiSpeed

Channel Power Channel Opt                    rsfsu_confSAMMarkChPowChanOpt

Channel Power Channel Set                    rsfsu_confSAMMarkChPowChanSet

Channel Power Meas Mode                      rsfsu_confSAMMarkChPowMeas

Channel Power Meas Limit                     rsfsu_confSAMMarkChPowLimExt

Channel Power Trace                          rsfsu_confSAMPowerTrace

Channel Power Noise Corr                     rsfsu_confSAMNoiseCorr

Channel Power Channel Spacing                rsfsu_confSAMChannelSpacing

Channel Power Carrier Count                  rsfsu_confSAMCarrierCount

Channel Power Reference Auto                 rsfsu_confSAMReferenceAuto

Channel Power Reference Man                  rsfsu_confSAMReferenceMan

Occupied BW Measurement                         rsfsu_confSAMMarkChPowOccPowBW

Signal Statistics

Signal Statistics                            rsfsu_confSAMSigStat

Signal Statistics Scaling                    rsfsu_confSAMSigStatScal

AM Mod Depth Measurement                        rsfsu_confSAMMarkModDepth

Third Order Intercept                           rsfsu_confTOI

Tracking Generator Mode

Tracking Generator Src Values                      rsfsu_confTrackSrc

Tracking Generator Ref Level                       rsfsu_confTrackRefLevel

Tracking Generator Ref Position                    rsfsu_confTrackRefPosition

Tracking Generator Modulation                      rsfsu_confTrackModulation

Tracking Generator Ext Src                         rsfsu_confTrackExtSrc

Tracking Generator Ext Select                      rsfsu_confTrackExtSel

Tracking Generator Ext Sweep                       rsfsu_confTrackExtFreqSweep

Tracking Generator Ext Power                       rsfsu_confTrackExtPower

Tracking Generator Ext Src Ref                     rsfsu_confTrackExtSrcRef

Vector Signal Analysis Mode

Vector Signal Analysis Mode                        rsfsu_confVSAMDemodMode

Vector Signal Analysis Display                     rsfsu_confVSAMDisp

Config Group - Digital Demod

Digital Demod Filter                            rsfsu_confVSAMDigiDemodFilter

Digital Meas Search Time                        rsfsu_confVSAMDigiSearchTime

Config Group - Analog Demod

Analog Demodulation                             rsfsu_confVSAMADemod

Analog Demodulation Parameters                  rsfsu_confVSAMADemodParam

Analog Demodulation Meas Time                   rsfsu_confVSAMADemodMeasTime

Analog Demodulation Set                         rsfsu_confVSAMADemodSet

Analog Demodulation Type                        rsfsu_confVSAMADemodType

Analog Demod RF Param                           rsfsu_confVSAMADemodRFParam

Analog Demodulation Demod BW                    rsfsu_confVSAMDemodBW

Analog Demodulation Filter                      rsfsu_confVSAMDemodFilt

Analog Demodulation Zoom                        rsfsu_confVSAMDemodZoom

PM Modulation Add-on

Analog Demod Zero Phase Ref Point            rsfsu_confVSAMADemodZeroPhase

Analog Demod Phase Wrap                      rsfsu_confVSAMADemodPhaseWrap

Analog Demod PM Units                        rsfsu_confVSAMADemodPMUnits

FM Analog Demodulation

FM Demodulation                              rsfsu_confFMDemod

FM Demodulation Output Filter                rsfsu_confFMDemodFilt

FM Demodulation LowPass Filter Auto          rsfsu_confFMDemodLPFiltAuto

FM Demodulation Output Range                 rsfsu_confFMDemodRange

FM Demodulation Output Range Auto            rsfsu_confFMDemodRangeAuto

Sweep Group

Digital Trigger                                 rsfsu_confVSAMDigiTrg

Digital Sequence Offset                         rsfsu_confVSAMDigiSeqOffset

Digital Sequence Pulse Offset                   rsfsu_confVSAMDigiSeqPulseOffset

GSM / EDGE MS/BTS Analysis Mode

MS Modulation Type                                 rsfsu_confMSModType

MS Set User Midamble                               rsfsu_confMSUserMid

MS Restore Limit Lines                             rsfsu_actMSRestore

GSM Samples Per Symbol                             rsfsu_confGSMSamplesPerSymbol

Configure Slots In Mobile                          rsfsu_confSlotsMulti

GSM Sync Search                                    rsfsu_confGSMSyncSearch

GSM Burst Search                                   rsfsu_confGSMBurstSearch

GSM Burst Search Threshold                         rsfsu_confGSMBurstSearchThreshold

GSM Burst Time Meas High Resolution                rsfsu_confGSMBurstTimeMeasHighResolution

GSM Burst Zoom Transition Number                   rsfsu_confGSMBurstZoomTransitionNumber

GSM Burst Meas Filter                              rsfsu_confGSMBurstMeasFilter

GSM Multi Carrier Mode State                       rsfsu_confGSMMultiCarrierModeState

Cdma2000 / 1xEV-DO / 3GPP WCDMA / TD-SCDMA MS/BTS

Configure WCDPower Measurement                     rsfsu_confWCDPMeas

Configure WCDPower MS Meas                         rsfsu_confWCDPMSMeas

WCDP Measurement Mode                              rsfsu_confWCDPMeasMode

Configure C2k Measurement                          rsfsu_confC2kCDPMeas

Configure C2k Band Class                           rsfsu_confC2kBandClass

Configure CDP Measurement                          rsfsu_confCDPMeas

CDP Measurement Mode                               rsfsu_confC2kCDPMeasMode

CDP Measurement Setting

CDP Set to Standard                             rsfsu_confCDPSetStd

CDP Standard                                    rsfsu_confCDPStandard

CDP Inactive Channel Treshold                   rsfsu_confCDPICT

CDP Side Band                                   rsfsu_confCDPSBand

CDP RRC Filter                                  rsfsu_confCDPRrcFilter

CDP Eliminate Tail Chips                        rsfsu_confCDPEliminateTailChips

CDP Spreading Factor                            rsfsu_confCDPSFactor

CDP Code Number                                 rsfsu_confCDPCodeNum

CDP Slot                                        rsfsu_confCDPSlot

CDP Slot Difference                             rsfsu_confCDPSlotDifference

CDP Slot Sets Count                             rsfsu_confCDPSlotSetsCount

CDP Slot Set To Analyze                         rsfsu_actCDPSlotSetToAnalyze

CDP Averaging                                   rsfsu_confCDPAveraging

CDP Subframes                                   rsfsu_confCDPSubframes

CDP RF Slot                                     rsfsu_confCDPRFSlot

CDP Scrambling Code                             rsfsu_confCDPLCode

CDP Long Code Mask                              rsfsu_confCDPLCodeMask

CDP Long Code Offset                            rsfsu_confCDPLCodeOffset

CDP Long Code Mode                              rsfsu_confCDPLCodeMode

CDP Scrambling Type                             rsfsu_confCDPLType

CDP Normalize                                   rsfsu_confCDPNormalize

CDP Q Invert                                    rsfsu_confCDPQInvert

CDP Analysis Base                               rsfsu_confCDPAnalysisBase

CDP Overview Display                            rsfsu_confCDPOverviewDisplay

CDP Preference                                  rsfsu_confCDPPref

CDP Sync Type                                   rsfsu_confCDPSyncType

CDP Channel Type                                rsfsu_confCDPChannelType

CDP Signal Mapping                              rsfsu_confCDPSigMap

CDP Signal Mapping Mode                         rsfsu_confCDPSigMapMode

CDP Constellation Parameter B                   rsfsu_confCDPConstellationParameterB

CDP Operation Mode                              rsfsu_confCDPOperationMode

CDP Antenna Type                                rsfsu_confCDPAntennaType

CDP Power Control                               rsfsu_confCDPPControl

CDP Frame to Analyze                            rsfsu_confCDPFrameAnalyze

CDP PN Offset                                   rsfsu_confCDPPNOffset

CDP IQ Length                                   rsfsu_confCDPIQLength

CDP Order                                       rsfsu_confCDPOrder

CDP Timing And Phase Offset                     rsfsu_confCDPTPM

CDP Switching Point                             rsfsu_confCDPSwitchingPoint

CDP Midamble Shift                              rsfsu_confCDPMidambleShift

CDP Channel Table (MS/BTS)

CDP Channel Table Order                         rsfsu_confCDPChTableOrder

CDP Channel Table                               rsfsu_confCDPChTable

CDP Channel Table File                          rsfsu_confCDPChTableFile

CDP Channel Table Name                          rsfsu_confCDPChTableName

CDP Channel Table Copy                          rsfsu_confCDPChTableCopy

CDP Channel Table Delete                        rsfsu_confCDPChTableDelete

CDP Channel Table Comment                       rsfsu_confCDPChTableComment

CDP Channel Table Data                          rsfsu_confCDPChTableData

CDP Channel Table Catalog                       rsfsu_confCDPChTableCatalog

CDP Channel Table Midamble Shift                rsfsu_confCDPChTableMidambleShift

CDP C2k Channel Table (MS/BTS)

CDP C2k Channel Table                           rsfsu_confC2kCDPChTable

CDP C2k Channel Table File                      rsfsu_confC2kCDPChTableFile

CDP C2k Channel Table Name                      rsfsu_confC2kCDPChTableName

CDP C2k Channel Table Copy                      rsfsu_confC2kCDPChTableCopy

CDP C2k Channel Table Delete                    rsfsu_confC2kCDPChTableDelete

CDP C2k Channel Table Restore                   rsfsu_confC2kCDPChTableRestore

CDP C2k Channel Table Comment                   rsfsu_confC2kCDPChTableComment

CDP C2k Channel Table Data                      rsfsu_confC2kCDPChTableData

CDP C2k Channel Table Catalog                   rsfsu_confC2kCDPChTableCatalog

WCDP Channel Table (MS)

WCDP MS Channel Table                           rsfsu_confWCDPMSChTable

WCDP MS Channel Table File                      rsfsu_confWCDPMSChTableFile

WCDP MS Channel Table Name                      rsfsu_confWCDPMSChTableName

WCDP MS Channel Table Copy                      rsfsu_confWCDPMSChTableCopy

WCDP MS Channel Table Delete                    rsfsu_confWCDPMSChTableDelete

WCDP MS Channel Table Comment                   rsfsu_confWCDPMSChTableComment

WCDP MS Channel Table Data                      rsfsu_confWCDPMSChTableData

WCDP MS Channel Table Catalog                   rsfsu_confWCDPMSChTableCatalog

WCDP MS Channel Evaluation                      rsfsu_confWCDPMSChEvaluation

WCDP MS Channel HS-DPCCH                        rsfsu_confWCDPMSChHSDPCCH

WCDP Channel Table (BTS)

WCDP Channel Table                              rsfsu_confWCDPChTable

WCDP Channel Table File                         rsfsu_confWCDPChTableFile

WCDP Channel Table Name                         rsfsu_confWCDPChTableName

WCDP Channel Table Copy                         rsfsu_confWCDPChTableCopy

WCDP Channel Table Delete                       rsfsu_confWCDPChTableDelete

WCDP Channel Table Comment                      rsfsu_confWCDPChTableComment

WCDP Channel Table Data                         rsfsu_confWCDPChTableData

WCDP Channel Table Catalog                      rsfsu_confWCDPChTableCatalog

Bluetooth Analysis Mode

BTooth Channel                                     rsfsu_confBToothChannel

BTooth Packet Type                                 rsfsu_confBToothPacketType

BTooth Num Adjacent Channels                       rsfsu_confBToothAdjacentChannels

BTooth Samples Per Symbol                          rsfsu_confBTSamplesPerSymbol

BTooth Output Power Class                          rsfsu_confBToothOutpPowerClass

BTooth Average Power                               rsfsu_confBToothAveragePower

Set Sync LAP                                       rsfsu_confVSAMDigiSetSyncLap

Bandwidth Group

BTooth Resolution BW                            rsfsu_confBToothResBW

BTooth Video BW                                 rsfsu_confBToothViBW

Trace Group

BTooth Trace Setup                              rsfsu_confBToothTrace

BTooth Trace Select                             rsfsu_confBToothTraceSelect

BTooth Detector                                 rsfsu_confBToothDetector

Sweep Group

BTooth Sweep                                    rsfsu_confBToothSweep

BTooth Sweep Time                               rsfsu_confBToothSweepTime

Lines Group

SEM Limit Line                                     rsfsu_confSEMLimitLine

PVT Limit Line                                     rsfsu_confPVTLimitLine

Power Meter

PWR Meter Frequency                                rsfsu_confPMetFrequency

PWR Meter Frequency Coupling                       rsfsu_confPMetFrequencyCoupling

PWR Meter Units                                    rsfsu_confPMetUnits

PWR Meter Meas Time                                rsfsu_confPMetMeasTime

PWR Meter Result Display                           rsfsu_confPMetResultDisplay

Noise Measurement

Frequency Settings

Noise Frequency Measurement                     rsfsu_confNoiseFrequencyMeasurement

Noise Frequency                                 rsfsu_confNoiseFrequency

Noise Start Frequency                           rsfsu_confNoiseStartFrequency

Noise Stop Frequency                            rsfsu_confNoiseStopFrequency

Noise Step Frequency                            rsfsu_confNoiseStepFrequency

Noise Frequency Table                           rsfsu_confNoiseFrequencyTable

Noise Fixed IF Frequency                        rsfsu_confNoiseFixedIFFrequency

Noise LO Frequency                              rsfsu_confNoiseLOFrequency

Noise Image Rejection                           rsfsu_confNoiseImageRejection

Noise DUT Type                                  rsfsu_confNoiseDUTType

Measurement Settings

Calibration

Noise 2nd Stage Correction                   rsfsu_confNoise2ndStageCorrection

Noise 2nd Stage Correction State             rsfsu_confNoise2ndStageCorrectionState

Analyzer Settings

Noise Resolution Bandwidth                   rsfsu_confNoiseRBW

Noise Sweep Time                             rsfsu_confNoiseSweepTime

Noise DUT Settling Time                      rsfsu_confNoiseDUTSettlingTime

Noise DUT Range                              rsfsu_confNoiseDUTRange

Noise Average                                rsfsu_confNoiseAverage

Noise RF Attenuation                         rsfsu_confNoiseRFAttenuation

Noise Ref Level                              rsfsu_confNoiseRefLevel

Noise Pre-selector                           rsfsu_confNoisePreselector

Noise Pre-amplifier                          rsfsu_confNoisePreamplifier

Generator Settings

Noise Generator Automatic Control            rsfsu_confNoiseGeneratorAuto

Noise Generator Settings                     rsfsu_confNoiseGeneratorSettings

Noise Generator Level                        rsfsu_confNoiseGeneratorLevel

Noise Generator Frequency                    rsfsu_confNoiseGeneratorFrequency

ENR Settings

Noise ENR Settings                              rsfsu_confNoiseENRSettings

Noise ENR Table                                 rsfsu_confNoiseENRTable

Loss Settings

Noise Loss Input Settings                       rsfsu_confNoiseLossInputSettings

Noise Loss Input Table                          rsfsu_confNoiseLossInputTable

Noise Loss Output Settings                      rsfsu_confNoiseLossOutputSettings

Noise Loss Output Table                         rsfsu_confNoiseLossOutputTable

Graphic Settings

Noise Trace Display                             rsfsu_confNoiseTraceDisplay

Noise Trace Settings                            rsfsu_confNoiseTraceSettings

Noise Gain Trace Settings                       rsfsu_confNoiseGainTraceSettings

Spurious Emissions

SE Resolution Bandwidth                            rsfsu_confSEResolutionBW

SE Video Bandwidth                                 rsfsu_confSEVideoBW

SE Break Sweep                                     rsfsu_confSEBreakSweep

SE Detector                                        rsfsu_confSEDetector

SE Filter                                          rsfsu_confSEFilter

SE Start And Stop Freq                             rsfsu_confSEStartStopFrq

SE Attenuator                                      rsfsu_confSEAtt

SE Attenuator Auto                                 rsfsu_confSEAttAuto

SE Pre-amplifier                                   rsfsu_confSEPreamplifier

SE Sweep Points                                    rsfsu_confSESweepPoints

SE Ref Level                                       rsfsu_confSERefLevel

SE Sweep Mode                                      rsfsu_confSESweepMode

SE Sweep Time                                      rsfsu_confSESweepTime

SE Sweep Time Auto                                 rsfsu_confSESweepTimeAuto

SE Transducer                                      rsfsu_confSETransducer

SE Delete Range                                    rsfsu_confSEDeleteRange

SE Search Peaks                                    rsfsu_confSESearchPeaks

Action/Status Functions

General Device Settings

Marker Group

Marker Search                                   rsfsu_actMarkSearch

Delta Marker Search                             rsfsu_actDMarkSearch

Marker To                                       rsfsu_actMarkTo

Delta Marker To                                 rsfsu_actDMarkTo

Delta Marker Link                               rsfsu_actDMarkLink

Get Marker Value                                rsfsu_actMarkValue

Get Delta Marker Value                          rsfsu_actMarkDeltaValue

Get Power Of Signal Pulses                      rsfsu_getPowerOfSignalPulses

Measure Group

Get Summary Marker Value                        rsfsu_actMarkerSumValue

Set Reference                                   rsfsu_actMarkerSumRef

Peak Group

Search Peaks                                    rsfsu_searchPeaks

Get Peaks Values                                rsfsu_getPeaksValues

Get Peaks Count                                 rsfsu_getPeaksCount

Lines Group

Get Limit Check Result                          rsfsu_actLimitCheck

Limit Check Result Clear                        rsfsu_actLimitCheckClear

Trigger Group

Send Trigger                                    rsfsu_actSendTrg

Send Trigger and Wait for OPC                   rsfsu_actSendTrgWopc

Window Trigger                                  rsfsu_actSendTrgWindow

Window Trigger and Wait for OPC                 rsfsu_actSendTrgWopcWindow

Abort                                           rsfsu_actAbort

Continue Measurement                            rsfsu_actContinueMeasurement

Trace Group

Copy Trace                                      rsfsu_actSAMCopyTrace

Calibration Group

Calibration                                     rsfsu_actCalibrationExt

Calibration Correction                          rsfsu_actCalCorr

System Setup Group

Service Source                                  rsfsu_actServiceSourceExt

Service Source Cal Signal                       rsfsu_actServiceSourceCalSignal

Get Module Info                                 rsfsu_actServiceModuleInfo

Get Time And Date                               rsfsu_actGetSystemTime

File Group

File Decimal Separator                          rsfsu_confFileDecSep

Store Trace to File                             rsfsu_actSAMStoreTraceToFile

Store Spurious Emissions to File                rsfsu_actSAMStoreSEToFile

File Operations                                 rsfsu_actMemOp

File Select Items                               rsfsu_actMemSel

File Directory                                  rsfsu_actFileCat

File Directory Path                             rsfsu_actFileCatPath

Read To File From Instrument                    rsfsu_readToFile

Write From File To Instrument                   rsfsu_writeFromFile

Hardcopy Group

Initiate Hardcopy                               rsfsu_actHCopy

Initiate Hardcopy To File                       rsfsu_actHCopyToFile

List Power Group

Get List Power Result                           rsfsu_actListPwrResult

Burst Power Group

Get Burst Power Result                          rsfsu_actBurstPwrResult

Get Burst Power Result Min                      rsfsu_actBurstPwrResultMin

Signal Analysis Mode

Analyzer Mode                                      rsfsu_actSAMMode

Marker Group

Get Marker Counter Value                        rsfsu_actSAMMarkCntValue

Get N dB Down Marker Value                      rsfsu_actSAMMarkNdBDValue

Get Noise Marker Value                          rsfsu_actSAMMarkNoiseValue

Get Phase Noise Result                          rsfsu_actSAMPhaseNoiseResult

Measure Group

Channel Power / ACP

Get Channel Power Value                      rsfsu_actSAMMarkPowerValueExt

Get Occupied Bandwidth Value                 rsfsu_actSAMMarkPowerBandValue

Channel Power To Ref Value                   rsfsu_actSAMMarkChPowToRef

Adjust Channel Power Settings                rsfsu_actSAMCPSet

ACP Limit Checking                           rsfsu_actSAMACPLimitCheckState

Get ACP Limit Check                          rsfsu_actSAMACPLimitCheck

Adapt to Signal

Channel Power Start Slot                  rsfsu_actSAMChannelPowerStartSlot

Channel Power Stop Slot                   rsfsu_actSAMChannelPowerStopSlot

Channel Power Autorange                   rsfsu_actSAMChannelPowerAutorange

Channel Power Autorange Result            rsfsu_actSAMChannelPowerAutorangeResult

Channel Power Auto Adjust                 rsfsu_actSAMChannelPowerAutoAdjust

Channel Power Auto Adjust Result          rsfsu_actSAMChannelPowerAutoAdjustResult

Get Signal Statistics                           rsfsu_actSAMSigStat

Get AM Mod Depth Value                          rsfsu_actSAMMarkModDepthValue

Get Third Order Intercept                       rsfsu_actTOI

Tracking Generator Mode

Tracking Generator Mode                            rsfsu_actTrackMode

Tracking Generator Corr State                      rsfsu_actTrackCorrState

Tracking Generator Correction                      rsfsu_actTrackCorr

Vector Signal Analysis Mode

Vector Analyzer Mode                               rsfsu_actVSAMMode

Get Analog Demod Value                             rsfsu_actVSAMMarkerADemod

Get Analog Demod Result Values                     rsfsu_getADemodResultValues

Get Analog Demod Param                             rsfsu_getVSAMParam

Get FM Offset                                      rsfsu_getVSFMOffset

TV-Demodulation Mode

TV Trigger Source                                  rsfsu_actTVTriggerSrc

GSM / EDGE MS/BTS Analysis Mode

GSM EDGE Mode                                      rsfsu_actGSMMSMode

MS Set Channel                                     rsfsu_actMSChannel

MS Channel Multi Slot                              rsfsu_actMSChannelMultiSlot

MS Limit Restore                                   rsfsu_actMSLimitRestore

GSM Burst Section                                  rsfsu_actGSMBurstSection

GSM Trigger Adjust                                 rsfsu_actGSMTrgAdjust

GSM Trigger Free Run                               rsfsu_actGSMTriggerFreeRun

GSM Measurement Mode

GSM Burst Measurement Mode                         rsfsu_actGSMBurstMeasMode

GSM Spectrum Measurement Mode                      rsfsu_actGSMSpectMeasMode

GSM Spurious Measurement Mode                      rsfsu_actGSMSpurMeasMode

GSM Modulation Accuracy                            rsfsu_actGSMModAccuracy

Bluetoooth Analysis Mode

Bluetooth Mode                                     rsfsu_actBToothMode

Bluetooth Measurement Mode                         rsfsu_actBToothMeasMode

Get BTooth Output Power                            rsfsu_actVSBToothOutputPower

Get BTooth Parameter                               rsfsu_actVSBToothParameter

Get BTooth Power Control                           rsfsu_actVSBToothPowerControl

Get BTooth Power Of Channels                       rsfsu_actVSBToothPowerOfChannels

Get BTooth Exceptions                              rsfsu_actVSBToothException

Get BTooth Digital Meas                            rsfsu_actVSBToothDM

Get BTooth Percentage Dev                          rsfsu_actVSBToothPercentageDev

Get BTooth Carrier Freq Param                      rsfsu_actVSBToothCarrierFreqParam

Get BTooth Packet Length                           rsfsu_actVSBToothPacketLength

Get BTooth Packet Type                             rsfsu_actVSBToothPacketType

Cdma2000 / 1xEV-DO / 3GPP WCDMA / TD-SCDMA MS/BTS

WCDPower Mode                                      rsfsu_actWCDPMode

WCDPower MS Mode                                   rsfsu_actWCDPMSMode

C2k CDPower Mode                                   rsfsu_actC2kCDPMode

C2k CDPower MS Mode                                rsfsu_actC2kCDPMSMode

1xEV-DO CDPower Mode                               rsfsu_actEVDOCDPMode

1xEV-DO CDPower MS Mode                            rsfsu_actEVDOCDPMSMode

TD-SCDMA CDPower Mode                              rsfsu_actTDSCDMACDPMode

TD-SCDMA CDPower MS Mode                           rsfsu_actTDSCDMACDPMSMode

CDP Level Auto Adjust                              rsfsu_actCDPAutoAdj

Marker Group

CDP Marker To                                   rsfsu_actCDPMarkTo

Get WCDP Measurement                            rsfsu_actWCDPMarkMeas

Get WCDP MS Measurement                         rsfsu_actWCDPMSMarkMeas

Get C2k CDP Measurement                         rsfsu_actC2kCDPMarkMeas

Get 1xEV-DO CDP Measurement                     rsfsu_actEVDOCDPMarkMeas

Get 1xEV-DO CDP MS Measurement                  rsfsu_actEVDOCDPMSMarkMeas

Get TD-SCDMA CDP Measurement                    rsfsu_actTDSCDMACDPMarkMeas

Lines Group

SEM Limit Line Check                               rsfsu_actSEMLimitLineCheck

Power Meter

PWR Meter State                                    rsfsu_actPMetState

PWR Meter Sensor Zeroing                           rsfsu_actPMetSensorZeroing

PWR Meter Reference Value                          rsfsu_actPMetReferenceValue

Noise Measurement

Noise Measurement Mode                             rsfsu_actNoiseMeasMode

Spurious Emissions

SE Start Measurement                               rsfsu_actSEStartMeasurement

SE Start Measurement And Wait for OPC              rsfsu_actSEStartMeasurementWopc

SE Stop Measurement                                rsfsu_SEStopMeasurement

Data Functions

Read Trace Data                                       rsfsu_dataReadTrace

Write Trace Data                                      rsfsu_dataWriteTrace

Read Trace IQ Data                                    rsfsu_dataReadTraceIQ

Read Memory IQ Data                                   rsfsu_dataReadMemoryIQ

Cdma2000  / 1xEV-DO / 3GPP WCDMA / TD-SCDMA MS/BTS

Read WCDMA Trace Data                              rsfsu_dataReadTraceWCDMA

Read WCDP Trace Data                               rsfsu_dataReadTraceWCDP

Read C2k Trace Data                                rsfsu_dataReadTraceC2k

Read C2k CDP Trace Data                            rsfsu_dataReadTraceC2kCDP

Read CDP Trace Data                                rsfsu_dataReadTraceCDP

Read Multi Frame Data                              rsfsu_readMultiFrameData

Read GSM / EDGE MS/BTS Results

Read Burst Values                                  rsfsu_dataReadBurst

Read Burst MACC Values                             rsfsu_dataReadBurstMACC

Read Burst IQ Values                               rsfsu_dataReadBurstIQ

Read Burst Ptemplate Ref                           rsfsu_dataReadPtempRef

Read Spectrum Mod Ref                              rsfsu_dataReadModRef

Read Spectrum                                      rsfsu_dataReadSpectrum

Read Level Time Values                             rsfsu_dataReadLevelTime

Fetch GSM / EDGE MS/BTS Results

Fetch Burst Values                                 rsfsu_dataFetchBurst

Fetch Burst MACC Values                            rsfsu_dataFetchBurstMACC

Fetch Burst IQ Values                              rsfsu_dataFetchBurstIQ

Fetch Burst Ptemplate Ref                          rsfsu_dataFetchPtempRef

Fetch Spectrum Mod Ref                             rsfsu_dataFetchModRef

Fetch Spectrum                                     rsfsu_dataFetchSpectrum

Power Meter

Fetch PWR Meter Result                             rsfsu_dataFetchPMetResult

Read PWR Meter Result                              rsfsu_dataReadPMetResult

Noise Measurement

Fetch Noise Measurement Result (Array)             rsfsu_dataFetchNoiseMeasArray

Fetch Noise Measurement Result (Scalar)            rsfsu_dataFetchNoiseMeasScalar

Spurious Emissions Measurement

SE Measurement Results                             rsfsu_dataSEMeasurementResults

Utility Functions

Display Update                                        rsfsu_dispUpdate

Instrument Options                                    rsfsu_OPTList

Set Status Register                                   rsfsu_setStatusRegister

Get Status Register                                   rsfsu_getStatusRegister

State Checking                                        rsfsu_errorCheckState

Write To Instrument                                   rsfsu_writeInstrData

Read Instrument Data                                  rsfsu_readInstrData

Reset                                                 rsfsu_reset

Self Test                                             rsfsu_self_test

Self Test Result                                      rsfsu_self_test_result

Error Query                                           rsfsu_error_query

Error Message                                         rsfsu_error_message

Revision Query                                        rsfsu_revision_query

FSEx Compatibility Group

Compatibility Checking                                rsfsu_compatibilityChecking

Configuration Functions

Limit Lines Parameters                             rsfsu_confLimitLineParam

Detector                                           rsfsu_confSAMDetector

Trace Math                                         rsfsu_confSAMTraceMath

Coupling Settings                                  rsfsu_confCoup

NT Hardcopy Device                                 rsfsu_confNTHCDevice

Marker Search Parameter                            rsfsu_confSAMMarkSearchParam

Channel Power Meas Limit                           rsfsu_confSAMMarkChPowLim

Channel Power / ACP

Channel Power Trigger Spacing                   rsfsu_confSAMTrigSpacing

Channel Power Trigger Count                     rsfsu_confSAMTrigCount

Action/Status Functions

Calibration                                        rsfsu_actCalibration

Read GSM/EDGE/BTS Results

GSM Burst Reference Meas                        rsfsu_dataReadBurstRef

Service Source                                     rsfsu_actServiceSource

Get Marker Power Value                             rsfsu_actSAMMarkPowerValue

Misc

Read Trace IQ Data From Mem                        rsfsu_dataReadTraceIQFromMemory

Obsolete

CDP C2k PN Offset                                  rsfsu_confC2kCDPPNOffset

CDP C2k IQ Length                                  rsfsu_confC2kCDPIQLength

CDP C2k Order                                      rsfsu_confC2kCDPOrder

CDP C2k Timing And Phase Offs                      rsfsu_confC2kCDPTPM

Limit Lines Parameters                             rsfsu_confLimitLineParamExt

Set Limit Lines Offset                             rsfsu_actSetLimitLinesOffset

Close                                                    rsfsu_close

________________________________________________________________________________

R&S Spectrum Analyzer FSU (VISA)

This instrument module provides programming support for the Rohde &

Schwarz FSU Spectrum Analyzer series FSU 3, FSU 8, FSU 26, FSU 46, FSU 50

included hardware options B3, B4, B6, B9, B10, B16, B25 and firmware

options K5, K7, K8, K9, K30, K72, K73, K74, K76, K77, K82, K83, K84, K85

and K90.

This driver is also in some level interchnageable for applications based

on FSEx (FSP, ESPI, FSQ) series driver with restrictions done by

functionality of FSU.

The driver module is divided into the following functions:

Functions/Classes:

(1) Initialize:

This function initializes the instrument and sets it to a default

configuration.

(2) Application Functions: (Class)

This class contains high-level, test and measurement routines.  These

examples call other instrument driver functions to configure, start, and

read from the instrument.

(3) Configuration Functions: (Class)

This class of functions configures the instrument by setting acquisition

and system configuration parameters.

(4) Action/Status Functions: (Class)

This class of functions begins or terminates an acquisition. It also

provides functions which allow the user to determine the current status

of the instrument.

(5) Data Functions: (Class)

This class of functions transfers data to or from the instrument.

(6) Utility Functions: (Class)

This class of functions provides lower level functions to communicate

with the instrument, and change instrument parameters.

(7) FSEx Compatibility Group: (Class)

This class provides interchangeability support for the FSEx series

functions. This driver is interchnageable for  applications based on FSEx

series driver with restrictions done by functionality of FSU. Functions

which are not fully  compatible to FSEx provides compatibility checking

and returns an error if parameter is not available.

(8) Close:

This function takes the instrument offline.

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

rsfsu_actAbort

ViStatus rsfsu_actAbort (ViSession instrumentHandle);

Purpose

This function aborts a current measurement and resets the trigger system.

Remote-control command(s):

ABORt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actBToothMeasMode

ViStatus rsfsu_actBToothMeasMode (ViSession instrumentHandle,

ViInt32 mode);

Purpose

This function sets the Bluetooth measurement mode.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:MEASurement OPOWer | ACLR | MCHar | ICFTolerance |

CFDRift

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

mode

Variable Type       ViInt32

Determines the bluetooth measurement mode.

Valid Range:

0 - Output Power (Default Value)

1 - RESERVED (Power Density)

2 - RESERVED (Power Control)

3 - RESERVED (TX Output Spectrum Freq. Range )

4 - RESERVED (TX Output Spectrum 20 dB Bandwidth)

5 - Adjacent Channel Power

6 - Modulation Characteristics

7 - Initial Carrier Frequency Tolerance

8 - Carrier Frequency Drift

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actBToothMode

ViStatus rsfsu_actBToothMode (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function selects the Bluetooth analyzer mode.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

INSTrument<1|2>:SELect BTOoth

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actBurstPwrResult

ViStatus rsfsu_actBurstPwrResult (ViSession instrumentHandle,

ViInt32 window, ViInt32 no_ofResults,

ViReal64 _VI_FAR burstPowerResults[]);

Purpose

This function queries the results of a multiple burst power measurement

as configured and initiated with Get Burst Power Result

(actBurstPwrResult). The unit used for the return values is always dBm.

This function may be used to obtain measurement results in an

asynchronous way using the service request mechanism for synchronization

with the end of the measurement. If no measurement results are available,

the function will return a query error.

Remote-control command(s):

SENSe<1|2>:MPOWer:RESult:LIST?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

no_ofResults

Variable Type       ViInt32

Expected number of values to be returned from the instrument. This

value must be equal to the Burst Power results array number of items.

Valid Range: > 0

Default Value: 256

Note:

This function will read data from the instrument till instruments

buffer is empty or number of results is reached.

burstPowerResults

Variable Type       ViReal64[]

This control returns array of elements as the results of a multiple

burst power measurement as configured and initiated with Get Burst

Power Result (actBurstPwrResult).

Note:

Array size must be => then number of elements sent from the

instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (No of Results) out of range.

BFFC0004  Parameter 4 (Burst Power Results) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actBurstPwrResultMin

ViStatus rsfsu_actBurstPwrResultMin (ViSession instrumentHandle,

ViInt32 window,

ViPReal64 powerResultMin);

Purpose

This function queries the minimum power value in a multiple burst power

measurement as configured and initiated with Burst Power Sequence

(confBurstPwrSeq). The unit used for the return values is always dBm. If

no measurement result is available, the function will return a query

error.

Remote-control command(s):

SENSe<1|2>:MPOWer:RESult:MIN?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

powerResultMin

Variable Type       ViReal64 (passed by reference)

This control queries the minimum power value in a multiple burst

power measurement as configured and initiated with Burst Power

Sequence (confBurstPwrSeq).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Power Result Min) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actC2kCDPMarkMeas

ViStatus rsfsu_actC2kCDPMarkMeas (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 measurement,

ViPReal64 measurementResult);

Purpose

This function queries the measured and calculated values of the Code

Domain Power analysis. The channel results are given for the channel to

which the code selected using the command CDPower:CODe belongs.

Where types of measurements are:

Total Power

Frequency Error in Hz

Trigger to Frame

Timing Offset

Modulation Accuracy

Peak Code Domain Error

Error Vector Magnitude RMS

Error Vector Magnitude Peak

Chip Rate Error

Channel Slot Number

Symbol Rate

Channel Number

Channel Power Absolute

Channel Power Relative

IQ Offset

IQ imbalance

Note:

(1) This function is available only in conjunction with option FS-K82 or

FS-K83.

Remote-control command(s):

CALCulate<1..4>:MARKer<1..4>:FUNCtion:CDPower:RESult? PTOTal | FERRor |

TFRame | MACCuracy | PCDerror | EVMRms | EVMPeak | CERRor | SRATe |

CHANnel | CDPabsolute | CDPRelative | IQOFfset | IQIMbalance | SLOT |

PPICh | RHO | FERPpm | ACTive | SFACtor | TOFFset | POFFset

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 4

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

Note:

1) This value is ignored. Default value has to be used.

measurement

Variable Type       ViInt32

Selects value to get.

Valid Range:

0 - Total Power (Default Value)

1 - Frequency Error in Hz

2 - Trigger to Frame

3 - Modulation Accuracy

4 - Peak Code Domain Error

5 - Error Vector Magnitude RMS

6 - Error Vector Magnitude Peak

7 - Chip Rate Error

8 - Symbol Rate

9 - Channel Number

10 - Channel Power Absolute

11 - Channel Power Relative

12 - IQ Offset

13 - IQ imbalance

14 - Channel Slot Number

15 - Pilot Power

16 - RHO

17 - Frequency Error in ppm

18 - Number Of Active Channels

19 - Spreading Factor

20 - Timing Offset

21 - Phase Offset

measurementResult

Variable Type       ViReal64 (passed by reference)

CDP measurement result. Format of the results depends on the selected

measurment.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Measurement) out of range.

BFFC0005  Parameter 5 (Measurement Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actC2kCDPMode

ViStatus rsfsu_actC2kCDPMode (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function activates cdma2000 BTS operating mode.

Note:

(1) This function is available only with option FS-K82.

Remote-control command(s):

INSTrument<1|2>:SELect BC2K

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actC2kCDPMSMode

ViStatus rsfsu_actC2kCDPMSMode (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function activates cdma2000 MS operating mode.

Note:

(1) This function is available only with option FS-K83.

Remote-control command(s):

INSTrument<1|2>:SELect MC2K

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actCalCorr

ViStatus rsfsu_actCalCorr (ViSession instrumentHandle,

ViBoolean calibrationData);

Purpose

This function determines whether (ON) or not (OFF) the current

calibration data are taken into consideration.

Remote-control command(s):

CALibration:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

calibrationData

Variable Type       ViBoolean

Determines whether or not the current calibration data are taken into

consideration.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Calibration Data) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actCalibration

ViStatus rsfsu_actCalibration (ViSession instrumentHandle,

ViInt32 calibrationType,

ViInt32 timeoutms,

ViPInt32 calibrationResult);

Purpose

This function initiates the instrument's calibration routines.

Note:

(1) This function is reserved for compatibility with FSEx driver. Prefer

to use function rsfsu_actCalibrationExt.

(2) If compatibility checking is enabled, the function will return an

error if incorrect value of Calibration Type parameter is used

Remote-control command(s):

CALibration:ALL?

CALibration:BANDwidth?

CALibration:LDETector?

CALibration:SHORt?

CALibration:IQ?

CALibration:LOSuppression?

CALibration:PPEak?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

calibrationType

Variable Type       ViInt32

Determines the type of calibration routine that will be run.

FSU Compatible Value:

0 - All (Default)

Valid Range:

0 - All (Default)

1 to 6 - reserved for compatibility with FSEx

Notes:

(1) All:

This command performs a complete calibration of the analyzer.

(2) Values 1 to 6 are accepted, but will also perform a complete

calibration of the analyzer, as the FSU does not support different

calibration subsets.

timeoutms

Variable Type       ViInt32

Sets the timeout for the calibration routine to be finished and have

obtained calibration result. If the length of time required for

calibration exceeds the timeout value, then the function will return

with a timeout error and the instrument will continue with

calibration.

Valid Range: 0 ms to 600000 ms

Default Value: 180000 ms

calibrationResult

Variable Type       ViInt32 (passed by reference)

This indicator is set to '0' when the calibration routine completed

sucessfully.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Calibration Type) out of range.

BFFC0003  Parameter 3 (Timeout [ms]) out of range.

BFFC0004  Parameter 4 (Calibration Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actCalibrationExt

ViStatus rsfsu_actCalibrationExt (ViSession instrumentHandle,

ViInt32 timeoutms,

ViChar _VI_FAR calibrationResult[]);

Purpose

This function initiates the instrument's calibration routine and returns

calibration result.

Note:

If the calibration fails, this function returns an error code.

Remote-control command(s):

CALibration?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

timeoutms

Variable Type       ViInt32

Sets the timeout for the calibration routine to be finished and have

obtained calibration result. If the length of time required for

calibration exceeds the timeout value, then the function will return

with a timeout error and the instrument will continue with

calibration.

Valid Range: 0 ms to 600000 ms

Default Value: 180000 ms

calibrationResult

Variable Type       ViChar[]

This control outputs the results of the correction data determination

from the instrument (buffer size must be at least 25000 characters).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Timeout [ms]) out of range.

BFFC0003  Parameter 3 (Calibration Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actCDPAutoAdj

ViStatus rsfsu_actCDPAutoAdj (ViSession instrumentHandle,

ViInt32 time_out_ms);

Purpose

This function is used for setting automatically the RF attenuation and IF

gain to the level of the applied signal. The instrument is to be switched

to the ATTEN MANUAL mode so as to set the RF attenuation and IF gain

separately to optimum values. This mode is maintained after changing from

code-domain power measurements to the analyzer modes.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K76, FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:LEVel:ADJust

*OPC?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

time_out_ms

Variable Type       ViInt32

Specifies timeout value for auto adjust function.

Valid Range: any positive number. Units are miliseconds.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Time out (ms)) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actCDPMarkTo

ViStatus rsfsu_actCDPMarkTo (ViSession instrumentHandle, ViInt32 window,

ViInt32 markerNumber, ViInt32 mkrFunction);

Purpose

This function changes the marker position to certain channels in code

domain power measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K82, FS-K83 or FS-K85.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:CPICh | PCCPch | DPCCh | PICH |

TDPICH

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 2

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 to 1

Default Value:

1

mkrFunction

Variable Type       ViInt32

Sets the marker function.

Valid Range:

5 - CPIC (Common Pilot Channel)  (Default Value)

6 - PCCP (Channel 1)

7 - DPCC

8 - PICH

9 - TDPICH

Notes:

(0) CPIC (Common Pilot Channel):

Sets the marker to channel 0 (Common Pilot Channel). Valid only for

option FS-K72.

(1) PCCP (Channel 1):

Sets the marker to channel 1. Valid only for option FS-K72.

(2) DPCCH:

Sets the marker 1 to channel 1. Only marker 1 is permissible. Window

1: CDP absolute and relative and window 2: modulation accuracy, peak

code domain error, power versus slot, bit stream, symbol

constellation and EVM.

(3) PICH

Sets the marker to channel 0.64. Valid only for option FS-K82, FS-K83

or FS-K85.

(4) TDPICH

Sets the marker to channel 16.128. Valid only for option FS-K82.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Mkr Function) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actCDPSlotSetToAnalyze

ViStatus rsfsu_actCDPSlotSetToAnalyze (ViSession instrumentHandle,

ViInt32 set);

Purpose

This function defines the set for which the results are calculated.

Note:

(1) This function is available only in conjunction with option FS-K76,

FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

SENSe:CDPower:SET:[VALue]

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

set

Variable Type       ViInt32

Specifies the set for which the results are calculated can be

defined.

Valid Range: 0 to SET COUNT -1

Note:

(1) SET COUNT is defined with SENSe:CDPower:SET:COUNt command

(rsfsu_confCDPSlotSetsCount).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Set) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actContinueMeasurement

ViStatus rsfsu_actContinueMeasurement (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function continues a stopped measurement at the current position in

single sweep mode. The function is useful especially for trace functions

MAXHold, MINHold and AVERage if the previous results are not to be

cleared with Sweep Count > 0 or Average Count > 0 on restarting the

measurement (INIT:IMMediate resets the previous results on restarting the

measurement).

The single-sweep mode is automatically switched on. Synchronization to

the end of the indicated number of measurements can then be performed

with the command *OPC, *OPC? or *WAI. In the continuous-sweep mode,

synchronization to the sweep end is not possible since the overall

measurement "never" ends.

Remote-control command(s):

INITiate<1|2>:CONMeas

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actDMarkLink

ViStatus rsfsu_actDMarkLink (ViSession instrumentHandle, ViInt32 window,

ViInt32 deltaMarkerNumber,

ViBoolean state);

Purpose

This command links the delta marker1 to the marker1

Note:

(1) Only Delta Marker 1 can be linked

Remote-control command(s):

CALCulate<1|2>:DELTamarker<1...4>:LINK ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 2

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

deltaMarkerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

state

Variable Type       ViBoolean

Switches the link ON/OFF

Valid Range:

0 - Off (Default Value)

1 - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Delta Marker Number) out of range.

BFFC0004  Parameter 4 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actDMarkSearch

ViStatus rsfsu_actDMarkSearch (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 searchMode);

Purpose

This function is useful for min/max delta marker peak searching.

Remote-control command(s):

CALCulate<1|2|3>:DELTamarker<1..4>:MAX:PEAK | NEXT | LEFT | RIGHt

CALCulate<1|2|3>:DELTamarker<1..4>:MIN:PEAK | NEXT | LEFT | RIGHt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 3

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

searchMode

Variable Type       ViInt32

This control positions the marker to the desired position in the

trace memory.

Valid Range:

0 - Peak Maximum (Default Value)

1 - Next Maximum

2 - Left Maximum

3 - Right Maximum

4 - Peak Minimum

5 - Next Minimum

6 - Left Minimum

7 - Right Minimum

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Search Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actDMarkTo

ViStatus rsfsu_actDMarkTo (ViSession instrumentHandle, ViInt32 window,

ViInt32 markerNumber, ViInt32 trace);

Purpose

This command assigns the selected delta marker (1 to 4) to the indicated

test curve.

Remote-control command(s):

CALCulate<1|2|3>:DELTamarker<1..4>:TRACe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 3

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

trace

Variable Type       ViInt32

Selects trace number.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Trace) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actEVDOCDPMarkMeas

ViStatus rsfsu_actEVDOCDPMarkMeas (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 measurement,

ViPReal64 measurementResult);

Purpose

This function queries the measured and calculated values of the Code

Domain Power analysis. The channel type can be set by means of the

CDPower:CTYPe command, the slot number by means of the CDPower:SLOT

command and the code number by means of the CDPower:CODe command.

Note:

(1) This function is available only in conjunction with option FS-K84.

Remote-control command(s):

CALCulate<1|2>:MARKer<1>:FUNCtion:CDPower[:BTS]:RESult? RHO | MACCuracy |

PCDerror | FERRor | FERPpm | CERRor | TFRame | IQOFfset | IQIMbalance |

SRATe | CHANnel | SFACtor | TOFFset | POFFset | CDPabsolute | CDPRelative

| EVMRms | EVMPeak | RHOPilot | RHO1 | RHO2 | PPILot | PMAC | PDATa |

PPReamble | MACTive | DACTive | PLENgth | MTYPe | DMTYpe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 4

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

Note:

1) This value is ignored. Default value has to be used.

measurement

Variable Type       ViInt32

Selects value to get.

Valid Range:

Global results for all slots:

0 - Frequency error in Hz

1 - Frequency error in ppm

2 - Chip rate error in ppm

3 - Trigger to Frame

4 - RHO over all slots for the pilot area

5 - RHOoverall-1 over all slots over all chips with start of

averaging at the half-slot limit

6 - RHOoverall-2 over all slots over all chips with start of

averaging at the quarter-slot limit

Results for the selected slot:

7  - Absolute power in the PILOT channel type

8  - Absolute power in the MAC channel type

9  - Absolute power in the DATA channel type

10 - Absolute power in the PREAMBLE channel type

11 - Composite EVM in %

12 - Modulation type of the DATA channel type:

2 = QPSK

3 = 8-PSK

4 = 16-QAM

13 - Number of active MAC channels

14 - Number of active DATA channels

15 - Length of preamble in chips

16 - RHO value for the selected channel type/slot

Results for the selected channel type in the selected slot:

17 - Peak code domain error in dB

18 - IQ imbalance in %

19 - IQ offset in %

Channel results:

20 - Symbol rate in ksps

21 - Channel number

22 - Spreading factor of the channel

23 - Timing offset in s

24 - Phase offset in rad

25 - Channel power, relative, in dB

26 - Channel power, absolute, in dBm

27 - Error vector magnitude RMS in %

28 - Error vector mag. peak in %

29 - Modulation type:

0 = BPSK-I

1 = BPSK-Q

2 = QPSK

3 = 8-PSK

4 = 16-QAM

5 = 2BPSK

measurementResult

Variable Type       ViReal64 (passed by reference)

CDP measurement result. Format of the results depends on the selected

measurment.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Measurement) out of range.

BFFC0005  Parameter 5 (Measurement Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actEVDOCDPMode

ViStatus rsfsu_actEVDOCDPMode (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function activates 1xEV-DO BTS (BDO = base station 1xEV-DO)

operating mode.

Note:

(1) This function is available only with option FS-K84.

Remote-control command(s):

INSTrument<1|2>:SELect BDO

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actEVDOCDPMSMarkMeas

ViStatus rsfsu_actEVDOCDPMSMarkMeas (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViInt32 measurement,

ViPReal64 measurementResult);

Purpose

This function queries the measured and calculated values of the code

domain power analysis. The channel results are provided for the channel

to which the code selected by means of the CDPower:CODe command belongs.

Note:

(1) This function is available only in conjunction with option FS-K85.

Remote-control command(s):

CALCulate<1|2>:MARKer<1>:FUNCtion:CDPower:RESult? SLOT | PTOTal | PPICh |

PRRI | RHO | MACCuracy | PCDerror | ACTive | FERRor | FERPpm | DRPich |

RHOVerall | CERRor | TFRame | IQOFfset | IQIMbalance | SRATe | CHANnel |

SFACtor | TOFFset | POFFset | CDPabsolute | CDPRelative | EVMRms |

EVMPeak

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 4

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

Note:

1) This value is ignored. Default value has to be used.

measurement

Variable Type       ViInt32

Selects value to get.

Valid Range:

Global results of selected half slot:

0 - Half-slot number

1 - Total power in dBm

2 - Pilot power in dBm

3 - RRI power in dBm

4 - RHO

5 - Composite EVM in %

6 - Peak code domain error in dB

7 - IQ imbalance in %

8 - IQ offset in %

9 - Number of active channels

Global results of all half slots:

10 - Frequency error in Hz

11 - Frequency error in ppm

12 - Delta RRI/PICH in dB

13 - RHO overall

14 - Trigger to frame

15 - Chip rate error in ppm

Channel results:

16 - Symbol rate in ksps

17 - Channel number

18 - Spreading factor of channel

19 - Timing offset in s

20 - Phase offset in rad

21 - Channel power relative in dB

22 - Channel power absolute in dBm

23 - Error vector magnitude RMS in %

24 - Error vector magnitude peak in %

measurementResult

Variable Type       ViReal64 (passed by reference)

CDP measurement result. Format of the results depends on the selected

measurment.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Measurement) out of range.

BFFC0005  Parameter 5 (Measurement Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actEVDOCDPMSMode

ViStatus rsfsu_actEVDOCDPMSMode (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function activates 1xEV-DO MS (MDO) operating mode.

Note:

(1) This function is available only with option FS-K85.

Remote-control command(s):

INSTrument<1|2>:SELect MDO

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actFileCat

ViStatus rsfsu_actFileCat (ViSession instrumentHandle,

ViString directory, ViInt32 bufferSize,

ViChar _VI_FAR output[]);

Purpose

This function reads the indicated directory. According to DOS convention,

wild card characters can be entered in order to query eg. a list of all

file of a certain type.

Remote-control command(s):

MMEMory:CATalog?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

directory

Variable Type       ViString

This control specifies the directory in which the command will be

performed.

Valid Range: not tested

Default Value: "D:\\user\\data"

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the ViChar array you specify for the

Output parameter.

If the current value of the attribute, including the terminating NUL

byte, contains more bytes that you indicate in this parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an

ASCII NUL byte at the end of the buffer, and returns the buffer size

you must pass to get the entire value.  For example, if the value is

"123456" and the Buffer Size is 4, the function places "123" into the

buffer and returns 7.

If you pass a negative number, the function copies the value to the

buffer regardless of the number of bytes in the value.

If you pass 0, you can pass VI_NULL for the Attribute Value buffer

parameter.

output

Variable Type       ViChar[]

This indicator returns the files in the indicated directory as comma

separated strings, for example:

�SPOOL1.WMF�,�SPOOL2.WMF�,�SPOOL3.WMF�

The buffer must be of type ViChar and have at least as many bytes as

indicated in the Buffer Size parameter.

If the current value of the attribute, including the terminating NUL

byte, contains more bytes that you indicate in this parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an

ASCII NUL byte at the end of the buffer, and returns the buffer size

you must pass to get the entire value.  For example, if the value is

"123456" and the Buffer Size is 4, the function places "123" into the

buffer and returns 7.

If you specify 0 for the Buffer Size parameter, you can pass VI_NULL

for this parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Directory) out of range.

BFFC0003  Parameter 3 (Buffer Size) out of range.

BFFC0004  Parameter 4 (Output) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actFileCatPath

ViStatus rsfsu_actFileCatPath (ViSession instrumentHandle,

ViString directory, ViInt32 bufferSize,

ViChar _VI_FAR output[],

ViPInt32 returnedValues);

Purpose

This function queries the directories and files in the given path.

According to DOS convention, wild card characters can be entered in order

to query eg. a list of all file of a certain type.

Remote-control command(s):

MEMory:CATalog:LONG? <path>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

directory

Variable Type       ViString

This control specifies the directory in which the command will be

performed.

Valid Range: not tested

Default Value: "D:\\user\\data"

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the ViChar array you specify for the

Output parameter.

If the current value of the attribute, including the terminating NUL

byte, contains more bytes that you indicate in this parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an

ASCII NUL byte at the end of the buffer, and returns the buffer size

you must pass to get the entire value in Returned Values.  For

example, if the value is "123456" and the Buffer Size is 4, the

function places "123" into the buffer and returns 7 in Returned

Values.

If you specify 0 for the Buffer Size parameter, you can pass VI_NULL

for this parameter.

Valid Values: any number

Default Value: 512

output

Variable Type       ViChar[]

This indicator returns the file/directory information in the

indicated directory as comma separated strings, in format:

<used_bytes_in_this_directory>,<free_bytes_on_this_disk>,

"<file_name>,<file_type>,<filesize_in_bytes>",

"<file_name>,<file_type>,<filesize_in_bytes>", ..

with

<file_name> name of file or directory

<file_type> file type: DIR (directory), ASCii (ASCII file),

BINary (binary file) and STATe (file with device

settings)

<filesize_in_bytes> size of file, 0 for directories

The buffer must be of type ViChar and have at least as many bytes as

indicated in the Buffer Size parameter.

If the current value of the attribute, including the terminating NUL

byte, contains more bytes that you indicate in this parameter, the

function copies Buffer Size - 1 bytes into the buffer, places an

ASCII NUL byte at the end of the buffer, and returns the buffer size

you must pass to get the entire value in Returned Values.  For

example, if the value is "123456" and the Buffer Size is 4, the

function places "123" into the buffer and returns 7 in Returned

Values.

If you specify 0 for the Buffer Size parameter, you can pass VI_NULL

for this parameter.

returnedValues

Variable Type       ViInt32 (passed by reference)

Returns the size of string returned from the instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Directory) out of range.

BFFC0003  Parameter 3 (Buffer Size) out of range.

BFFC0004  Parameter 4 (Output) is null pointer.

BFFC0005  Parameter 5 (Returned Values) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actGetSystemTime

ViStatus rsfsu_actGetSystemTime (ViSession instrumentHandle,

ViBoolean timeDate, ViPInt32 hour,

ViPInt32 minute, ViPInt32 second,

ViPInt32 day, ViPInt32 month,

ViPInt32 year);

Purpose

This function gets the instrument's internal clock or calendar.

Remote-control command(s):

SYSTem:TIME?

SYSTem:DATE?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

timeDate

Variable Type       ViBoolean

This control selects TIME or DATE to get.

Valid Range:

VI_FALSE (0) - Date

VI_TRUE  (1) - Time (Default Value)

hour

Variable Type       ViInt32 (passed by reference)

Gets the clock hour.

Note:

Return -1 means not a valid value.

minute

Variable Type       ViInt32 (passed by reference)

Gets the clock minutes.

Note:

Return -1 means not a valid value.

second

Variable Type       ViInt32 (passed by reference)

Gets the clock seconds.

Note:

Return -1 means not a valid value.

day

Variable Type       ViInt32 (passed by reference)

Gets the calendar day.

Note:

Return -1 means not a valid value.

month

Variable Type       ViInt32 (passed by reference)

Gets calendar month.

Note:

Return -1 means not a valid value.

year

Variable Type       ViInt32 (passed by reference)

Gets the year for the date function.

Note:

Return -1 means not a valid value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Time/Date) out of range.

BFFC0003  Parameter 3 (Hour) is null pointer.

BFFC0004  Parameter 4 (Minute) is null pointer.

BFFC0005  Parameter 5 (Second) is null pointer.

BFFC0006  Parameter 6 (Day) is null pointer.

BFFC0007  Parameter 7 (Month) is null pointer.

BFFC0008  Parameter 8 (Year) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actGSMBurstMeasMode

ViStatus rsfsu_actGSMBurstMeasMode (ViSession instrumentHandle,

ViInt32 measurementMode);

Purpose

This function selects measuremet to be performed (carrier power,

phase/frequency error, power vs. time).

Note:

(1) This function is available only in conjunction with option GSM/EDGE

Analyzer, FS-K5.

Remote-control command(s):

CONFigure:BURSt:PFERror | POWer | PTEMplate

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measurementMode

Variable Type       ViInt32

This control selects measurement mode to be performed.

Valid Range:

0 - Phase/Freq Error (Default Value)

1 - Carrier Power

2 - Power Vs. Time

3 - GSM Modulation Accuracy

Notes:

(1) Phase/Freq Error:

This command selects measurement of the phase and frequency error of

the base station or mobile.

(2) Carrier Power:

This command selects measurement of the average carrier power of the

base station or mobile.

(3) Power Vs. Time:

This command selects measurement of the power of the base station or

mobile vs. time.

(4) GSM Modulation Accuracy

This command selects measurement of the GSM Modulation Accuracy.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Measurement Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actGSMBurstSection

ViStatus rsfsu_actGSMBurstSection (ViSession instrumentHandle,

ViInt32 burstSection);

Purpose

This function defined the burst section to be measured for Power vs Time.

Note:

This function is available only in conjunction with option FS-K5.

Remote-control command(s):

CONFigure:BURSt:PTEMplate:SELect FULL | TOP | RISing | FALLing |FRZoom

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

burstSection

Variable Type       ViInt32

This control defines the burst section to be measured.

Valid Range:

0 - Full Burst (Default Value)

1 - Burst High Resolution

2 - Rising

3 - Falling

4 - Fall/Rise Zoom

Note(s):

(1) Fall/Rise Zoom is available if more than one slot is active.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Burst Section) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actGSMModAccuracy

ViStatus rsfsu_actGSMModAccuracy (ViSession instrumentHandle);

Purpose

This function selects modulation accuracy measurement of the  mobile in

the GSM/EDGE analyzer mode and when modulation type EDGE is selected.

Note:

(1) It is available only in conjunction with option (FS-K5).

Remote-control command(s):

CONFigure:BURSt:MACCuracy:IMMediate

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actGSMMSMode

ViStatus rsfsu_actGSMMSMode (ViSession instrumentHandle);

Purpose

This function sets the GSM/EDGE operating mode.

Note:

This function is available only in conjunction with option GSM/EDGE

Analyzer, FS-K5.

Remote-control command(s):

INSTrument<1|2>:SELect MGSM

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actGSMSpectMeasMode

ViStatus rsfsu_actGSMSpectMeasMode (ViSession instrumentHandle,

ViBoolean measurementMode);

Purpose

This function selects measurement to be performed (used to determine the

power of the spectral contributions due to modulation switching spectrum,

transient spectrum).

Note:

(1) This function is available only in conjunction with option GSM/EDGE

Analyzer, FS-K5.

Remote-control command(s):

CONFigure:SPECtrum:MODulation | SWITching

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measurementMode

Variable Type       ViBoolean

This control selects measurement mode to be performed.

Valid Range:

VI_FALSE (0) - Modulation (Default Value)

VI_TRUE  (1) - Switching

Notes:

(1) Modulation:

This command selects measurement of the spectrum due to modulation.

(2) Switching:

This command selects measurement of the spectrum due to switching

transients.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Measurement Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actGSMSpurMeasMode

ViStatus rsfsu_actGSMSpurMeasMode (ViSession instrumentHandle);

Purpose

This function selects measurement used for measuring the power of

spurious emissions.

Note:

(1) This function is available only in conjunction with option GSM/EDGE

Analyzer, FS-K5.

Remote-control command(s):

CONFigure:SPURious

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actGSMTrgAdjust

ViStatus rsfsu_actGSMTrgAdjust (ViSession instrumentHandle,

ViInt32 window, ViInt32 triggerAdjust,

ViReal64 adjustValue);

Purpose

This function performes two operations.

1) The determined GSM trigger is selected. For all GSM measurements for

which determined trigger is possible, the  trigger setting is used. If

determined trigger is not possible, the IMMediate trigger setting is used

(see table of triggers in FS-K5 manual)

2) The correction value for the time offset of the trigger from the

beginning of the first active slot is defined (GSM trigger offset, see

section "Trigger and Time References" in Chapter 1).

This correction value is needed in order to establish an exact time

reference between the trigger event and the beginning of the slot if

there is no midamble triggering. The correction value for the DUT in

question can be determined in the GENERAL SETTINGS menu with the TRIGGER

OFFSET softkey.

Note:

(1) This function is available only with installed options GSM/ESGE

Analyzer - FS-K5 and FS-B6 (TV and RF Trigger) .

Remote-control command(s):

TRIGger<1|2>:SEQuence:SYNC:ADJust:RFPower

TRIGger<1|2>:SEQuence:SYNC:ADJust:EXTernal

TRIGger<1|2>:SEQuence:SYNC:ADJust:IFPower

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

triggerAdjust

Variable Type       ViInt32

Choose desired operation.

Valid Range:

0 - External GSM Trigger (Default Value)

1 - IF Power GSM Trigger

2 - RESERVED (RF Power GSM Trigger)

Notes:

(1) External GSM Trigger:

This command is a combination of 2 commands:

For one, the �Extern� GSM trigger is selected. For all GSM

measurements requiring a trigger signal and for which an external

trigger is possible, the EXTernal trigger setting is used. If an

external trigger is not possible, the IMMediate trigger setting is

used (see manual of option).

For another, the correction value for the time offset of the external

trigger from the beginning of the measured slot is defined (see

manual of option). This correction value is needed in order to

establish an exact time reference between the trigger event and the

beginning of the slot if there is no midamble triggering. The

correction value for the DUT in question can be determined in the

GENERAL SETTINGS menu with the TRIGGER OFFSET softkey.

If the instrument is set to external trigger on switchover to

GSM/EDGE mode, the 'Extern' GSM trigger is used, otherwise the 'IF

Power' GSM trigger (default) is set.

When switching from GSM/EDGE mode to Analyzer mode, the GSM/EDGE

trigger setting is maintained, i.e. IF power if 'IF Power' was set

before and external trigger, if 'Extern' was set before.

(2) IF Power GSM Trigger:

This command is a combination of 2 commands:

For one, the �IF Power� GSM trigger is selected. For all GSM

measurements for which an IF power trigger is possible, the IFPower

trigger setting is used. If an IF power trigger is not possible, the

IMMediate trigger setting is used (see manual of option).

For another, the correction value for the time offset of the IF power

trigger from the beginning of the measured slot is defined (see

manual of option). This correction value is needed in order to

establish an exact time reference between the trigger event and the

beginning of the slot if there is no midamble triggering. The

correction value for the DUT in question can be determined in the

GENERAL SETTINGS menu with the TRIGGER OFFSET softkey.

If the instrument is set to external trigger on switchover to

GSM/EDGE mode, the �Extern� GSM trigger is used, otherwise the �IF

Power� GSM trigger (default) is set.

adjustValue

Variable Type       ViReal64

This control sets selected adjust value (See also Trigger Adjust

help).

Valid Range: -460.0e-6 to 100.0 s

Default Value: 0.0 s

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Trigger Adjust) out of range.

BFFC0004  Parameter 4 (Adjust Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actGSMTriggerFreeRun

ViStatus rsfsu_actGSMTriggerFreeRun (ViSession instrumentHandle);

Purpose

This function selects the GSM Trigger Free Run.

Remote-control command(s):

TRIGger[:SEQuence]:SYNChronize:ADJust:IMMediate

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actHCopy

ViStatus rsfsu_actHCopy (ViSession instrumentHandle,

ViInt32 deviceNumber);

Purpose

This function starts a hardcopy output for selected device.

Remote-control command(s):

HCOPy:IMMediate

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

deviceNumber

Variable Type       ViInt32

Selects hardcopy device.

Valid Range:

1 - 1 (Default Value)

2 - 2

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Device Number) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actHCopyToFile

ViStatus rsfsu_actHCopyToFile (ViSession instrumentHandle,

ViInt32 deviceNumber,

ViString destination, ViInt32 dataFormat,

ViBoolean pageOrientation);

Purpose

This function starts stores a hardcopy output for selected device to file

on local disk.

Remote-control command(s):

MMEMory:MSIS

MMEMory:CDIRectory

MMEMory:NAME

HCOPy:IMMediate

MMEMory:DATA? <file_name>

HCOPy:DEVice:LANGuage<1|2> WMF | GDI | EWMF | BMP

HCOPy:PAGE:ORIentation<1|2> PORTrait | LANDscape

HCOPy:DESTination<1|2> 'MMEM' | 'SYST:COMM:PRIN' | 'SYST:COMM:CLIP'

MMEMory:DELete

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

deviceNumber

Variable Type       ViInt32

Selects hardcopy device.

Valid Range:

1 - 1 (Default Value)

2 - 2

destination

Variable Type       ViString

This control specifies the destination of the operation to be

performed.

Valid Range: not tested

Default Value: ""

dataFormat

Variable Type       ViInt32

This control determines the data format of the printout.

Valid Range:

0 - WMF

1 - GDI

2 - EWMF

3 - BMP (Default Value)

Notes:

WMF           (WINDOWS Metafile Format) and

EWMF          Data formats for output files which can at a later

time be integrated in corresponding programs for

documentation purposes.

BMP           (Bitmap) Data format for output files.

GDI           (Graphics Device Interface) Default format for the

output to a printer configured under Windows. For

the output file the printer driver configured

under Windows is used and thus a printer-specific

format is generated.

pageOrientation

Variable Type       ViBoolean

The command selects the format of the output (portrait and landscape)

(hardcopy unit 1 or 2).

Valid Range:

VI_TRUE  (1) - Portrait (Default Value)

VI_FALSE (0) - Landscape

Note:

(1)  This setting will be ignored for

Destination <> 'SYST:COMM:PRIN'.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Device Number) out of range.

BFFC0003  Parameter 3 (Destination) out of range.

BFFC0004  Parameter 4 (Data Format) out of range.

BFFC0005  Parameter 5 (Page Orientation) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actLimitCheck

ViStatus rsfsu_actLimitCheck (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViPInt32 checkResult);

Purpose

This function queries the result of the limit check.

Remote-control command(s):

CALCulate<1|2>:LIMit<1..8>:FAIL?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

checkResult

Variable Type       ViInt32 (passed by reference)

This control returns the result of the limit check responds with 0 in

case of PASS and with 1 in case of FAIL.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Check Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actLimitCheckClear

ViStatus rsfsu_actLimitCheckClear (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo);

Purpose

This function deletes the result of the current limit check for all limit

lines.

After this command has been issued the function Get Limit Check Result

(rsfsu_actLimitCheck) will return 0 until the next measurement has been

run. This function also clears the STATus:QUEStionable:LIMit condition

and event registers.

Note this function does not effect the display of the limit check on the

display for the appropriate limit line.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:CLEar[:IMMediate]

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actListPwrResult

ViStatus rsfsu_actListPwrResult (ViSession instrumentHandle,

ViInt32 window, ViInt32 no_ofResults,

ViReal64 _VI_FAR listPowerResults[]);

Purpose

This function queries the result of a previous list measurement as

configured and initiated with List Power Sequence (confListPwrSeq). This

function may be used to obtain measurement results in an asynchronous

way, using the service request mechanism for synchronization with the end

of the measurement. If no measurement results are available, the function

will return a query error.

Remote-control command(s):

SENSe<1|2>:LIST:POWer:RESult?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

no_ofResults

Variable Type       ViInt32

Expected number of values to be returned from the instrument. This

value must be equal to the List Power results array number of items.

Valid Range: > 0

Default Value: 256

Note:

This function will read data from the instrument till instruments

buffer is empty or number of results is reached.

listPowerResults

Variable Type       ViReal64[]

This control returns array of elements as the result of a previous

list measurement as configured and initiated with List Power Sequence

(confListPwrSeq).

Note:

Array size must be => then number of elements sent from the

instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (No of Results) out of range.

BFFC0004  Parameter 4 (List Power Results) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMarkDeltaValue

ViStatus rsfsu_actMarkDeltaValue (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 deltaMarker,

ViPReal64 deltaMarkerValue,

ViPReal64 deltaMarkerValue_RESERVED);

Purpose

This function gets delta marker values from the selected marker position.

Remote-control command(s):

CALCulate<1|2|3>:DELTamarker<1..4>:X?

CALCulate<1|2|3>:DELTamarker<1..4>:X:RELative?

CALCulate<1|2|3>:DELTamarker<1..4>:Y?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 3

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

deltaMarker

Variable Type       ViInt32

This control sets delta marker value to be returned.

Valid Range:

0 - X value ABS (Default Value)

1 - X value REL

2 - Y value

Notes:

(1) X value ABS:

This query always returns the absolute value of frequency or time.

(2) X value REL:

This command queries the frequency (span > 0) or time (span = 0) of

the selected delta marker relative to the reference marker.

deltaMarkerValue

Variable Type       ViReal64 (passed by reference)

This control returns real part of the selected delta marker value.

Depending on the unit defined with CALC:UNIT or on the activated

measuring functions, the query result is output in the units below:

DBM | DBPW | DBUV | DBMV | DBUA:        Output unit DB

WATT | VOLT | AMPere:                   Output unit W | V | A

Statistics function (APD or CCDF) on:   Dimensionless output

Result Display FM (FS-K7):              Hz     (Y value only)

Result Display RF POWER (FS-K7):        dB     (Y value only)

Result Display SPECTRUM (FS-K7):        dB     (Y value only)

Result Display OUTPUT POWER (FS-K8):    dB     (Y value only)

Result Display TX SPECTRUM ACP (FS-K8): dB     (Y value only)

Result Display MODUL. CHAR. (FS-K8):    Hz     (Y value only)

Result Display INIT. CARR FREQ (FS-K8): Hz     (Y value only)

Result Display CARR.FREQ DRIFT (FS-K8): Hz     (Y value only)

deltaMarkerValue_RESERVED

Variable Type       ViReal64 (passed by reference)

This control returns imaginary part of the selected delta marker

value.

Note:

(1) This control is reserved for compatibility with FSEx driver  and

does not apply to the instrument.

(2) To avoid compatibility checking pass VI_NULL as a value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Delta Marker) out of range.

BFFC0005  Parameter 5 (Delta Marker Value) is null pointer.

BFFC0006  Parameter 6 (Delta Marker Value (RESER) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMarkerSumRef

ViStatus rsfsu_actMarkerSumRef (ViSession instrumentHandle,

ViInt32 window);

Purpose

With this function the power values currently measured by means of the

calculation of the average (..:SUMMary:MEAN) and of the rms value

(..:SUMMary:RMS) are declared as reference values for

relative measurements in the indicated measurement window.

Notes:

(1) It is only available in the time domain (span = 0).

(2) If the measurement of rms value and average is not activated, the

reference value 0 dBm is used.

(3) If the function ...:SUMMary:AVERage or ...:SUMMary:PHOLd is switched

on, the current value is the measured value summed at the time

considered.

Remote-control command(s):

CALCulate<1|2>:MARKer:FUNCtion:SUMMary:REFerence:AUTO ONCE

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 2

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMarkerSumValue

ViStatus rsfsu_actMarkerSumValue (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 summaryMarker,

ViPReal64 measuredValue);

Purpose

This function queries the results of the time domain power measurement.

Remote-control command(s):

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:MAXimum:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:PPEak:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:MPEak:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:MIDDle:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:RMS:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:MEAN:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:SDEViation:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:PPEak:AVERage:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:PPEak:PHOLd:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:RMS:AVERage:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:RMS:PHOLd:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:MEAN:AVERage:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:MEAN:PHOLd:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:SDEViation:AVER:RESult?

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:SUMmary:SDEViation:SDEViation:PHOL

d:RESult?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 3

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

summaryMarker

Variable Type       ViInt32

This control selects value to be measured.

Valid Range:

0 - RESERVED

1 - Positive Peak (Default Value)

2 - RESERVED

3 - RESERVED

4 - RMS

5 - Mean

6 - Standard Deviation

7 - Positive Peak Average

8 - Positive Peak Hold

9 - RMS Average

10 - RMS Peak Hold

11 - Mean Average

12 - Mean Peak Hold

13 - Standard Deviation Average

14 - Standard Deviation Peak Hold

Notes:

(1) Positive Peak:

This command is used to query the result of the measurement of the

positive peak value in the selected measurement window. The

measurement may have to be switched on previously. It is only

available in the time domain (span = 0). A complete sweep with

synchronization to sweep end should be performed between switching on

the function and querying the measured value to obtain a valid query

result. This is only possible in the

single-sweep mode.

(2) RMS:

This command queries the result of the measurement of the rms value

of the total trace in the selected measurement window. It is only

available in the time domain (span = 0). A complete sweep with

synchronization to sweep end should be performed between switching on

the function and querying the measured value to obtain a valid query

result. This is only possible in the single-sweep mode.

(3) Mean:

This command queries the result of the measurement of the mean value

of the total trace in the selected measurement window. It is only

available in the time domain (span = 0). A complete sweep with

synchronization to sweep end should be performed between switching on

the function and querying the measured value to obtain a valid query

result. This is only possible in the single-sweep mode.

(4) Standard Deviation:

This command queries the results of the standard deviation

measurement. It is only available in the time domain (span = 0).

A complete sweep with synchronization to sweep end should be

performed between switching on the function and querying the measured

value to obtain a valid query result. This is only possible in the

single-sweep mode.

(5) Positive Peak Average:

This command is used to query the result of the measurement of the

positive peak value in the selected measurement window. It is only

available in the time domain (span = 0). A complete sweep with

synchronization to sweep end should be performed between switching on

the function and querying the measured value to obtain a valid query

result. This is only possible in the single-sweep mode.

(6) Positive Peak Peak Hold:

This command is used to query the result of the measurement of the

positive peak value with the peak hold function switched on.  It is

only available in the time domain (span = 0). A complete sweep with

synchronization to sweep end should be performed between switching on

the function and querying the measured value to obtain a valid query

result. This is only possible in the single-sweep mode.

(7) RMS Average:

This command queries the result of the measurement of the rms value

of the total trace in the selected measurement window with the

average calculated. It is only available in the time domain (span =

0). A complete sweep with synchronization to sweep end should be

performed between switching on the function and querying the measured

value to obtain a valid query result. This is only possible in the

single-sweep mode.

(8) RMS Peak Hold:

This command queries the result of the measurement of the rms value

of the total trace in the selected measurement window with the peak

hold function switched on. It is only available in the time domain

(span = 0). A complete sweep with synchronization to sweep end should

be performed between switching on the function and querying the

measured value to obtain a valid query result. This is only possible

in the single-sweep mode.

(9) Mean Average:

This command queries the result of the measurement of the mean value

of the total trace in the selected measurement window with the

average calculated. It is only available in the time domain (span =

0). A complete sweep with synchronization to sweep end should be

performed between switching on the function and querying the measured

value to obtain a valid query result. This is only possible in the

single-sweep mode.

(10) Mean Peak Hold:

This command queries the result of the measurement of the mean value

of the total trace in the selected measurement window with the peak

hold function switched on. It is only available in the time domain

(span = 0). A complete sweep with synchronization to sweep end should

be performed between switching on the function and querying the

measured value to obtain a valid query result. This is only possible

in the single-sweep mode.

(11) Standard Deviation Average:

This command queries the result of the standard deviation determined

in several sweeps in the selected measurement window. The query is

possible only if the calculation of average is active. It is only

available in the time domain (span = 0). A complete sweep with

synchronization to sweep end should be performed between switching on

the function and querying the measured value to obtain a valid query

result. This is only possible in the single-sweep mode.

(12) Standard Deviation Peak Hold:

This command queries the maximum standard deviation value determined

in several sweeps in the selected measurement window. The query is

possible only if the peak hold function is active.  It is only

available in the time domain (span = 0). A complete sweep with

synchronization to sweep end should be performed between switching on

the function and querying the measured value to obtain a valid query

result. This is only possible in the single-sweep mode.

measuredValue

Variable Type       ViReal64 (passed by reference)

Returns measured value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Summary Marker) out of range.

BFFC0005  Parameter 5 (Measured Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMarkSearch

ViStatus rsfsu_actMarkSearch (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 searchMode);

Purpose

This function is useful for min/max peak searching.

Remote-control command(s):

CALCulate<1|2|3>:MARKer<1..4>:MAX:PEAK | NEXT | LEFT | RIGHt

CALCulate<1|2|3>:MARKer<1..4>:MIN:PEAK | NEXT | LEFT | RIGHt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 3

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

searchMode

Variable Type       ViInt32

This control positions the marker to the desired position in the

trace memory.

Valid Range:

0 - Peak Maximum (Default Value)

1 - Next Maximum

2 - Left Maximum

3 - Right Maximum

4 - Peak Minimum

5 - Next Minimum

6 - Left Minimum

7 - Right Minimum

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Search Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMarkTo

ViStatus rsfsu_actMarkTo (ViSession instrumentHandle, ViInt32 window,

ViInt32 markerNumber, ViInt32 mkrFunction,

ViInt32 trace);

Purpose

This function changes instrument parameters via currently active marker.

Remote-control command(s):

CALCulate<1|2|3>:MARKer<1..4>:TRACe

CALCulate<1|2|3>:MARKer<1..4>:FUNCtion:CENTer | CSTep | STARt | STOP |

MSTep | REFerence | TRACe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 3

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

mkrFunction

Variable Type       ViInt32

Sets the marker function.

Valid Range:

0 - Center (Default Value)

1 - CF Step

2 - RESERVED

3 - RESERVED

4 - RESERVED

5 - Reference

6 - Trace

Notes:

(1) Center:

This command sets the center frequency to that of the current marker.

(2) CF Step:

This command sets the step width of the center frequency to the

x-value of the current marker.

(3) Reference:

This command sets the reference level to that of the current marker.

(4) Trace:

This command assigns the selected marker (1 to 4) to the indicated

test curve.

trace

Variable Type       ViInt32

Selects trace number.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Mkr Function) out of range.

BFFC0005  Parameter 5 (Trace) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMarkValue

ViStatus rsfsu_actMarkValue (ViSession instrumentHandle, ViInt32 window,

ViInt32 markerNumber,

ViPReal64 markerValue,

ViPReal64 markerValue_RESERVED);

Purpose

This function gets the marker value from the selected marker position.

Remote-control command(s):

CALCulate<1|2|3>:MARKer<1..4>:Y?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 3

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

markerValue

Variable Type       ViReal64 (passed by reference)

This control returns real part of marker value.

The unit of returned value depends on CALC:UNIT setting.

With linear level scaling is the output in %.

If the FM Demodulator (FS-K7) is activated, the query result is

output in the following units:

Result Display FM:                      Hz

Result Display RF POWER LOG:            dBm

Result Display RF POWER LIN:            %

Result Display SPECTRUM LOG:            dBm

Result Display SPECTRUM LIN:            %

Result Display OUTPUT POWER (FS-K8):    dB

Result Display TX SPECTRUM ACP (FS-K8): dB

Result Display MODUL. CHAR. (FS-K8):    Hz

Result Display INIT. CARR FREQ (FS-K8): Hz

Result Display CARR.FREQ DRIFT (FS-K8): Hz

markerValue_RESERVED

Variable Type       ViReal64 (passed by reference)

This control returns imaginary part of marker value.

Note:

(1) This control is reserved for compatibility with FSEx driver  and

does not apply to the instrument.

(2) To avoid compatibility checking pass VI_NULL as a value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Marker Value) is null pointer.

BFFC0005  Parameter 5 (Marker Value (RESERVED)) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMeasBurstPwrSeq

ViStatus rsfsu_actMeasBurstPwrSeq (ViSession instrumentHandle,

ViInt32 window,

ViReal64 analyzerFrequencyHz,

ViReal64 resolutionBandwidthHz,

ViReal64 measTimes,

ViInt32 triggerSource,

ViReal64 triggerLevel,

ViReal64 triggerOffsets,

ViInt32 type_ofMeas,

ViInt32 no_ofMeasurements,

ViInt32 timeout,

ViReal64 _VI_FAR burstPowerResults[]);

Purpose

This function configures the instrument setup for multiple burst power

measurement, starts a measurement sequence and reads the results. To

reduce the setting time, the setup is performed simultaneously for all

selected parameters.

Remote-control command(s):

SENSe<1|2>:MPOWer:SEQuence?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

analyzerFrequencyHz

Variable Type       ViReal64

Receive frequency for the burst signals to be measured (= center

frequency in manual operation).

Valid Range: 0.0 Hz to max. analyzer frequency, depending on

instrument model

Default Value: 0.0 Hz

resolutionBandwidthHz

Variable Type       ViReal64

Resolution bandwidth for the measurement.

Valid Range: 10.0 Hz to 10.0e6 Hz in steps of 1, 3, 10

Default Value: 10.0e6 Hz

measTimes

Variable Type       ViReal64

Time span during which measurement samples are sampled for RMS / peak

measurement The type of measurement is selected by Type of Meas.

Valid Range: 1.0e-6 s to 30.0 s

Default Value: 1.0e-6

triggerSource

Variable Type       ViInt32

Trigger signal source.

Valid Range:

0 - External (Default Value)

1 - Video

Notes:

(1) External:

The trigger signal is fed from the "Ext. Trigger/Gate" input on the

rear of the unit.

(2) Video:

The internal video signal is used as trigger signal.

triggerLevel

Variable Type       ViReal64

Signal level of the video signal as a percentage of the diagram

height at which the trigger becomes active. If Trigger Source =

External is selected, the value entered here is ignored, as in this

case the trigger input uses TTL levels.

Valid Range: 0.0 % to 100.0 %

Default Value: 50.0 %

triggerOffsets

Variable Type       ViReal64

Offset between the detection of the trigger signal and the start of

the measurement.

Valid Range: 125.0e-9 s to 100.0 s

Default Value: 125.0e-6 s

type_ofMeas

Variable Type       ViInt32

Determines whether mean power (RMS) or peak power (PEAK) is to be

measured. The FSU detector is selected accordingly.

Valid Range:

0 - MEAN (Default Value)

1 - PEAK

no_ofMeasurements

Variable Type       ViInt32

Number of individual bursts to be measured.

Valid Range: 1 to 501

Default Value: 1

timeout

Variable Type       ViInt32

This control sets the timeout value for data readout in milliseconds.

Valid Range: 0 to 2^31

Default Value: 10000 ms

Note:

This control is not used.

burstPowerResults

Variable Type       ViReal64[]

This control returns array of elements as the result of a previous

list measurement as configured and initiated with List Power Sequence

(confListPwrSeq).

Note:

Array size must be => then number of elements sent from the

instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Analyzer Frequency [Hz]) out of range.

BFFC0004  Parameter 4 (Resolution Bandwidth [Hz]) out of range.

BFFC0005  Parameter 5 (Meas Time [s]) out of range.

BFFC0006  Parameter 6 (Trigger Source) out of range.

BFFC0007  Parameter 7 (Trigger Level [%]) out of range.

BFFC0008  Parameter 8 (Trigger Offset [s]) out of range.

BFFC0009  Parameter 9 (Type of Meas) out of range.

BFFC000A  Parameter 10 (No of Measurements) out of range.

BFFC000B  Parameter 11 (Timeout) out of range.

BFFC000C  Parameter 12 (Burst Power Results) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMeasListPwrSeq

ViStatus rsfsu_actMeasListPwrSeq (ViSession instrumentHandle,

ViInt32 window,

ViInt32 no_ofListItems,

ViReal64 _VI_FAR analyzerFrequencyHz[],

ViReal64 _VI_FAR referenceLeveldBm[],

ViReal64 _VI_FAR RFInputAttenuationdB[],

ViReal64 _VI_FAR RFInputElectronicAttndB[],

ViInt32 _VI_FAR filterType[],

ViReal64 _VI_FAR resolutionBandwidthHz[],

ViReal64 _VI_FAR videoBandwidthHz[],

ViReal64 _VI_FAR measTimes[],

ViReal64 _VI_FAR triggerLevel[],

ViInt32 timeout,

ViReal64 _VI_FAR listPowerResults[]);

Purpose

This function configures the list of settings (max. 100 entries for an

individual frequency point) for the multiple power measurement, starts a

measurement sequence and returns the measured data. To reduce the setting

time, all indicated parameters are set up simultaneously at each test

point.

Notes:

(1) Each one test point is defined on the same position in all parameter

arrays.

(2) The measurement is performed in the time domain and therefore the

span is set to 0 Hz. If the time domain is quit, the function is

automatically switched off. The measurement is not compatible with other

measurements, especially as far as marker, adjacent channel power

measurement or statistics are concerned. The corresponding commands thus

automatically deactivate the function. The function is only available in

REMOTE operation. It is deactivated when switching the instrument back to

LOCAL.

Remote-control command(s):

SENSe<1|2>:LIST:POWer:SEQuence?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

no_ofListItems

Variable Type       ViInt32

Set number of list items. This value must be equal to array size of

each list parameter.

Valid Range: 1 to 100

Default Value: 1

analyzerFrequencyHz

Variable Type       ViReal64[]

Array of receive frequency for the signal to be measured (= center

frequency in manual operation).

Valid Range: 0.0 Hz to max. analyzer frequency, depending on the

instrument model.

Default Value: none

referenceLeveldBm

Variable Type       ViReal64[]

Array of reference level values for each list item.

Valid Range: -70.0 dBm to 30.0 dBm in 10 dB steps

-75.0 dBm to 30.0 dBm in 5 dB steps (Electronic

Attenuator option FS-B25)

Default Value: none

RFInputAttenuationdB

Variable Type       ViReal64[]

Array of RF input attenuation values for each list item.

Valid Range: 0.0 dB to 70.0 dB in 10 dB steps

0.0 dB to 75.0 dB in 5 dB steps (Electronic

Attenuator Option FS-B25)

Default Value: none

RFInputElectronicAttndB

Variable Type       ViReal64[]

Array of RF input attenuation values of electronic attenuator for

each list item.

Valid Range: 0.0 dB to 30.0 dB in 10 dB steps

Default Value: none

Note:

If value is < 0.0 (negative value) then electronic attenuator is not

in the signal path (OFF). If option FS-B25 is missing, -1 (OFF) is to

be given.

filterType

Variable Type       ViInt32[]

Array of filter types for each list item.

Valid Range:

0 - Normal

1 - CFilter

2 - RRC

Default Value: none

Notes:

(1) Normal:

Normal resolution filter.

(2) CFilter:

Channel filter. These are especially steep-edged filters, which are

used for example in Fast ACP measurement to ensure the band-limiting

of a transmission channel in the time domain.

(3) RRC:

Root Raised Cosine filter. This special filter form is used to

determine the channel power for some mobile radio standards.

resolutionBandwidthHz

Variable Type       ViReal64[]

Array of resolution bandwidth values for each list item.

Valid Range: 10.0 Hz to 10.0e6 Hz in 1, 3, 10 steps

Default Value: 0.0 Hz

Note:

Valid range also depends on the selected filter type.

videoBandwidthHz

Variable Type       ViReal64[]

Array of video bandwidth values for each list item.

Valid Range: 1.0 Hz to 10.0e6 Hz in 1, 3, 10 steps.

Default Value: none

Note:

The value is ignored for <filter type> = CFilter or RRC.

measTimes

Variable Type       ViReal64[]

Array of measurement time values for each list item.

Valid Range: 1.0e-6 s to 30.0 s

Default Value: none

triggerLevel

Variable Type       ViReal64[]

Array of trigger level values for each list item.

This value is Reserved. Must be set to 0 for each list item.

timeout

Variable Type       ViInt32

This control sets the timeout value for data readout in milliseconds.

Valid Range: 0 to 2^31

Default Value: 10000 ms

listPowerResults

Variable Type       ViReal64[]

This control returns array of elements as the result of a previously

configured and initiated list measurement.

Notes:

(1) Array size must be => No of List Items

(2) The result list can be extended to 3 results per frequency point

if the Peak, RMS and Average detectors are activated.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (No of List Items) out of range.

BFFC0004  Parameter 4 (Analyzer Frequency [Hz]) out of range.

BFFC0005  Parameter 5 (Reference Level [dBm]) out of range.

BFFC0006  Parameter 6 (RF Input Attenuation [dB]) out of range.

BFFC0007  Parameter 7 (RF Input Electronic Attn ) out of range.

BFFC0008  Parameter 8 (Filter Type) out of range.

BFFC0009  Parameter 9 (Resolution Bandwidth [Hz]) out of range.

BFFC000A  Parameter 10 (Video Bandwidth [Hz]) out of range.

BFFC000B  Parameter 11 (Meas Time [s]) out of range.

BFFC000C  Parameter 12 (Trigger Level) out of range.

BFFC000D  Parameter 13 (Timeout) out of range.

BFFC000E  Parameter 14 (List Power Results) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMemOp

ViStatus rsfsu_actMemOp (ViSession instrumentHandle,

ViInt32 operation_to_bePerformed,

ViString source, ViString destination);

Purpose

The MMEMory (mass memory) subsystem function provides commands which

allow for access to the storage media of the instrument and for storing

and loading various instrument settings.

The File to be Printed/Plotted command stores the HCOPy outputs in a

file. The various drives can be addressed via the mass storage unit

specifier <msus> using the conventional DOS syntax. The internal hard

disk is addressed by "C:", the floppy-disk drive installed by "A:". The

file names <file_name> are indicated as string parameters with the

commands being enclosed in quotation marks. They correspond to the DOS

conventions.

DOS file names consist of max. 8 ASCII characters and an extension of up

to three characters separated from the file name by a colon "." Both, the

colon and the extension are optional. The colon is not part of the file

name. DOS file names do not differ between uppercase and lowercase

notation. All letters and digits are permitted as well as the special

characters "_", "^", "$", "~", "!", "#", "%", "&", "-", "{","}", "(",

")", "@" and "� ". Reserved file names are CLOCK$, CON, AUX, COM1 to

COM4, LPT1 to LPT3, NUL and PRN. The two characters "*" and "?" have the

function of so-called "wildcards", i.e., they are variables for

selection of several files. The question mark "?" replaces exactly one

character which may be any, the asterisk means any of the remaining

characters in the file name. "*.*" thus means all files in a directory.

Remote-control command(s):

MMEMory:CDIRectory

MMEMory:COPY

MMEMory:DELete

MMEMory:INITialize

MMEMory:LOAD:STATe

MMEMory:MDIRectory

MMEMory:MOVE

MMEMory:MSIS

MMEMory:NAME

MMEMory:RDIRectory

MMEMory:STORe:STATe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

operation_to_bePerformed

Variable Type       ViInt32

This control selects the type of memory operation to be performed.

Valid Range:

0 - Change Directory (Default Value)

1 - Copy

2 - Delete

3 - Format Disk

4 - Load

5 - Make a Directory

6 - Rename a File

7 - Change Drives

8 - File to be Printed/Plotted

9 - Delete Directory

10 - Store Instrument Settings

Notes:

(1) Change Directory:

This command changes the current directory. In addition to the path

name, the indication of the directory may contain the drive name. The

path name complies with the DOS conventions. Put the directory name

into Source control (Example �D:\\USER\\DATA�).

(2) Copy:

This command copies the files indicated. The indication of the file

name may include the path and the drive. The file names and path

information must be in accordance with the DOS conventions. Use

Source (Example �D:\\USER\\DATA\\SETUP.CFG�) and Destination (Example

�A:�) controls to define file names.

(3) Delete:

This command deletes the files indicated. The indication of the file

name contains the path and, optionally, the drive. Indication of the

path corresponds to the DOS conventions. The file name includes

indication of the path and may also include the drive. The path name

corresponds to the DOS conventions. Use Source control to define

files to be deleted (Example �TEST01.HCP�).

(4) Format Disk:

This command formats the disk in drive A. Formatting deletes all data

stored on the floppy disk. Use Source control to define disk drive

(Example �A:�).

(5) Load:

This command loads instrument settings from files. The contents of

the file is loaded and set as new instrument state. Use Source

control to define file name (Example �A:TEST�).

(6) Make a Directory:

This command creates a new directory. Put the directory name into

Source control (Example �D:\\USER\\DATA�).

(7) Rename a File:

This command renames existing files. Use Source (Example

�TEST01.CFG�) and Destination (Example �SETUP.CFG�) controls to

define file names.

(8) Change Drives:

This command changes to the drive indicated. The drive may be the

internal hard disk 'D:' or the floppy-disk drive 'A:'. The drive is

indicated according to the DOS conventions. Use Source control to

define disk drive.

(9) File to be Printed/Plotted:

This command specifies a file which is printed or plotted to. The

file name includes indication of the path and may also include the

drive. The file name and path information correspond to the DOS

conventions. The output to the printer is routed into a file using

the command "HCOP:DEST �MMEM�". Use Source control to define file

name (Example �PLOT1.HPG�).

(10) Delete Directory:

This command deletes the directory indicated. The directory name

includes indication of the path and may also include the drive. The

path name corresponds to the DOS conventions. Use Source control to

define directory name (Example �D:\\TEST�).

(11) Store Instrument Settings:

This command stores the current instrument setting in a file. The

current instrument state is stored as a file. The file name includes

indication of the path and may also include the drive. The path name

corresponds to the DOS conventions.  Use Source control to define

path and file name (Example �D:\\USER\\DATA\\SET_1�)

source

Variable Type       ViString

This control specifies the source of the operation to be performed.

If the operation requires a single parameter, this control specifies

that parameter.

Valid Range: not tested

Default Value: "D:\\user\\data"

destination

Variable Type       ViString

This control specifies the destination of the operation to be

performed.

Valid Range: not tested

Default Value: ""

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Operation to be Performed) out of range.

BFFC0003  Parameter 3 (Source) out of range.

BFFC0004  Parameter 4 (Destination) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMemSel

ViStatus rsfsu_actMemSel (ViSession instrumentHandle,

ViInt32 itemSelector, ViBoolean itemState);

Purpose

This function add/remove settings to/from the list of partial datasets of

a device setting to be stored/loaded. This function is working in

conjunction with function File Operations (rsfsu_actMemOp).

Remote-control command(s):

MMEMory:SELect:NONE

MMEMory:SELect:DEFault

MMEMory:SELect:ALL

MMEMory:SELect:GSETup  ON | OFF

MMEMory:SELect:HWSettings ON | OFF

MMEMory:SELect:TRACE  ON | OFF

MMEMory:SELect:LINes  ON | OFF

MMEMory:SELect:LINes:ALL ON | OFF

MMEMory:SELect:TRANsducer ON | OFF

MMEMory:SELect:TRANsducer:ALL ON | OFF

MMEMory:SELect:CVL ON | OFF

MMEMory:SELect:CVL:ALL ON | OFF

MMEMory:SELect:SCData ON | OFF

MMEMory:SELect:CSETup  ON | OFF

MMEMory:SELect:HCOPy ON | OFF

MMEMory:SELect:MACRos ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

itemSelector

Variable Type       ViInt32

This control selects the items to save or recall.

Valid Range:

0 - Select None

1 - Select Default (Default Value)

2 - Select All

3 - RESERVED

4 - HW Settings

5 - Available Trace

6 - Available Trace

7 - Available Trace

8 - RESERVED

9 - RESERVED

10 - All Lines

11 - Active Transducer

12 - All Transducers

itemState

Variable Type       ViBoolean

This control sets if selected item is active or not for to

save/recall.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Item Selector) out of range.

BFFC0003  Parameter 3 (Item State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMSChannel

ViStatus rsfsu_actMSChannel (ViSession instrumentHandle,

ViInt32 midamble);

Purpose

This function selects the midamble used by the mobile or base station.

Notes:

(1) This function is available only in conjunction with option GSM/EDGE

Analyzer, FS-K5.

(2) The selected midamble is only significant for GSM/EDGE measurements

modulation accuracy (MAC), phase/frequency error (PFE) and power versus

time (PVT).

Remote-control command(s):

CONFigure:CHANnel:TSC

CONFigure:CHANnel:TSC USER | AB0 | AB1 | B2

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

midamble

Variable Type       ViInt32

This control selects the midamble used by the mobile.

Valid Range:

0 - 0 (Default Value)

1 - 1

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - User Defined

9 - AB0

10 - AB1

11 - AB2

Notes:

(1) 0 to 7:

Training sequence for the Normal Burst.

(2) 8:

With selection USER, the user defined midamble set with command

CONFigure:CHANnel:TSC:USER is used.

(3) 9 to 10:

The Access Burst Mode is activated using the access burst midamble

pattern and slot structure from the GSM standard. When the Access

Burst Mode is entered the current measurement (like PVT, PFE, MOD,

etc.) is left and the modulation is set to GSM. In Access Burst Mode

the measurements PVT or PFE can be selected (other measurements are

not available).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Midamble) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMSChannelMultiSlot

ViStatus rsfsu_actMSChannelMultiSlot (ViSession instrumentHandle,

ViInt32 slots);

Purpose

This function defines the used slots of the mobile.

The multislot setting defines how many adjacent slots are active and

which of the active slots should be used for synchronization.

Note:

(1) This function is available only in conjunction with option GSM/EDGE

Analyzer FS-K5 option.

Remote-control command(s):

CONFigure:CHANnel:SLOT:MULTi ACT1SYNC1 | ACT2SYNC1 | ACT2SYNC2 |

ACT3SYNC1 | ACT3SYNC2 | ACT3SYNC3 | ACT4SYNC1 | ACT4SYNC2 | ACT4SYNC3 |

ACT4SYNC4 | ACT8SYNC1 | ACT8SYNC2 |              ACT8SYNC3 | ACT8SYNC4 |

ACT8SYNC5 | ACT8SYNC6 | ACT8SYNC7 | ACT8SYNC8

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

slots

Variable Type       ViInt32

This control selects the midamble used by the mobile.

Valid Range:

0 - 1 active slot, sync to 1st  (Default Value)

1 - 2 active slots, sync to 1st

2 - 2 active slots, sync to 2st

3 - 3 active slots, sync to 1st

4 - 3 active slots, sync to 2st

5 - 3 active slots, sync to 3st

6 - 4 active slots, sync to 1st

7 - 4 active slots, sync to 2st

8 - 4 active slots, sync to 3st

9 - 4 active slots, sync to 4st

10 - 8 active slots, sync to 1st

11 - 8 active slots, sync to 2st

12 - 8 active slots, sync to 3st

13 - 8 active slots, sync to 4st

14 - 8 active slots, sync to 5st

15 - 8 active slots, sync to 6st

16 - 8 active slots, sync to 7st

17 - 8 active slots, sync to 8st

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Slots) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMSLimitRestore

ViStatus rsfsu_actMSLimitRestore (ViSession instrumentHandle);

Purpose

Restores the GSM/EDGE limit lines. Any modifications made to the GSM/EDGE

limit lines are lost and the status upon delivery is restored.

Note:

(1) This function is available only in conjunction with option  FS-K5.

Remote-control command(s):

CONFigure:RESTore

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actMSRestore

ViStatus rsfsu_actMSRestore (ViSession instrumentHandle);

Purpose

This function restores the GSM limit lines. All previous changes to the

GSM limit lines are lost, and the default Limit Lines from the delivery

state are restored.

Note:

(1) This function is available only in conjunction with option GSM MS

Analyzer, FS-K5.

Remote-control command(s):

CONFigure:RESTore

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actNoiseMeasMode

ViStatus rsfsu_actNoiseMeasMode (ViSession instrumentHandle);

Purpose

This function selects active operation of the R&S FS-K30 option.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

INSTrument:SELect NOIS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actPMetReferenceValue

ViStatus rsfsu_actPMetReferenceValue (ViSession instrumentHandle,

ViInt32 window,

ViInt32 referenceValue,

ViReal64 manualReferenceValue);

Purpose

This function is used to obtain the currently measured power as

a reference value for the relative display or the reference value can

also be set manually. Manual entry of a reference value for relative

measurements is in the unit dBm.

Note:

(1) This function is available only in conjunction with option FS-K9.

Remote-control command(s):

CALCulate<1|2>:PMETer:RELative[:MAGNitude]

CALCulate<1|2>:PMETer:RELative[:MAGNitude]:AUTO ONCE

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

referenceValue

Variable Type       ViInt32

Selects whether to set reference value automatically from currently

measured power or manually.

Valid Range:

0 - Auto

1 - Manual

Default Value: 0

manualReferenceValue

Variable Type       ViReal64

Sets the reference value for relative measurements.

Valid Range: -200.0 dBm to 200.0 dBm

Default Value: 0.0 dBm

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Reference Value) out of range.

BFFC0004  Parameter 4 (Manual Reference Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actPMetSensorZeroing

ViStatus rsfsu_actPMetSensorZeroing (ViSession instrumentHandle);

Purpose

This function zeroes the power sensor.

Note:

(1) This function is available only in conjunction with option FS-K9.

Remote-control command(s):

CALibration:PMETer:ZERO:AUTO ONCE

*WAI

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actPMetState

ViStatus rsfsu_actPMetState (ViSession instrumentHandle, ViInt32 window,

ViBoolean powerMeterState);

Purpose

This function switches measurements with a power sensor on or off.

Note:

(1) This function is available only in conjunction with option FS-K9.

Remote-control command(s):

[SENSe<1|2>:]PMETer[:STATe] ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

powerMeterState

Variable Type       ViBoolean

Switches measurements with a power sensor on or off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Power Meter State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMACPLimitCheck

ViStatus rsfsu_actSAMACPLimitCheck (ViSession instrumentHandle,

ViInt32 window, ViInt32 ACPResult,

ViPInt32 firstResult,

ViPInt32 secondResult);

Purpose

This function queries result of the limit check for the first/second

alternate or upper/lower adjacent channel when adjacent channel power

measurement is performed.

Remote-control command(s):

CALCulate<1|2>:LIMit:ACPower:ACHannel:RESult?

CALCulate<1|2>:LIMit:ACPower:ALTernate<1|2>:RESult?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be used.

Valid Range:

1 - Window A (Default Value)

2 - Window B

ACPResult

Variable Type       ViInt32

This control selects result value.

Valid Range:

0 - Up/Lo Adj Ch (Default Value)

1 - Alternate 1 Ch

2 - Alternate 2 Ch

Notes:

(1) Up/Lo Adj Ch:

Upper/Lower adjacent channel power measurement result. First returned

value denotes the lower, the second denotes the upper adjacent

channel.

(2) Alternate 1, 2 Ch:

First/Second alternate adjacent channel power measurement result.

firstResult

Variable Type       ViInt32 (passed by reference)

The result is displayed in character data form. Possible values are:

0 - PASSED  ... limit not exceeded

1 - FAILED  ... limit exceeded

secondResult

Variable Type       ViInt32 (passed by reference)

The result is displayed in character data form. Possible values are:

0 - PASSED  ... limit not exceeded

1 - FAILED  ... limit exceeded

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (ACP Result) out of range.

BFFC0004  Parameter 4 (First Result) is null pointer.

BFFC0005  Parameter 5 (Second Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMACPLimitCheckState

ViStatus rsfsu_actSAMACPLimitCheckState (ViSession instrumentHandle,

ViInt32 window,

ViBoolean state);

Purpose

This function switches On or Off limit checking of ACP measurement.

Remote-control command(s):

CALCulate<1|2>:LIMit:ACPower ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

state

Variable Type       ViBoolean

Enables or disables limit checking.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMChannelPowerAutoAdjust

ViStatus rsfsu_actSAMChannelPowerAutoAdjust (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function adapts the reference level to the measured channel power.

Remote-control command(s):

[SENSe<1|2>:]POWer:ACHannel:PRESet:RLEVel

*WAI

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMChannelPowerAutoAdjustResult

ViStatus rsfsu_actSAMChannelPowerAutoAdjustResult

(ViSession instrumentHandle, ViInt32 window,

ViPInt32 adjustmentResult);

Purpose

This command returns PASSED on succsessful level adjustment or FAILED if

no optimum settings could be found.

Remote-control command(s):

[SENSe<1|2>:]POWer:ACHannel:PRESet:RLEVel?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

adjustmentResult

Variable Type       ViInt32 (passed by reference)

The result of adjustment.

Where:

0 - PASSED

1 - FAILED

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Adjustment Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMChannelPowerAutorange

ViStatus rsfsu_actSAMChannelPowerAutorange (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function starts the autorange routine for the reference level. This

also creates the relationship between trigger and subframe start.

Note:

This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

[SENSe<1|2>:]POWer:ACHannel:AUTO:LTIMe

*WAI

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMChannelPowerAutorangeResult

ViStatus rsfsu_actSAMChannelPowerAutorangeResult

(ViSession instrumentHandle, ViInt32 window,

ViPInt32 adjustmentResult, ViPReal64 trigger_toFrame,

ViPReal64 reserved);

Purpose

This function returns status of level adjustment.

Note:

This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

[SENSe<1|2>:]POWer:ACHannel:AUTO:LTIMe?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

adjustmentResult

Variable Type       ViInt32 (passed by reference)

The result of adjustment.

Where:

0 - PASSED

1 - FAILED

trigger_toFrame

Variable Type       ViReal64 (passed by reference)

Returns trigger to frame in seconds.

reserved

Variable Type       ViReal64 (passed by reference)

Reserved.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Adjustment Result) is null pointer.

BFFC0004  Parameter 4 (Trigger to Frame) is null pointer.

BFFC0005  Parameter 5 (Reserved) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMChannelPowerStartSlot

ViStatus rsfsu_actSAMChannelPowerStartSlot (ViSession instrumentHandle,

ViInt32 window,

ViInt32 startSlot);

Purpose

This function allows entry of the start slot for gated sweep mode. The

gated mode is on between START SLOT and STOP SLOT.

Note:

This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

[SENSe<1|2>:]POWer:ACHannel:SLOT:STARt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

startSlot

Variable Type       ViInt32

Sets start slot for gated sweep mode.

Valid Range: 1 to 7

Default Value: 4

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Start Slot) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMChannelPowerStopSlot

ViStatus rsfsu_actSAMChannelPowerStopSlot (ViSession instrumentHandle,

ViInt32 window,

ViInt32 stopSlot);

Purpose

This function allows entry of the stop slot for gated sweep mode. The

gated mode is on between START SLOT and STOP SLOT.

Note:

This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

[SENSe<1|2>:]POWer:ACHannel:SLOT:STOP

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

stopSlot

Variable Type       ViInt32

Sets stop slot for gated sweep mode.

Valid Range: 1 to 7

Default Value: 6

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Stop Slot) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMCopyTrace

ViStatus rsfsu_actSAMCopyTrace (ViSession instrumentHandle,

ViInt32 sourceTrace,

ViInt32 destinationTrace);

Purpose

This function copies data from one trace to another. The second operand

designates the source, the first operand the destination of the data to

be copied.

Note:

The associated measurement window is selected with the function Set

Active Window (rsfsu_confSetActiveWindow).

Remote-control command(s):

TRACe<1|2>:COPY TRACE1 | TRACE2 | TRACE3 , TRACE1 | TRACE2 | TRACE3

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

sourceTrace

Variable Type       ViInt32

This control sets the source trace.

Valid Range:

0 - Trace 1 (Default Value)

1 - Trace 2

2 - Trace 3

destinationTrace

Variable Type       ViInt32

This control sets the destination trace.

Valid Range:

0 - Trace 1

1 - Trace 2 (Default Value)

2 - Trace 3

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Source Trace) out of range.

BFFC0003  Parameter 3 (Destination Trace) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMCPSet

ViStatus rsfsu_actSAMCPSet (ViSession instrumentHandle, ViInt32 window,

ViInt32 channelPowerType);

Purpose

This function optimises automatically the settings of the analyzer for

the selected power measurement.

Remote-control command(s):

SENSe<1|2>:POWer:ACHannel:PRESet ACPower | CPOWer | OBANdwidth | CN | CN0

| MCACpower

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

channelPowerType

Variable Type       ViInt32

Specifies the channel power standard.

Valid Range:

0 - Adjacent Channel (Default Value)

1 - Channel Power

2 - Occupied Bandwidth

3 - C / N

4 - C / No

5 - Multi Carrier Adjacent Channel Power

Notes:

(1) Adjacent Channel:

Adjacent channel power measurement.

(2) Channel Power:

Channel power measurement.

(3) Occupied Bandwidth:

Occupied bandwidth power measurement.

(4) C / N

Carrier/Noise measurement.

(5) C / No

Carrier/Noise measurement normalized to 1 Hz bandwidth.

(6) Multi Carrier Adjacent Channel Power:

Adjacent channel power measurement with multiple triggers.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Channel Power Type) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMMarkChPowToRef

ViStatus rsfsu_actSAMMarkChPowToRef (ViSession instrumentHandle,

ViInt32 window);

Purpose

This sets the reference value to the currently measured channel power.

Remote-control command(s):

SENSe<1|2>:POWer:ACHannel:REFerence:AUTO ONCE

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMMarkCntValue

ViStatus rsfsu_actSAMMarkCntValue (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViPReal64 markerFrequency);

Purpose

This function gets frequency counter value from the marker position.

Note:

Frequency counter must be turned to On!

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:COUNt:FREQuency?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

markerFrequency

Variable Type       ViReal64 (passed by reference)

This control returns result of the frequency counter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Marker Frequency) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMMarkModDepthValue

ViStatus rsfsu_actSAMMarkModDepthValue (ViSession instrumentHandle,

ViInt32 window,

ViPReal64 AMModDepthResult);

Purpose

This function queries the AM modulation depth in the indicated

measurement window. A complete sweep with synchronization to sweep end

should be performed between switching on the

function and querying the measured value to obtain a valid query result.

This is only possible in the single sweep mode.

Remote-control command(s):

CALCulate<1|2>:MARKer:FUNCtion:MDEPth:RESult?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

AMModDepthResult

Variable Type       ViReal64 (passed by reference)

This control returns AM modulation depth measurement result.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (AM Mod Depth Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMMarkNdBDValue

ViStatus rsfsu_actSAMMarkNdBDValue (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViBoolean n_dBDownQuery,

ViReal64 _VI_FAR returnedValue[]);

Purpose

This function queries the frequency spacing (bandwidth) of the "N dB

Down" markers or queries the frequencies (time) of the "N dB Down"

marker. In zero span mode, the pulse width or marker x-values in seconds

are returned.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:NDBDown:RESult?

CALCulate<1|2>:MARKer<1..4>:FUNCtion:NDBDown:FREQuency?

CALCulate<1|2>:MARKer<1..4>:FUNCtion:NDBDown:TIME?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

n_dBDownQuery

Variable Type       ViBoolean

This control selects N dB Down query mode.

Valid Rnage:

VI_FALSE (0) - Result

VI_TRUE  (1) - Frequency (Time)

Default Value: VI_FALSE (0)

Notes:

(1) Result:

Queries the frequency spacing (bandwidth) of the "N dB Down" markers.

In zero span mode, the pulse width result (PWIDTH) is returned.

(2) Frequency (Time):

Queries the frequencies of the "N dB Down" marker. In zero span mode,

the two marker x-values in seconds are returned.

returnedValue

Variable Type       ViReal64[]

This control returns N dB Down value(s).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (N dB Down Query) out of range.

BFFC0005  Parameter 5 (Returned Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMMarkNoiseValue

ViStatus rsfsu_actSAMMarkNoiseValue (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViPReal64 noiseResult);

Purpose

This function queries the result of the noise measurement.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:NOISe:RESult?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

noiseResult

Variable Type       ViReal64 (passed by reference)

This control returns noise measurement result.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Noise Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMMarkPowerBandValue

ViStatus rsfsu_actSAMMarkPowerBandValue (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViReal64 _VI_FAR powerValue[]);

Purpose

This function queries the results of the occupied bandwidth power

measurement.

Remote-control command(s):

*CLS

CALCulate<1|2>:MARKer<1..4>:FUNCtion:POWer:RESult? OBANdwidth

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

powerValue

Variable Type       ViReal64[]

Measurement result of occupied bandwidth. The occupied bandwidth in

Hz is returned.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Power Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMMarkPowerValue

ViStatus rsfsu_actSAMMarkPowerValue (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViInt32 channelPowerType,

ViReal64 _VI_FAR powerValue[]);

Purpose

This function queries the results of the power measurement.

Note:

(1) This function is reserved for compatibility with FSEx driver. Prefer

to use function rsfsu_actSAMMarkPowerValueExt.

Remote-control command(s):

*CLS

CALCulation<1|2>:MARKer<1..4>:FUNCtion:POWer:RESult? ACPower | CPOWer |

OBANdwidth

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

channelPowerType

Variable Type       ViInt32

Specifies the channel power function to measure.

Valid Range:

0 - Adjacent Channel (Default Value)

1 - Channel Power

2 - Occupied Bandwidth

Notes:

(1) Adjacent Channel:

Adjacent channel power measurement. Results are output in the

following order:

Power of main channel

Power of lower adjacent channel 1

Power of upper adjacent channel 1

Power of lower adjacent channel 2

Power of upper adjacent channel 2

...

The number of results depends on the number of adjacent channels

selected. With logarithmic scaling (RANGE LOG), the power is output

in dBm, with linear scaling (RANGE LIN dB or LIN %) in W. If

SENSe:POWer:ACHannel:MODE REL is selected, adjacent channel power is

output in dB.

(2) Channel Power:

Channel power measurement. With logarithmic scaling (RANGE LOG), the

channel power is output in dBm, with linear scaling (RANGE LIN dB or

LIN %) in W.

(3) Occupied Bandwidth:

Occupied bandwidth power measurement. The return value is the

occupied bandwidth in Hz.

powerValue

Variable Type       ViReal64[]

This control returns array of results for the power measurement.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Channel Power Type) out of range.

BFFC0005  Parameter 5 (Power Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMMarkPowerValueExt

ViStatus rsfsu_actSAMMarkPowerValueExt (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViInt32 channelPowerType,

ViBoolean HzPowerResult,

ViReal64 _VI_FAR powerValue[]);

Purpose

This function queries the results of the time domain power measurement.

Remote-control command(s):

*CLS

CALCulate<1|2>:MARKer<1..4>:FUNCtion:POWer:RESult:PHZ ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:POWer:RESult? ACPower | CPOWer |

OBANdwidth | CN | CN0 | MCACpower

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

channelPowerType

Variable Type       ViInt32

Specifies the channel power function to measure.

Valid Range:

0 - Adjacent Channel Power (Default Value)

1 - Channel Power

2 - Occupied Bandwidth Power

3 - C / N

4 - C / No

5 - Multi Carrier Adjacent Channel Power

Notes:

(1) Adjacent Channel:

Adjacent channel power measurement. Results are output in the

following order:

Power of transmission channel

Power of lower adjacent channel

Power of upper adjacent channel

Power of lower alternate channel 1

Power of upper alternate channel 1

Power of lower alternate channel 2

Power of upper alternate channel 2

The number of results depends on the number of adjacent channels

selected. With logarithmic scaling (RANGE LOG), the power is output

in dBm, with linear scaling (RANGE LIN dB or LIN %) in W. If

SENSe:POWer:ACHannel:MODE REL is selected, adjacent channel power is

output in dB.

(2) Channel Power:

Channel power measurement. With logarithmic scaling (RANGE LOG), the

channel power is output in dBm, with linear scaling (RANGE LIN dB or

LIN %) in W.

(3) Occupied Bandwidth Power:

Occupied bandwidth power measurement. The occupied bandwidth in Hz is

returned.

(4) C / N:

The carrier-to-noise ratio in dB is returned.

(5) C / No:

Measurement of carrier-to-noise ratio referenced to 1 Hz bandwidth.

The carrier-to-noise ratio in dB/Hz is returned.

(6) Multi Carrier Adjacent Channel Power:

Channel/adjacent-channel power measurement with several carrier

signals Results are output in the following sequence:

Power of carrier signal 1

Power of carrier signal 2

Power of carrier signal 3

Power of carrier signal 4

.

.

.

Power of carrier signal 12

Total power of all carrier signals

Power of lower adjacent channel

Power of upper adjacent channel

Power of lower alternate channel 1

Power of upper alternate channel 1

Power of lower alternate channel 2

Power of upper alternate channel 2

The number of measured values returned depends on the number of

carrier signals and adjacent/alternate channels selected with

SENSe:POWer:ACHannel:TXCHannel:COUNt and

SENSe:POWer:ACHannel:ACPairs.

If only one carrier signal is measured, the total value of all

carrier signals will not be output.

With logarithmic scaling (RANGE LOG), the power is output in dBm;

with linear scaling (RANGE LIN dB or LIN %), the power is output in

W. If SENSe:POWer:ACHannel:MODE REL is selected, the

adjacent/alternate-channel power is output in dB.

HzPowerResult

Variable Type       ViBoolean

This control switches the query of the channel power measurement

results in the indicated measurement window between output of

absolute values (OFF) and output referred to the measurement bandwith

(ON).

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Notes:

(1) Off:

Results output in absolute values.

(2) On:

Results output referred to measurement bandwidth.

powerValue

Variable Type       ViReal64[]

This control returns array of results for the power measurement.

Notes:

(1) The array must contain at least 19 elements ViReal64[19].

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Channel Power Type) out of range.

BFFC0005  Parameter 5 (/Hz Power Result) out of range.

BFFC0006  Parameter 6 (Power Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMMode

ViStatus rsfsu_actSAMMode (ViSession instrumentHandle, ViInt32 window);

Purpose

This function sets the instrument to the spectrum analyzer mode.

Remote-control command(s):

INSTrument<1|2>:SELect SANalyzer

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be used.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMPhaseNoiseResult

ViStatus rsfsu_actSAMPhaseNoiseResult (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViPReal64 noiseResult);

Purpose

This function queries the result of the phase noise measurement.

Remote-control command(s):

CALCulate<1|2>:DELTamarker<1..4>:FUNCtion:PNOise:RESult?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

noiseResult

Variable Type       ViReal64 (passed by reference)

This control returns phase noise measurement result.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Noise Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMSigStat

ViStatus rsfsu_actSAMSigStat (ViSession instrumentHandle,

ViInt32 window_RESERVED,

ViInt32 traceNumber,

ViReal64 _VI_FAR results[]);

Purpose

This function reads out the results of statistical measurements of a

recorded trace.

Remote-control command(s):

CALCulate:STATistic:RESult<1|2|3>? ALL

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window_RESERVED

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

(2) This parameter is reserved for future use.

traceNumber

Variable Type       ViInt32

Selects the trace.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

results

Variable Type       ViReal64[]

Results of all measurements <mean power>,<peak power>,<crest factor>.

MEAN - Average power in dBm measured over the measurement time.

PEAK - Peak power in dBm measured over the measurement time.

CFACtor - Determined CREST factor (= ratio of peak power to

average power) in dB.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window (RESERVED)) out of range.

BFFC0003  Parameter 3 (Trace Number) out of range.

BFFC0004  Parameter 4 (Results) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMStoreSEToFile

ViStatus rsfsu_actSAMStoreSEToFile (ViSession instrumentHandle,

ViInt32 window, ViString fileName);

Purpose

This function stores the peak list of selected trace in ASCII format to a

file.

Note(s):

(1) This command is only available in GSM/EDGE mode (FS-K5).

Remote-control command(s):

MMEMory:STORe:SPUR <file_name>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 2 (screen A or B)

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

fileName

Variable Type       ViString

Specifies full path of the file, where spurios emissions data will be

stored.

Valid Range: not checked

Default Value: "a:\\test.txt"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (File Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSAMStoreTraceToFile

ViStatus rsfsu_actSAMStoreTraceToFile (ViSession instrumentHandle,

ViInt32 window,

ViInt32 sourceTrace,

ViString fileName);

Purpose

This function stores the selected trace (1 to 3) in the measurement

window indicated by window in a file with ASCII format.

Remote-control command(s):

MMEMory:STORe<1|2>:TRACe 1 to 3,<file_name>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 2 (screen A or B)

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

sourceTrace

Variable Type       ViInt32

This control sets the source trace.

Valid Range:

0 - Trace 1 (Default Value)

1 - Trace 2

2 - Trace 3

fileName

Variable Type       ViString

Specifies full path of the file, where trace data will be stored.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Source Trace) out of range.

BFFC0004  Parameter 4 (File Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSEMLimitLineCheck

ViStatus rsfsu_actSEMLimitLineCheck (ViSession instrumentHandle,

ViInt32 measValue,

ViPReal64 measResult);

Purpose

This function queries peak values for spectrum emission mask. If a signal

of high leakage power is measured with the spectrum emission mask

measurement, the power or frequency value of the worst FAIL of SEM is

read.

Note:

(1) This function is available only in conjunction with options FS-K72,

FS-K73, FS-K74, FS-K82, FS-K83, FS-K84, FS-K85.

Remote-control command(s):

CALCulate:LIMit:ESPectrum:CHECk:X?

CALCulate:LIMit:ESPectrum:CHECk:Y?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measValue

Variable Type       ViInt32

Selects value of the worst fail of a SEM measurement to be queried.

Valid Range:

0 - Frequency

1 - Power

Default Value: 0 - Frequency

measResult

Variable Type       ViReal64 (passed by reference)

Returns freqeuncy (in Hz) or power (in dBm) value of the worst fail

of a SEM measurement.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Meas Value) out of range.

BFFC0003  Parameter 3 (Meas Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSendTrg

ViStatus rsfsu_actSendTrg (ViSession instrumentHandle);

Purpose

This function triggers all actions waiting for a trigger event in the

currently active window.

Remote-control command(s):

*TRG

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSendTrgWindow

ViStatus rsfsu_actSendTrgWindow (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function triggers all actions waiting for a trigger event in the

selected window.

Remote-control command(s):

INITiate<1|2>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSendTrgWopc

ViStatus rsfsu_actSendTrgWopc (ViSession instrumentHandle,

ViInt32 timeoutms);

Purpose

This function triggers all actions waiting for a trigger event in the

currently active window and waits for operation completed (OPC) before

returning the status code.

Remote-control command(s):

*TRG

*OPC?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

timeoutms

Variable Type       ViInt32

Sets the timeout for the triggering routine to be finished. If the

length of time required for triggering exceeds the timeout value,

then the function will return with a timeout error.

Valid Range: 0 ms to 600000 ms

Default Value: 180000 ms

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Timeout [ms]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSendTrgWopcWindow

ViStatus rsfsu_actSendTrgWopcWindow (ViSession instrumentHandle,

ViInt32 window, ViInt32 timeoutms);

Purpose

This function triggers all actions waiting for a trigger event in the

selected window and waits for operation completed (OPC) before returning

the status code.

Remote-control command(s):

INITiate<1|2>

*OPC?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

timeoutms

Variable Type       ViInt32

Sets the timeout for the triggering routine to be finished. If the

length of time required for triggering exceeds the timeout value,

then the function will return with a timeout error.

Valid Range: 0 ms to 600000 ms

Default Value: 180000 ms

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Timeout [ms]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actServiceModuleInfo

ViStatus rsfsu_actServiceModuleInfo (ViSession instrumentHandle,

ViInt32 stringSize,

ViChar _VI_FAR modules[],

ViPInt32 returnedValues);

Purpose

This function queries the contents of the module info table.

Remote-control command(s):

DIAGnostic:SERVice:HWINfo?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

stringSize

Variable Type       ViInt32

Thsi control sets the size of string.

modules

Variable Type       ViChar[]

Returns the  modules description. The module descriptions are divided

by comma. The structure of module description is:

<component> <serial #> <order #> <model> <HWC> <rev> <sub rev>

returnedValues

Variable Type       ViInt32 (passed by reference)

Returns number of returned characters.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (String Size) out of range.

BFFC0003  Parameter 3 (Modules) is null pointer.

BFFC0004  Parameter 4 (Returned Values) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actServiceSource

ViStatus rsfsu_actServiceSource (ViSession instrumentHandle,

ViBoolean input,

ViBoolean noiseSource);

Purpose

This function supports the instrument's signal sources.

Note:

(1) This function is reserved for compatibility with FSEx driver. Prefer

to use function rsfsu_actServiceSourceExt.

Remote-control command(s):

DIAGnostic:SERVice:INPut:SELect CALibration | RF

DIAGnostic:SERVice:NSOurce ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

input

Variable Type       ViBoolean

This control selects the reference signal as either the RF input on

the front panel, or the internal 120-MHz reference signal.

Valid Range:

VI_FALSE(0) - RF (Default Value)

VI_TRUE (1) - CAL

noiseSource

Variable Type       ViBoolean

This control switches the 28V supply at the rear connector of the

noise source either on or off.

Valid Range:

VI_FALSE(0) - Off (Default Value)

VI_TRUE (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Input) out of range.

BFFC0003  Parameter 3 (Noise Source) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actServiceSourceCalSignal

ViStatus rsfsu_actServiceSourceCalSignal (ViSession instrumentHandle,

ViInt32 window,

ViInt32 pulseState,

ViInt32 pulseRate);

Purpose

This function configures the calibration signal in the selected

measurement window.

Remote-control command(s):

DIAGnostic<1|2>:SERVice:INPut:PULSe:STATe ON | OFF

DIAGnostic<1|2>:SERVice:INPut:PULSe:PRATe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

pulseState

Variable Type       ViInt32

This control toggles the calibration signal in the selected

measurement window between pulsed and non-pulsed.

Valid Range:

0 - Non Pulsed (Default Value)

1 - Pulsed

Note:

(1) The selection takes effect only if the RF input has been set to

the internal reference signal.

pulseRate

Variable Type       ViInt32

This control selects the pulse rate for the pulsed calibration signal

in the selected measurement window.

Valid Range:

0 - 10 kHz

1 - 62.5 kHz

2 - 128 MHz (Default Value)

3 - 640 MHz

4 - 100 kHz

5 - 1 MHz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Pulse State) out of range.

BFFC0004  Parameter 4 (Pulse Rate) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actServiceSourceExt

ViStatus rsfsu_actServiceSourceExt (ViSession instrumentHandle,

ViBoolean input,

ViReal64 calSourceLeveldBm,

ViBoolean noiseSource);

Purpose

This function supports the instrument's signal sources.

Remote-control command(s):

DIAGnostic:SERVice:INPut:SELect RF | CALibration

DIAGnostic:SERVice:CSOurce

DIAGnostic:SERVice:NSOurce ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

input

Variable Type       ViBoolean

This control toggles between the RF input on the front panel and the

internal 128-MHz reference signal. The level of the 128-MHz signals

can be selected by control Cal Source Level.

Valid Range:

VI_FALSE(0) - RF (Default Value)

VI_TRUE (1) - CAL

calSourceLeveldBm

Variable Type       ViReal64

This control switches the level of the 128 MHz reference signal

source if selected.

Valid Range: -30.0 dBm to 0.0 dBm

Default Value: -30.0 dBm

noiseSource

Variable Type       ViBoolean

This control switches the 28V supply at the rear connector of the

noise source either on or off.

Valid Range:

VI_FALSE(0) - Off (Default Value)

VI_TRUE (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Input) out of range.

BFFC0003  Parameter 3 (Cal Source Level [dBm]) out of range.

BFFC0004  Parameter 4 (Noise Source) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSEStartMeasurement

ViStatus rsfsu_actSEStartMeasurement (ViSession instrumentHandle);

Purpose

This function starts a spurious measurement.

Remote-control command(s):

INITiate<1|2>:SPURious

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSEStartMeasurementWopc

ViStatus rsfsu_actSEStartMeasurementWopc (ViSession instrumentHandle,

ViInt32 timeoutms);

Purpose

This function starts a spurious measurement and waits for OPC.

Remote-control command(s):

INITiate<1|2>:SPURious

*WAI

*OPC?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

timeoutms

Variable Type       ViInt32

Sets the timeout for the triggering routine to be finished. If the

length of time required for triggering exceeds the timeout value,

then the function will return with a timeout error.

Valid Range: 0 ms to 600000 ms

Default Value: 180000 ms

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Timeout [ms]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSetLimitLinesOffset

ViStatus rsfsu_actSetLimitLinesOffset (ViSession instrumentHandle,

ViInt32 window, ViReal64 xOffset,

ViReal64 yOffset);

Purpose

This function sets the X and Y limit line offsets.

Remote-control command(s):

CALCulate<1|2>:LIMit:CONTrol:OFFSET

CALCulate<1|2>:LIMit:UPPer:OFFSET

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

xOffset

Variable Type       ViReal64

This control defines an offset for the X-axis value of all active

relative limit lines.

Valid Range: not range checked

Default Value: 0.0

yOffset

Variable Type       ViReal64

This control defines an offset for the Y-axis value of all active

relative limit lines.

Valid Range: not range checked

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (X Offset) out of range.

BFFC0004  Parameter 4 (Y Offset) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSetLimitLineSpacing

ViStatus rsfsu_actSetLimitLineSpacing (ViSession instrumentHandle,

ViInt32 window,

ViInt32 limitLineNo,

ViInt32 spacing);

Purpose

This function makes a selection between linear and logarithmic

interpolation for determining the limit line from the frequency points.

Remote-control command(s):

CALCulate<1|2>:LIMit<1..8>:CONTrol:SPACing LINear | LOGarithmic

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

spacing

Variable Type       ViInt32

This control selects the spacing type.

Valid Values:

RSFSU_LOG (0) - Logarithmic Spacing

RSFSU_LIN (1) - Linear Spacing (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Spacing) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actSetParamLimitLineSpacing

ViStatus rsfsu_actSetParamLimitLineSpacing (ViSession instrumentHandle,

ViInt32 window,

ViInt32 limitLineNo,

ViBoolean lineParameter,

ViInt32 spacing);

Purpose

This function makes a selection between linear and logarithmic

interpolation for the upper or lower limit line.

Remote-control command(s):

CALCulate<1|2>:LIMit<1..8>:LOWer:SPACing LINear | LOGarithmic

CALCulate<1|2>:LIMit<1..8>:UPPer:SPACing LINear | LOGarithmic

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

lineParameter

Variable Type       ViBoolean

This control selects which set of parameters will be actually

configured.

Valid Range:

VI_FALSE (0) - Lower (Default Value)

VI_TRUE  (1) - Upper

spacing

Variable Type       ViInt32

This control selects the spacing type.

Valid Values:

RSFSU_LOG (0) - Logarithmic Spacing

RSFSU_LIN (1) - Linear Spacing (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Line Parameter) out of range.

BFFC0005  Parameter 5 (Spacing) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actTDSCDMACDPMarkMeas

ViStatus rsfsu_actTDSCDMACDPMarkMeas (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViInt32 measurement,

ViPReal64 measurementResult);

Purpose

This function polls the measured and calculated values of code domain

power analysis. The results are provided for the channel to which the

code selected by the CDPower:CODe command belongs.

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CALCulate<1|2>:MARKer<1>:FUNCtion:CDPower:RESult? SLOT | PDATa | PD1 |

PD2 | PMIDamble | RHO | MACCuracy | PCDerror | FERRor | CERRor | TFRame I

IQOFfset | IQIMbalance | ACTive | SRATe | CHANnel | SFACtor | CDPabsolute

| CDPRelative | EVMRms | EVMPeak

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 4

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

Note:

1) This value is ignored. Default value has to be used.

measurement

Variable Type       ViInt32

Selects value to get.

Valid Range:

Global results of selected slot:

0 - Slot number

1 - Power data fields in dBm

2 - Power data field 1 in dBm

3 - Power data field 2 in dBm

4 - Power midamble in dBm

5 - RHO

6 - Composite EVM in %

7 - Peak code domain error in dB

8 - Frequency error in Hz

9 - Chip rate error in ppm

10 - Trigger to frame

11 - IQ imbalance in %

12 - IQ offset in %

13 - Number of active channels

Channel results:

14 - Data rate in kbps

15 - Channel number

16 - Spreading factor of channel

17 - Channel power relative in dB

18 - Channel power absolute in dBm

19 - Error vector magnitude RMS in %

20 - Error vector magnitude Peak in %

Default Value: 0 - Slot number

measurementResult

Variable Type       ViReal64 (passed by reference)

CDP measurement result. Format of the results depends on the selected

measurment.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Measurement) out of range.

BFFC0005  Parameter 5 (Measurement Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actTDSCDMACDPMode

ViStatus rsfsu_actTDSCDMACDPMode (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function activates TD-SCDMA BTS (BTDS) operating mode.

Note:

(1) This function is available only with option FS-K76.

Remote-control command(s):

INSTrument<1|2>:SELect BTDScdma

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actTDSCDMACDPMSMode

ViStatus rsfsu_actTDSCDMACDPMSMode (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function activates TD-SCDMA MS (MTDScdma) operating mode.

Note:

(1) This function is available only with option FS-K77.

Remote-control command(s):

INSTrument<1|2>:SELect MTDScdma

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actTOI

ViStatus rsfsu_actTOI (ViSession instrumentHandle, ViInt32 window,

ViPReal64 result);

Purpose

This function queries the result of the intercept point measurement in

the indicated measurement window. A complete sweep with synchronization

to sweep end should be performed between switching on the function and

querying the measured value to obtain a valid query result. This is only

possible in the single sweep mode.

Remote-control command(s):

CALCulate<1|2>:MARKer:FUNCtion:TOI:RESult?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

result

Variable Type       ViReal64 (passed by reference)

Returns TOI measurement result.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actTrackCorr

ViStatus rsfsu_actTrackCorr (ViSession instrumentHandle, ViInt32 window,

ViInt32 sourceCalibration,

ViInt32 timeoutms);

Purpose

This function calls the tracking generator calibrations.

Notes:

(1) This function is only valid in conjunction with option Tracking

Generator, FSU-B9.

Remote-control command(s):

SENSe<1|2>:CORRection:RECall

SENSe<1|2>:CORRection:METHod TRANsmission | REFLexion

SENSe<1|2>:CORRection:COLLect THRough | OPEN

*OPC?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

sourceCalibration

Variable Type       ViInt32

This control selects the kind of measurement with active tracking

generator (transmission/reflexion) and kind of measurement for the

reference values of the normalization (response calibration).

Valid Range:

0 - Transmission Calibration (Default Value)

1 - Reflection Calibration Open

2 - Reflection Calibration Short

3 - Recall Calibration

timeoutms

Variable Type       ViInt32

Sets the timeout for the source calibration routine to be finished.

If the length of time required for calibration exceeds the timeout

value, then the function will return with a timeout error.

Valid Range: 0 ms to 600000 ms

Default Value: 10000 ms

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Source Calibration) out of range.

BFFC0004  Parameter 4 (Timeout [ms]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actTrackCorrState

ViStatus rsfsu_actTrackCorrState (ViSession instrumentHandle,

ViInt32 window, ViBoolean corrState);

Purpose

This function activates/deactivates normalization of the measurement

results.

Notes:

(1) This function is only valid in conjunction with option Tracking

Generator, FSU-B9.

Remote-control command(s):

SENSe<1|2>:CORRection:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

corrState

Variable Type       ViBoolean

This control activates/deactivates normalization of the measurement

results.

Valid Range:

VI_TRUE  (1) - On

VI_FALSE (0) - Off (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Corr State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actTrackMode

ViStatus rsfsu_actTrackMode (ViSession instrumentHandle, ViInt32 window,

ViBoolean setMode);

Purpose

This function sets the tracking generator operating modes.

Notes:

(1) Switchover to tracking generator mode is only possible in conjunction

with option FSU-B9, Tracking Generator.

Remote-control command(s):

OUTPut<1|2>:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

setMode

Variable Type       ViBoolean

This control sets the output state of tracking generator.

Valid Range:

VI_TRUE  (1) - Tracking Generator On

VI_FALSE (0) - Tracking Generator Off (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Set Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actTVTriggerSrc

ViStatus rsfsu_actTVTriggerSrc (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function sets the trigger sequence source to TV mode.

Note:

(1) This function is only valid if option FS-B6, TV Demodulator is

fitted.

Remote-control command(s):

TRIGger<1|2>:SEQuence:SOURce TV

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSAMMarkerADemod

ViStatus rsfsu_actVSAMMarkerADemod (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViInt32 modulationType,

ViInt32 parameterToRead,

ViInt32 trace, ViPReal64 value);

Purpose

This function returns the FM measurement result in FM demodulation mode.

Notes:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:ADEMod:AM:RESult<1..3>? PPEak |

MPEak | MIDDle |RMS

CALCulate<1|2>:MARKer<1..4>:FUNCtion:ADEMod:FM:RESult<1..3>? PPEak |

MPEak | MIDDle |RMS

CALCulate<1|2>:MARKer<1..4>:FUNCtion:ADEMod:PM:RESult<1..3>? PPEak |

MPEak | MIDDle |RMS

CALCulate<1|2>:MARKer<1..4>:FUNCtion:ADEMod:AFRequency:RESult<1..3>?

CALCulate<1|2>:MARKer<1..4>:FUNCtion:ADEMod:FERRor:RESult<1..3>?

CALCulate<1|2>:MARKer<1..4>:FUNCtion:ADEMod:CARRier:RESult<1..3>?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

modulationType

Variable Type       ViInt32

This control sets modulation type.

Valid Range:

0 - AM

1 - FM (Default Value)

2 - PM

Notes:

(1) AM:

AM modulation measurement results of the analog demodulation.

(2) FM:

FM modulation measurement results of the analog demodulation.

(3) PM:

PM modulation measurement results of the analog demodulation.

parameterToRead

Variable Type       ViInt32

Specifies the parameter to read.

Valid Range:

0 - Pos PK (Default Value)

1 - Neg PK

2 - Average PK

3 - RMS

4 - Reserved

5 - Audio Frequency

6 - Frequency Error

7 - Carrier Power

trace

Variable Type       ViInt32

Selects the trace.

Valid Range:

1 to 3

Default Value:

1

value

Variable Type       ViReal64 (passed by reference)

Returns measured value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Modulation Type) out of range.

BFFC0005  Parameter 5 (Parameter To Read) out of range.

BFFC0006  Parameter 6 (Trace) out of range.

BFFC0007  Parameter 7 (Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSAMMode

ViStatus rsfsu_actVSAMMode (ViSession instrumentHandle, ViInt32 window,

ViInt32 setMode);

Purpose

This function selects the FM demodulation mode.

Notes:

(1) Switchover to analog demodulation is only possible in conjunction

with option FM Demodulator FS-K7.

Remote-control command(s):

INSTrument<1|2>:SELect ADEMod

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

setMode

Variable Type       ViInt32

This control switches between the operating modes.

Valid Values:

0 - Vector Analyzer - Analog Demod (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Set Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSBToothCarrierFreqParam

ViStatus rsfsu_actVSBToothCarrierFreqParam (ViSession instrumentHandle,

ViInt32 parameter,

ViPReal64 value);

Purpose

This function returns the carrier frequency parameters for Bluetooth

standard.

Notes:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CALCulate:BTOoth:ICFTolerance? AVERage

CALCulate:BTOoth:ICFTolerance? MINimum

CALCulate:BTOoth:ICFTolerance? MAXimum

CALCulate:BTOoth:CFDRift:MAXimum?

CALCulate:BTOoth:CFDRift:RATE?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

parameter

Variable Type       ViInt32

Selcts the parameter.

Valid Range:

0 - Initial Carr. Freq. tolerance Average(Default Value)

1 - Initial Carr. Freq. tolerance Min

2 - Initial Carr. Freq. tolerance Max

3 - Max. Carrier Frequncy Drift

4 - Carrier Frequency Drift

value

Variable Type       ViReal64 (passed by reference)

Returns measured value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Parameter) out of range.

BFFC0003  Parameter 3 (Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSBToothDM

ViStatus rsfsu_actVSBToothDM (ViSession instrumentHandle,

ViInt32 frequency, ViInt32 parameter,

ViInt32 modificator,

ViPReal64 parameterValue);

Purpose

This function returns the digital measurement parameters of the Bluetooth

standard.

Notes:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

(2) Only one parameter should be returned per call.

Remote-control command(s):

CALCulate:BTOoth:MCHar:DF<1|2>:AVERage? MINimum

CALCulate:BTOoth:MCHar:DF<1|2>:AVERage? MAXimum

CALCulate:BTOoth:MCHar:DF<1|2>:MAXimum? MINimum

CALCulate:BTOoth:MCHar:DF<1|2>:MAXimum? MAXimum

CALCulate:BTOoth:MCHar:DF<1|2>:MAXimum? AVERage

CALCulate:BTOoth:MCHar:RATio? MINimum

CALCulate:BTOoth:MCHar:RATio? MAXimum

CALCulate:BTOoth:MCHar:RATio? AVERage

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

frequency

Variable Type       ViInt32

Selects the frequency.

Valid Range:

1 to 2

Default Value:

1

Notes:

(1) This control has no effect for Parameter Ratio Frequency Rise.

parameter

Variable Type       ViInt32

Selects the measured value.

Valid Range:

0 - Average Frequency Deviation (Valid Value)

1 - Maximum Frequency Deviation

2 - Frequency Deviation Ratio

Notes:

(1) Average Frequency Deviation:

Returns the average frequency deviation for different bitmaps of the

payload. This selection is valid only if Modulation Characteristics

Measurement is selected with Bluetooth Measurement Mode

(actBToothMeasMode).

(2) Maximum Frequency Deviation:

Returns the maximum frequency deviation for different bitmaps of the

payload. This selection is valid only if Modulation Characteristics

Measurement is selected with Bluetooth Measurement Mode

(actBToothMeasMode).

(3) Frequency Deviation Ratio:

Returns the behaviour of the average frequency deviation for

different bitmaps of the payload. This selection is valid only if

Modulation Characteristics Measurement is selected with Bluetooth

Measurement Mode (actBToothMeasMode).

modificator

Variable Type       ViInt32

Selects the modificator of the command.

Valid Range:

0 - Minimum (Default Value)

1 - Maximum

2 - Average

Notes:

(1) For Parameter Average Frequency Deviation are valid only values

Minimum and Maximum.

parameterValue

Variable Type       ViReal64 (passed by reference)

Returns measured value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Frequency) out of range.

BFFC0003  Parameter 3 (Parameter) out of range.

BFFC0004  Parameter 4 (Modificator) out of range.

BFFC0005  Parameter 5 (Parameter Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSBToothException

ViStatus rsfsu_actVSBToothException (ViSession instrumentHandle,

ViPInt32 exceptions);

Purpose

This function returns the number of exceptions which occurred during

measurement on adjacent channels.

Notes:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

(2) This function returns correct values only if Adjacent Channel Power

Measurement is selected with Bluetooth Measurement Mode

(actBToothMeasMode).

Remote-control command(s):

CALCulate:BTOoth:ACLR:EXCeptions?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

exceptions

Variable Type       ViInt32 (passed by reference)

Returns the number of exceptions.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Exceptions) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSBToothOutputPower

ViStatus rsfsu_actVSBToothOutputPower (ViSession instrumentHandle,

ViInt32 outputPower,

ViPReal64 value);

Purpose

This function returns the output power measurement according Bluetooth

standard.

Notes:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

(2) The output power measurement must be set with function Bluetooth

Measurement Mode (actBToothMeasMode). In other cases the function returns

error.

Remote-control command(s):

CALCulate:BTOoth:OPOWer:PEAK?

CALCulate:BTOoth:OPOWer:AVERage? MINimum

CALCulate:BTOoth:OPOWer:AVERage? MAXimum

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

outputPower

Variable Type       ViInt32

Selects the output poer type.

Valid Range:

0 - Peak (Default Value)

1 - Average Min

2 - Average Max

value

Variable Type       ViReal64 (passed by reference)

Returns measured value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Output Power) out of range.

BFFC0003  Parameter 3 (Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSBToothPacketLength

ViStatus rsfsu_actVSBToothPacketLength (ViSession instrumentHandle,

ViPInt32 value);

Purpose

This function returns the length of the most recently measured packet for

the Bluetooth standard.

Notes:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CALCulate:BTOoth:PLENgth?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

value

Variable Type       ViInt32 (passed by reference)

Returns packet Length.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSBToothPacketType

ViStatus rsfsu_actVSBToothPacketType (ViSession instrumentHandle,

ViPInt32 value);

Purpose

This function returns the type of the mostly recently measured packet for

the Bluetooth standard.

Notes:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CALCulate:BTOoth:PTYPe?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

value

Variable Type       ViInt32 (passed by reference)

Returns packet type.

0 - AUX1

1 - DH1

2 - DH3

3 - DH5

4 - DM1

5 - DM3

6 - DM5

7 - FHS

8 - HV1

9 - HV2

10 - HV3

11 - DV

12 - NULL

13 - POLL

14 - UNDEF

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSBToothParameter

ViStatus rsfsu_actVSBToothParameter (ViSession instrumentHandle,

ViInt32 parameter,

ViPReal64 parameterValue,

ViPBoolean bluetoothStatus);

Purpose

This function returns the different parameters of the Bluetooth standard.

Notes:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

(2) Only one parameter should be returned in time.

Remote-control command(s):

*CLS

CALCulate:BTOoth:STATus?

CALCulate:BTOoth:PDENsity?

CALCulate:BTOoth:FRANge?

CALCulate:BTOoth:OBANdwidth?

CALCulate:BTOoth:OBANdwidth:FLOW?

CALCulate:BTOoth:OBANdwidth:FHIGh?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

parameter

Variable Type       ViInt32

Selects the bluetooth parameter.

Valid Range:

0 - Power Density (Valid Value)

1 - TX Output Spectrum Frequency Range

2 - TX Output Spectrum 20 dB Bandwidth

3 - TX Output Spectrum 20 dB Bandwidth High

4 - TX Output Spectrum 20 dB Bandwidth Low

Notes:

(1) Power Density:

Returns the power density of the measured signal. This selection is

valid only if Power Density measurement is selected with function

Bluetooth Measurement Mode (actBToothMeasMode).

(2) TX Output Spectrum Frequency Range:

Returns the TX Output Spectrum Frequency Range for selected channel.

This selection is valid only if TX Output Spectrum Frequency Range

measurement is selected with function Bluetooth Measurement Mode

(actBToothMeasMode).

(3) TX Output Spectrum 20 dB Bandwidth:

Returns the TX Output Spectrum 20 dB Bandwidth. This selection is

valid only if  TX Output Spectrum 20 dB Bandwidth measurement is

select with function Bluetooth Measurement Mode (actBToothMeasMode).

(4) TX Output Spectrum 20 dB Bandwidth High:

Returns the max. frequency of TX Output Spectrum 20 dB Bandwidth.

This selection is valid only if  TX Output Spectrum 20 dB Bandwidth

measurement is selected with function Bluetooth Measurement Mode

(actBToothMeasMode).

(5) TX Output Spectrum 20 dB Bandwidth Low:

Returns the min. frequency of TX Output Spectrum 20 dB Bandwidth.

This selection is valid only if  TX Output Spectrum 20 dB Bandwidth

measurement is selected with function Bluetooth Measurement Mode

(actBToothMeasMode).

parameterValue

Variable Type       ViReal64 (passed by reference)

Returns measured value.

bluetoothStatus

Variable Type       ViBoolean (passed by reference)

Returns the status of the last measurement.

VI_FALSE (0) -  PASS

VI_TRUE  (1) -  FAIL

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Parameter) out of range.

BFFC0003  Parameter 3 (Parameter Value) is null pointer.

BFFC0004  Parameter 4 (Bluetooth Status) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSBToothPercentageDev

ViStatus rsfsu_actVSBToothPercentageDev (ViSession instrumentHandle,

ViPReal64 value);

Purpose

This function returns the percentage deviation for Bluetooth standard.

Notes:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

(2) The Modulation Characteristics Measurement must be selected with

Bluetooth Measurement Mode (actBToothMeasMode).

Remote-control command(s):

CALCulate:BTOoth:MCHar:DF2:PERCent?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

value

Variable Type       ViReal64 (passed by reference)

Returns measured value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSBToothPowerControl

ViStatus rsfsu_actVSBToothPowerControl (ViSession instrumentHandle,

ViInt32 mode, ViInt32 arraySize,

ViReal64 _VI_FAR powerList[],

ViPInt32 returnedValues);

Purpose

This function returns the power control mode of the Bluetooth standard.

Notes:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CALCulate:BTOoth:PCONtrol:RELative?

CALCulate:BTOoth:PCONtrol:ABSolute?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

mode

Variable Type       ViInt32

Selects the mode of power control.

Valid Range:

0 - Relative

1 - Absolute (Default Value)

Note:

(1) Absolute:

Returns the list of power steps. The max. output power is returned as

first item.

(2) Relative:

Returns the list of relative power steps.

arraySize

Variable Type       ViInt32

This parameter determines the number of item in array list.

This value mustn't be great than items alloceated in Power List.

powerList

Variable Type       ViReal64[]

Returns the list of power values. These values are used for power

control.

Note:

(1) Rating Class 'Absolute':

Returns the list of power steps. The max. output power is returned as

first item.

(2) Rating Class 'Relative':

Returns the list of relative power steps. This means the list of

differences between steps.

returnedValues

Variable Type       ViInt32 (passed by reference)

Returns number of values, that were read from instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Mode) out of range.

BFFC0003  Parameter 3 (Array Size) out of range.

BFFC0004  Parameter 4 (Power List) is null pointer.

BFFC0005  Parameter 5 (Returned Values) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actVSBToothPowerOfChannels

ViStatus rsfsu_actVSBToothPowerOfChannels (ViSession instrumentHandle,

ViInt32 arraySize,

ViReal64 _VI_FAR power[],

ViPInt32 returnedValues);

Purpose

This function returns the power of selected adjacent channels of the

Bluetooth standard.

Notes:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

(2) This function returns correct values only if Adjacent Channel Power

Measurement is selected with function Bluetooth Measurement Mode

(actBToothMeasMode).

Remote-control command(s):

CALCulate:BTOoth:ACLR:LIST?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

arraySize

Variable Type       ViInt32

Defines the number of items in array Power.

power

Variable Type       ViReal64[]

Returns power of selected channels.

returnedValues

Variable Type       ViInt32 (passed by reference)

Returned number of values readed from instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Array Size) out of range.

BFFC0003  Parameter 3 (Power) is null pointer.

BFFC0004  Parameter 4 (Returned Values) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actWCDPMarkMeas

ViStatus rsfsu_actWCDPMarkMeas (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 measurement,

ViPReal64 measurementResult);

Purpose

This function queries the measured and calculated results of the WCDMA

code domain power measurement.

Where types of measurements are:

Total Power (Default Value)

Frequency Error in Hz

Trigger to Frame

Timing Offset

Modulation Accuracy

Peak Code Domain Error

Error Vector Magnitude RMS

Error Vector Magnitude Peak

Chip Rate Error

Channel Slot Number

Symbol Rate

Channel Number

Channel Power Absolute

Channel Power Relative

IQ Offset

IQ imbalance

Modulation Type

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73 or FS-K74.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:WCDPower:RESult? PTOTal | FERRor |

TFRame | TOFFset | MACCuracy | PCDerror | EVMRms | EVMPeak | CERRor |

CSLot | SRATe | CHANnel | CDPabsolute | CDPRelative | IQOFfset |

IQIMbalance | MTYPe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 4

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

Note:

1) This value is ignored. Default value has to be used.

measurement

Variable Type       ViInt32

Selects value to get.

Valid Range:

0 - Total Power (Default Value)

1 - Frequency Error in Hz

2 - Trigger to Frame

3 - Timing Offset

4 - Modulation Accuracy

5 - Peak Code Domain Error

6 - Error Vector Magnitude RMS

7 - Error Vector Magnitude Peak

8 - Chip Rate Error

9 - Channel Slot Number

10 - Symbol Rate

11 - Channel Number

12 - Channel Power Absolute

13 - Channel Power Relative

14 - IQ Offset

15 - IQ imbalance

16 - Modulation Type

Notes:

(1) Modulation Type:

Modulation type of the code channel at the selected channel slot of

the selected channel. The channel can be addressed by using SELECTED

CHANNEL. The channel slot can be adressed by using SELECT CPICH.

Where results are:

2  - QPSK modulation type QPSK

4  - 16QAM modulation type 16QAM

15 - NONE there is no power in the selected channel slot.

[slot is switched OFF.] (According to 3GPP the power of a

HSDPA channel can be switched every 2ms, which are 3

slots.)

measurementResult

Variable Type       ViReal64 (passed by reference)

WCDP measurement result. Format of the results depends on the

selected measurment.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Measurement) out of range.

BFFC0005  Parameter 5 (Measurement Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actWCDPMode

ViStatus rsfsu_actWCDPMode (ViSession instrumentHandle, ViInt32 window);

Purpose

This function sets the WCDP (Wideband Code Domain Power) BTS analyzer

mode.

Note:

(1) This function is available only with FS-K72 3GPP FDD Base Station

Test or FS-K74 3GPP HSDPA Base Station Test option.

Remote-control command(s):

INSTrument<1|2>:SELect WCDPower

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actWCDPMSMarkMeas

ViStatus rsfsu_actWCDPMSMarkMeas (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 measurement,

ViPReal64 measurementResult);

Purpose

This function queries the measured and calculated results of the WCDMA

code domain power measurement.

Where types of measurements are:

Total Power

Frequency Error in Hz

Trigger to Frame

Modulation Accuracy

Peak Code Domain Error

Error Vector Magnitude RMS

Error Vector Magnitude Peak

Chip Rate Error

Symbol Rate

Channel Number

Channel Power Absolute

Channel Power Relative

IQ Offset

IQ Imbalance

Channel Mapping

No of Pilot Symbols

Notes:

(1) Window 1 and marker 1 are only permissible values to be used.

(2) This function is available only in conjunction with option FS-K72 and

FS-K73.

Remote-control command(s):

CALCulate1:MARKer1:FUNCtion:WCDPower:MS:RESult?  PTOTal | FERRor | TFRame

| MACCuracy | PCDerror | EVMRms | EVMPeak    | CERRor | SRATe | CHANnel |

CDPabsolute | CDPRelative | IQOFfset | IQIMbalance | CMAPping | PSYMbo

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 1

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

Note:

1) This value is ignored. Default value has to be used.

measurement

Variable Type       ViInt32

Selects value to get.

Valid Range:

0 - Total Power (Default Value)

1 - Frequency Error in Hz

2 - Trigger to Frame

3 - Modulation Accuracy

4 - Peak Code Domain Error

5 - Error Vector Magnitude RMS

6 - Error Vector Magnitude Peak

7 - Chip Rate Error

8 - Symbol Rate

9 - Channel Number

10 - Channel Power Absolute

11 - Channel Power Relative

12 - IQ Offset

13 - IQ Imbalance

14 - Channel Mapping

15 - No of Pilot Symbols

measurementResult

Variable Type       ViReal64 (passed by reference)

WCDP measurement result. Format of the results depends on the

selected measurement.

Notes:

(1) If Channel Mapping is selected, then meaning of the result is

following:

0.0 = I

1.0 = Q

(2) If Trigger to Frame is selected, then result is returned as

microseconds.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Measurement) out of range.

BFFC0005  Parameter 5 (Measurement Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_actWCDPMSMode

ViStatus rsfsu_actWCDPMSMode (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function sets the WCDP (Wideband Code Domain Power) MS analyzer

mode.

Note:

(1) This function is available only with FS-K73 3GPP FDD User Equipment

Test option.

Remote-control command(s):

INSTrument<1|2>:SELect MWCDpower

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_appReadSpectrum

ViStatus rsfsu_appReadSpectrum (ViSession instrumentHandle,

ViInt32 measurementMode,

ViInt32 frequencyParameter,

ViReal64 startFrequencyHz,

ViReal64 stopFrequencyHz,

ViInt32 sourceTrace,

ViPInt32 no_ofPoints,

ViReal64 _VI_FAR traceData[]);

Purpose

This function shows how to use instrument driver functions. For

reasonable measurement is necessary to set up instrument to the required

state.

Remote-control command(s):

SENSe<1|2>:FREQuency:STARt

SENSe<1|2>:FREQuency:STOP

INITiate<1|2>:CONTinuous ON |OFF

INITiate<1|2>:DISPlay ON |OFF

SENSe<1|2>:SWEep:COUNt

FORMat ASCII

TRACe<1|2>? TRACE1 | TRACE2 | TRACE3

FORMat REAL,32

*CLS

*TRG

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measurementMode

Variable Type       ViInt32

This control sets the measurement mode.

Valid Range:

0 - Set Up Frequencies (Default Value)

1 - Continual Measurement

2 - Single Measurement

3 - Transfer Data Only

Notes:

(1) Set Up Frequencies:

Allows to set up start and stop frequencies.

(2) Continual Measurement:

Sets the continual sweeping.

(3) Single Measurement:

Sets the single sweeping and immediately sweeps.

(4) Transfer Data Only:

Tranfers data from the instrument without reconfiguring it.

frequencyParameter

Variable Type       ViInt32

Specifies which parameter will be set.

Valid Range:

0 - Start And Stop (Default Value)

1 - Start Only

2 - Stop Only

startFrequencyHz

Variable Type       ViReal64

Sets the start frequency of the analyzer.

Valid Range: 0.0 to 50.0e9 Hz

Default Value: 0.0 Hz

stopFrequencyHz

Variable Type       ViReal64

Sets the stop frequency of the analyzer.

Valid Range: 0.0 to 50.0e9 Hz

Default Value: 3.0E9 Hz

sourceTrace

Variable Type       ViInt32

This control sets the source trace data to be readed from instrument.

Valid Range:

0 - Trace 1 (Default Value)

1 - Trace 2

2 - Trace 3

no_ofPoints

Variable Type       ViInt32 (passed by reference)

Returns the number of trace data points.

traceData

Variable Type       ViReal64[]

Returns the trace data array.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Measurement Mode) out of range.

BFFC0003  Parameter 3 (Frequency Parameter) out of range.

BFFC0004  Parameter 4 (Start Frequency [Hz]) out of range.

BFFC0005  Parameter 5 (Stop Frequency [Hz]) out of range.

BFFC0006  Parameter 6 (Source Trace) out of range.

BFFC0007  Parameter 7 (No. of Points) is null pointer.

BFFC0008  Parameter 8 (Trace Data) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_close

ViStatus rsfsu_close (ViSession instrumentHandle);

Purpose

This function performs the following operations:

viClose (instrSession) and viClose (rmSession).

Notes:

(1) The instrument must be reinitialized to use it again.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_compatibilityChecking

ViStatus rsfsu_compatibilityChecking (ViSession instrumentHandle,

ViBoolean state);

Purpose

This function switches ON/OFF compatibility checking. If this function is

switched to ON, then each not fully compatible FSEx <=> FSU function

returns an error code RSFSU_ERROR_COMPATIBILITY_ERROR (This is due to

using DUMMY parameters which are not any more supported by FSU).

Note:

Switching this function to OFF may cause unexpected behaviour of

driver (instrument).

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

state

Variable Type       ViBoolean

This control selects ON or OFF compatibility checking.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothAdjacentChannels

ViStatus rsfsu_confBToothAdjacentChannels (ViSession instrumentHandle,

ViInt32 number);

Purpose

This function selects the number of adjacent channels if the adjacent

channels meas. is performed.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:ACLR:ACPairs

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

number

Variable Type       ViInt32

Sets the number of adjacent channels.

Valid Range:

1  to 78

Default Value:

78

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Number) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothAveragePower

ViStatus rsfsu_confBToothAveragePower (ViSession instrumentHandle,

ViInt32 start, ViInt32 stop);

Purpose

This function sets the start and stop points for counting the power

average.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:POWer:AVERage:STARt

CONFigure:BTOoth:POWer:AVERage:STOP

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

start

Variable Type       ViInt32

Selects start point.

Valid Range:

0 to 100 %

Default Value:

20%

stop

Variable Type       ViInt32

Selects stop point.

Valid Range:

0 to 100 %

Default Value:

80%

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Start) out of range.

BFFC0003  Parameter 3 (Stop) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothChannel

ViStatus rsfsu_confBToothChannel (ViSession instrumentHandle,

ViInt32 channel, ViInt32 region);

Purpose

This function selects the region and frequency for measurement.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:GEOGraphy EURope | USA | FRANce

CONFigure:BTOoth:CHANnel

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channel

Variable Type       ViInt32

Selects the frequency channel.

Valid Range:

0 to 78

Default Value:

0

Note:

(1) The channel range depends on selected region.

EURope, USA: 0...78

FRANce: 0...22

region

Variable Type       ViInt32

Selects the region.

Valid Range:

0 - Europe (Default Value)

1 - USA

2 - France

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel) out of range.

BFFC0003  Parameter 3 (Region) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothDetector

ViStatus rsfsu_confBToothDetector (ViSession instrumentHandle,

ViInt32 traceNumber,

ViInt32 traceDetector);

Purpose

This function controls the recording of measurement values via the type

of BlueTooth detector selected for each trace.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:DETector<1|2|3> APEak | NEGative | POSitive |

SAMPle | RMS | AVERage

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

traceNumber

Variable Type       ViInt32

Trace number to select detector.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

traceDetector

Variable Type       ViInt32

This command switches the detector for recording of the measured

value.

Valid Range:

2 - Auto Peak (Default Value)

3 - Negative Peak

4 - Positive Peak

5 - Sample

6 - RMS

7 - Average

(1) Auto Peak:

The value "APEak" (AutoPeak) displays both the positive peak value

and the negative peak value when noise is present. The positive peak

value is displayed when one signal is present.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Trace Number) out of range.

BFFC0003  Parameter 3 (Trace Detector) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothOutpPowerClass

ViStatus rsfsu_confBToothOutpPowerClass (ViSession instrumentHandle,

ViInt32 class);

Purpose

This function selects the power class for power output measurement. The

class defines the limits of Output Power measurement.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:PCLass

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

class

Variable Type       ViInt32

Selects the power class.

Valid Range:

1 to 3

Default Value:

1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Class) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothPacketType

ViStatus rsfsu_confBToothPacketType (ViSession instrumentHandle,

ViInt32 type);

Purpose

This function selects the packet type.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:PTYPe DH<1|3|5>

CONFigure:BTOoth:PTYPe AUTO

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

type

Variable Type       ViInt32

Selects the packet type.

Valid Range:

0 - Auto

1 - 1 slot packet (Default Value)

3 - 3 slot packet

5 - 5 slot packet

Note:

(0) Auto: Selects automatic packet type.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Type) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothResBW

ViStatus rsfsu_confBToothResBW (ViSession instrumentHandle,

ViBoolean type, ViReal64 bwValueHz);

Purpose

This function selects the resolution bandwidth for Bluetooth measurement.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:BWIDth:AUTO ON | OFF

CONFigure:BTOoth:BWIDth:RESolution

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

type

Variable Type       ViBoolean

This control selects the automatic and manual setting of resolution

bandwidth.

Valid Range:

VI_FALSE (0) - Manual (Default Value)

VI_TRUE  (1) - Auto

bwValueHz

Variable Type       ViReal64

Sets the resolution bandwidth in manual mode.

Valid Range:

300.0 kHz to 10.0 MHz

Default Value: 1.0e6 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Type) out of range.

BFFC0003  Parameter 3 (Bw Value [Hz]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothSweep

ViStatus rsfsu_confBToothSweep (ViSession instrumentHandle,

ViInt32 sweep, ViInt32 sweepCountValue);

Purpose

This function configures the sweep subsystem of BlueTooth.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

INITiate:CONTinuous ON | OFF

CONFigure:BTOoth:SWEep:COUNt

INITiate:DISPlay ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

sweep

Variable Type       ViInt32

This control selects the sweep parameter.

Valid Range:

0 - Continuous Sweep (Default Value)

1 - Single Sweep

2 - Sweep Count Value

3 - Single Sweep Disp On

4 - Single Sweep Disp Off

Notes:

(1) Continuous Sweep, Single Sweep:

This command determines if the trigger system is continuously

initiated ("Free Run").

(2) Sweep Count Value:

Allows to set up sweep count value.

(3) Single Sweep Disp On/Off:

This command switches the display on or off during a single sweep.

sweepCountValue

Variable Type       ViInt32

This control defines the number of sweeps started with single sweep.

Valid Range: 0 to 32767

Default Value: 0

Note:

(1) This parameter defines the number of sweeps or the number of

averaging procedures. In the average mode, the value 0 defines a

running averaging of measurement data over 10 sweeps.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Sweep) out of range.

BFFC0003  Parameter 3 (Sweep Count Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothSweepTime

ViStatus rsfsu_confBToothSweepTime (ViSession instrumentHandle,

ViBoolean type, ViReal64 sweepTime);

Purpose

This function controls the setting of the analyzer's BlueTooth sweeptime.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:SWEep:TIME:AUTO ON | OFF

CONFigure:BTOoth:SWEep:TIME

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

type

Variable Type       ViBoolean

This control selects the automatic and manual setting of sweep time.

Valid Range:

VI_FALSE (0) - Manual (Default Value)

VI_TRUE  (1) - Auto

sweepTime

Variable Type       ViReal64

Sets manually the sweeptime.

Valid Range: not checked

Default Value: 0.005 sec

Note:

(1) This value is not range checked

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Type) out of range.

BFFC0003  Parameter 3 (Sweep Time) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothTrace

ViStatus rsfsu_confBToothTrace (ViSession instrumentHandle,

ViInt32 traceNumber, ViInt32 traceMode);

Purpose

This function defines the type of display and the evaluation of the

traces.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:TRACe<1|2|3>:MODE WRITe | VIEW | AVERage | MAXHold |

MINHold | BLANk

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

traceNumber

Variable Type       ViInt32

Selects the trace to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

traceMode

Variable Type       ViInt32

This control defines the type of display and the evaluation of the

traces.

Valid Range:

0 - Clear/Write (Default Value)

1 - View

2 - Average

3 - Blank

4 - Max Hold

5 - Min Hold

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Trace Number) out of range.

BFFC0003  Parameter 3 (Trace Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothTraceSelect

ViStatus rsfsu_confBToothTraceSelect (ViSession instrumentHandle,

ViInt32 traceNumber);

Purpose

This function defines the type of display and the evaluation of the

traces.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:TRACe<1|2|3>:SELect

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

traceNumber

Variable Type       ViInt32

Selects the trace to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Trace Number) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBToothViBW

ViStatus rsfsu_confBToothViBW (ViSession instrumentHandle,

ViBoolean type, ViReal64 bwValueHz);

Purpose

This function controls the setting of the analyzer's BlueTooth video

filter bandwidth.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:BWIDth:VIDeo:AUTO ON | OFF

CONFigure:BTOoth:BWIDth:VIDeo

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

type

Variable Type       ViBoolean

This control selects the automatic and manual setting of video

bandwidth.

Valid Range:

VI_FALSE (0) - Manual (Default Value)

VI_TRUE  (1) - Auto

bwValueHz

Variable Type       ViReal64

Sets the video bandwidth in manual mode.

Valid Range: 1.0 Hz to 10.0E+6 Hz

Default Value: 1.0E+6 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Type) out of range.

BFFC0003  Parameter 3 (Bw Value [Hz]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBTSamplesPerSymbol

ViStatus rsfsu_confBTSamplesPerSymbol (ViSession instrumentHandle,

ViInt32 samplesPerSymbol);

Purpose

This function determines the number of samples per symbol (points per

symbol).

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Remote-control command(s):

CONFigure:BTOoth:PRATe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

samplesPerSymbol

Variable Type       ViInt32

This control selects the number of samples per symbol.

Valid Range:

2, 4, 8, 16, 32

Default Value:

4

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Samples Per Symbol) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confBurstPwrSeq

ViStatus rsfsu_confBurstPwrSeq (ViSession instrumentHandle,

ViInt32 window,

ViReal64 analyzerFrequencyHz,

ViReal64 resolutionBandwidthHz,

ViReal64 measTimes,

ViInt32 triggerSource,

ViReal64 triggerLevel,

ViReal64 triggerOffsets,

ViInt32 type_ofMeas,

ViInt32 no_ofMeasurements);

Purpose

This function configures the instrument setup for multiple burst power

measurement and starts a measurement sequence. To reduce the setting

time, the setup is performed simultaneously for all selected parameters.

Remote-control command(s):

SENSe<1|2>:MPOWer:SEQuence

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

analyzerFrequencyHz

Variable Type       ViReal64

Receive frequency for the burst signals to be measured (= center

frequency in manual operation).

Valid Range: 0.0 Hz to max. analyzer frequency, depending on

instrument model

Default Value: 0.0 Hz

resolutionBandwidthHz

Variable Type       ViReal64

Resolution bandwidth for the measurement.

Valid Range: 10.0 Hz to 10.0e6 Hz in steps of 1, 3, 10

Default Value: 10.0e6 Hz

measTimes

Variable Type       ViReal64

Time span during which measurement samples are sampled for RMS / peak

measurement The type of measurement is selected by Type of Meas.

Valid Range: 1.0e-6 s to 30.0 s

Default Value: 1.0e-6

triggerSource

Variable Type       ViInt32

Trigger signal source.

Valid Range:

0 - External (Default Value)

1 - Video

Notes:

(1) External:

The trigger signal is fed from the "Ext. Trigger/Gate" input on the

rear of the unit.

(2) Video:

The internal video signal is used as trigger signal.

triggerLevel

Variable Type       ViReal64

Signal level of the video signal as a percentage of the diagram

height at which the trigger becomes active. If Trigger Source =

External is selected, the value entered here is ignored, as in this

case the trigger input uses TTL levels.

Valid Range: 0.0 % to 100.0 %

Default Value: 50.0 %

triggerOffsets

Variable Type       ViReal64

Offset between the detection of the trigger signal and the start of

the measurement.

Valid Range: 125.0e-9 s to 100.0 s

Default Value: 125.0e-6 s

type_ofMeas

Variable Type       ViInt32

Determines whether mean power (RMS) or peak power (PEAK) is to be

measured. The FSU detector is selected accordingly.

Valid Range:

0 - MEAN (Default Value)

1 - PEAK

no_ofMeasurements

Variable Type       ViInt32

Number of individual bursts to be measured.

Valid Range: 1 to 501

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Analyzer Frequency [Hz]) out of range.

BFFC0004  Parameter 4 (Resolution Bandwidth [Hz]) out of range.

BFFC0005  Parameter 5 (Meas Time [s]) out of range.

BFFC0006  Parameter 6 (Trigger Source) out of range.

BFFC0007  Parameter 7 (Trigger Level [%]) out of range.

BFFC0008  Parameter 8 (Trigger Offset [s]) out of range.

BFFC0009  Parameter 9 (Type of Meas) out of range.

BFFC000A  Parameter 10 (No of Measurements) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kBandClass

ViStatus rsfsu_confC2kBandClass (ViSession instrumentHandle,

ViInt32 bandClass);

Purpose

This function selects the band class.

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:BCLass

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

bandClass

Variable Type       ViInt32

This control selects the band class.

Valid Range:

0 - 800 MHz band

1 - 1900 MHz band

2 - TACS band

3 - JTACS band

4 - Korean PCS band

5 - 450 MHz band

6 - 2 GHz band

7 - 700 MHz band

8 - 1800 MHz band

9 - 900 MHz band

10 - Secondary 800 MHz band

11 - 400 MHz European PAMR band

12 - 800 MHz PAMR band

Default Value: 0 - 800 MHz band

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Band Class) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPChTable

ViStatus rsfsu_confC2kCDPChTable (ViSession instrumentHandle,

ViBoolean channelTable);

Purpose

This function switches the channel table on or off in the code domain

power measurement mode.

On switching on, the measured channel table is stored under the name

RECENT and switched on. After the RECENT channel table is switched on,

another channel table can be selected with the function CDP C2k Channel

Table File (confC2kCDPChTableFile). The RECENT channel table must always

be switched on first with this function and then the required channel

table can be selected with the function CDP C2k Channel Table File

(confC2kCDPChTableFile).

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:CTABle:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTable

Variable Type       ViBoolean

This control switches the channel table on or off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPChTableCatalog

ViStatus rsfsu_confC2kCDPChTableCatalog (ViSession instrumentHandle,

ViPInt32 number_ofChannelTables,

ViInt32 bufferSize,

ViChar _VI_FAR channelTablesList[]);

Purpose

This function reads out the names of all channel tables stored on the

harddisk in the code domain power measurement mode.

Notes:

(1) Driver returns an instrument specific error, when catalog list is

empty.

(2) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:CTABle:CATalog?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

number_ofChannelTables

Variable Type       ViInt32 (passed by reference)

This control returns number of channel tables, stored on harddisk.

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the string you specify for the catalog

list parameter.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the value contains more

bytes than you specify with the Buffer Size parameter, the function

copies Buffer Size - 1 bytes into the buffer and returns the buffer

size you must pass to get the entire value. For example, if the value

is "file1,file2" and the Buffer Size is 6, the function places

"file1" into the buffer and returns 12.

channelTablesList

Variable Type       ViChar[]

This control returns array of channel tables, stored on harddisk.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the value contains more

bytes than you specify with the Buffer Size parameter, the function

copies Buffer Size - 1 bytes into the buffer and returns the buffer

size you must pass to get the entire value. For example, if the value

is "file1,file2" and the Buffer Size is 6, the function places

"file1" into the buffer and returns 12.

if you pass 0 for buffer Size, you can use VI_NULL for this

parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Number of Channel Tables) is null pointer.

BFFC0003  Parameter 3 (Buffer Size) out of range.

BFFC0004  Parameter 4 (Channel Tables List) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPChTableComment

ViStatus rsfsu_confC2kCDPChTableComment (ViSession instrumentHandle,

ViString channelTableComment);

Purpose

This function defines a comment for the selected channel table in the

code domain power measurement mode.

Prior to this function, the name of the channel table has to be defined

with function CDP C2k Channel Table Name (confC2kCDPChTableName) and the

values of the table have to be defined with function CDP C2k Channel

Table Data (confC2kCDPChTableData).

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:CTABle:COMMent

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableComment

Variable Type       ViString

This control defines a comment for the selected channel table.

Valid Range: up to 128 characters

Default Value: "Comment for table 1"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Comment) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPChTableCopy

ViStatus rsfsu_confC2kCDPChTableCopy (ViSession instrumentHandle,

ViString channelTableName);

Purpose

This function copies one channel table onto another one in the code

domain power measurement mode.

The channel table to be copied is selected with function CDP C2k Channel

Table Name (confC2kCDPChTableName).

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:CTABle:COPY

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableName

Variable Type       ViString

This control defines channel table onto which is copied channel table

selected with the function CDP C2k Channel Table Name

(confC2kCDPChTableName).

Valid Range: up to 8 characters

Default Value: "CTAB_2"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPChTableData

ViStatus rsfsu_confC2kCDPChTableData (ViSession instrumentHandle,

ViInt32 _VI_FAR channelType[],

ViInt32 _VI_FAR codeClass[],

ViInt32 _VI_FAR codeNumber[],

ViInt32 _VI_FAR radioConfig_Mapping[],

ViInt32 _VI_FAR activity[],

ViInt32 _VI_FAR reserved2[],

ViInt32 _VI_FAR status[],

ViReal64 _VI_FAR CDPRelative[],

ViInt32 num_ofElements);

Purpose

This function defines the values of the selected channel table in the

code domain power measurement mode. The whole table is defined in one go.

The inactive channels (INACtive) do not have to be defined. Eight values

are specified for a line of table.

Prior to this function, the name of the channel table has to be defined

with function CDP C2k Channel Table Name (confC2kCDPChTableName).

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:CTABle:DATA

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelType

Variable Type       ViInt32[]

Sets the channel type.

Valid Range: option dependent

Default Value: none

Note:

The channel type is numerically coded as follows:

Option FS-K82, FS-K83:

0 = PICH

1 = SYNC

2 = PCH

3 = TDPICH

4 = APICH

5 = ATDPICH

6 = BCH

7 = CPCCH

8 = CACH

9 = CCCH

10 = CHAN

11 = INACTIVE

Option FS-K84:

0 = PILOT

1 = MAC

2 = PREAMBLE with 64 chip length

3 = PREAMBLE with 128 chip length

4 = PREAMBLE with 256 chip length

5 = PREAMBLE with 512 chip length

6 = PREAMBLE with 1024 chip length

7 = DATA

Option FS-K85:

0 = PICH

1 = RRI

2 = DATA

3 = ACK

4 = DRC

5 = INACTIVE

codeClass

Variable Type       ViInt32[]

Sets code class.

Valid Range:

FS-K82: 2 to 7

FS-K83: 1 to 6

FS-K84: Depending on channel type, fixed: PILOT: 5, MAC: 6,

PREAMBLE: 5 and DATA: 4 (spreading factor = 2^code

class)

FS-K85: 2 to 4

Default Value: none

codeNumber

Variable Type       ViInt32[]

Sets code number.

Valid Range:

FS-K82: 0 to 127

FS-K83: 0 to 63

FS-K84: 0 to spreading factor - 1

FS-K85: 0 to 15

Default Value: none

radioConfig_Mapping

Variable Type       ViInt32[]

- For option FS-K82 specifies the radio configuration used.

Valid Range: 0 to 5

Note:

0 for all special channels; for the others (CHAN), the values 1, 2,

3, 4 and 5 are allowed.

- For option FS-K83, FS-K85 is set mapping.

Valid Range:

0: I branch

1: Q branch

- For option FS-K84 is set modulation type including mapping.

Valid Range:

0 = BPSK-I

1 = BPSK-Q

2 = QPSK

3 = 8-PSK

4 = 16-QAM

Modulation types QPSK/8-PSK/16-QAM have complex values.

Default Value: none

activity

Variable Type       ViInt32[]

Sets channel activity for FS-K85 option settings. The decimal number

interpreted as a binary number in 16 bits, determines the half slot

in which the channel is active (value 1) or inactive (value 0).

Default Valid Range: 0 to 65535

Default Value: none

Examples:

65535 dec. = 1111 1111 1111 1111 bin. (e.g. DATA)

Channel is active in each half slot

43690 dec. = 1010 1010 1010 1010 bin. (e.g. RRI)

Channel is active in half slot 0, 2, 4 etc

24576 dec. = 0110 0000 0000 0000 bin. (e.g. DRC)

Channel is active in half slot 1 and 2

reserved2

Variable Type       ViInt32[]

Reserved for future use.

status

Variable Type       ViInt32[]

Status setting. Can be used in a setting command to disable a channel

temporarily.

Default Valid Range:

0: not active

1: active

Default Value: none

CDPRelative

Variable Type       ViReal64[]

Sets CDP relative.

Valid Range: any value

Default Value: none

Note:

Tis parameter is not used for option FS-K85 settings.

num_ofElements

Variable Type       ViInt32

Sets number of elements for parameter arrays.

Valid Range: >= 1

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Type) out of range.

BFFC0003  Parameter 3 (Code Class) out of range.

BFFC0004  Parameter 4 (Code Number) out of range.

BFFC0005  Parameter 5 (Radio Config (Mapping)) out of range.

BFFC0006  Parameter 6 (Activity) out of range.

BFFC0007  Parameter 7 (Reserved 2) out of range.

BFFC0008  Parameter 8 (Status) out of range.

BFFC0009  Parameter 9 (CDP Relative) out of range.

BFFC000A  Parameter 10 (Num of Elements) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPChTableDelete

ViStatus rsfsu_confC2kCDPChTableDelete (ViSession instrumentHandle);

Purpose

This function deletes the selected channel table in the code domain power

measurement mode.

The channel table to be deleted is selected with function CDP C2k Channel

Table Name (confC2kCDPChTableName).

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:CTABle:DELete

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPChTableFile

ViStatus rsfsu_confC2kCDPChTableFile (ViSession instrumentHandle,

ViString channelTableFile);

Purpose

This function selects a predefined channel table file in the code domain

power measurement mode.

Prior to use this function, the RECENT channel table must be switched on

first with the function CDP C2k Channel Table (confC2kCDPChTable).

Note:

(1) This function is available only in conjunction with option FS-K82

,FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:CTABle:SELect

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableFile

Variable Type       ViString

This control selects a predefined channel table file.

Valid Range: up to 8 characters

Default Value: "RECENT"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table File) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPChTableName

ViStatus rsfsu_confC2kCDPChTableName (ViSession instrumentHandle,

ViString channelTableName);

Purpose

This function selects an existing channel table or creates the name of a

new channel table in the code domain power measurement mode. It is not

used for the analysis. Refer to the commands CONF:CDP:CTAB:STAT and

CONF:CDP:CTAB:SEL.

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:CTABle:NAME

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableName

Variable Type       ViString

This control selects an existing channel table or creates the name of

a new channel table.

Valid Range: up to 8 characters

Default Value: "NEW_TAB"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPChTableRestore

ViStatus rsfsu_confC2kCDPChTableRestore (ViSession instrumentHandle);

Purpose

This function restores the predefined tables. They are restored to the

values they had when the device was delivered. With this restore feature

an undo of overwriting these channel tables is possible.

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:CTABle:RESTore

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPIQLength

ViStatus rsfsu_confC2kCDPIQLength (ViSession instrumentHandle,

ViInt32 IQLength);

Purpose

This function sets the capture length in multiples of the power control

group.

Note:

(1) This function is used only for compatibility with older versions of

the driver. See new function rsfsu_confCDPIQLength.

Remote-control command(s):

[CONFigure:]CDPower:IQLength

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

IQLength

Variable Type       ViInt32

This control sets the capture length in multiples of the power

control group.

Valid Range: >= 2

Default Value: 3

Note:

The range might vary depending on analyzer model.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (IQ Length) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPMeas

ViStatus rsfsu_confC2kCDPMeas (ViSession instrumentHandle,

ViInt32 CDPMeasurement);

Purpose

This function selects the CDMA2000 / 1xEV-DO base (mobile) station

measurement tests (power, ACLR, occupied bandwidth, code domain power,

emission spectrum mask, complementary cumulative distribution function or

power versus time measurement).

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CONFigure:CDPower:MEASurement POWer | ACLR | ESPectrum | OBANdwith |

OBWidth | CDPower | CCDF | PVTime

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

CDPMeasurement

Variable Type       ViInt32

This control selects CDP Measurement mode.

Valid Range:

0 - POWer (Default Value)

1 - ACLR

2 - ESPectrum

3 - OBWidth

4 - CDPower

5 - CCDF

6 - PVTime

Notes:

(1) POWer:

Power channel measurement with predefined settings

(2) ACLR:

Adjacent channel power measurements with predefined settings.

(3) ESPectrum:

Check on signal power (Spectrum Emission Mask).

(4) OBWidth:

Measurement of occupied bandwidth.

(5) CDPower:

Code domain analyzer measurement.

(6) CCDF:

Measurement of the Complementary Cumulative Distribution

function (signal statistics measurement).

(7) PVTime:

Measurement of power versus time, here trace readout delivers 512

points. Available only with option FS-K84.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (CDP Measurement) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPMeasMode

ViStatus rsfsu_confC2kCDPMeasMode (ViSession instrumentHandle,

ViInt32 measurementMode);

Purpose

This function selects the CDP measurement mode (code domain power, code

domain power ratio, CDP error summary, CDP modulation accuracy, CDP power

versus slot, CDP power versus chip, CDP bit stream, CDP channel

assignment table, peak code domain error, symbol constellation, symbol

error vector magnitude, code power overview).

Note:

(1) This function is available only in conjunction with option FS-K76,

FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CALCulate1:FEED 'XPOW:CDP'

CALCulate1:FEED 'XPOW:CDP:RAT'

CALCulate1:FEED 'XPOW:CDEP'

CALCulate2:FEED 'XTIM:CDP:MACC'

CALCulate2:FEED 'XTIM:CDP:PVSL'

CALCulate2:FEED 'XTIM:CDP:PVSLot:RAT'

CALCulate2:FEED 'XTIM:CDP:PVCHip'

CALCulate2:FEED 'XTIM:CDP:PVSY'

CALCulate2:FEED 'XTIM:CDP:BSTR'

CALCulate2:FEED 'XTIM:CDP:ERR:SUMM'

CALCulate1:FEED 'XTIM:CDP:ERR:CTAB'

CALCulate2:FEED 'XTIM:CDP:ERR:PCD'

CALCulate2:FEED 'XTIM:CDP:SYMB:CONS'

CALCulate2:FEED 'XTIM:CDP:SYMB:EVM'

CALCulate2:FEED 'XTIM:CDP:COMP:CONS'

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measurementMode

Variable Type       ViInt32

This control selects the CDP measurement mode.

Valid Range:

0 - Code Domain Power Absolute (Default Value)

1 - Code Domain Power Ratio

2 - Code Domain Power

3 - CDP Modulation Accuracy

4 - Power vs Power control group

5 - Power vs Symbol

6 - CDP Bit Stream

7 - CDP Error Summary

8 - CDP Channel Assignment Table

9 - Peak Code Domain Error

10 - Symbol Constellation

11 - Symbol Error Vector Magnitude

12 - Composite Constellation

13 - CDP Power Versus Chip (FS-K84 only)

14 - CDP Power Versus Slot Rrelative (FS-K76, FS-K77 only)

Notes:

(1) Code Domain Power Absolute:

Result display of code domain power (absolute) as bargraph.

(2) Code Domain Power Ratio:

Result display of code domain power ratio.

(3) Code Domain Power:

Result display of code domain power as bargraph.

(4) CDP Modulation Accuracy:

Result display of modulation accuracy.

(5) CDP Power Versus Slot:

Result display of power versus slot.

(6) CDP Power Versus Symbol:

Result display of power versus symbol.

(7) CDP Bit Stream:

Result display of bit stream.

(8) CDP Error Summary:

Result display in tabular form.

(9) CDP Channel Assignment Table:

Result disp of channel assignment table.

(10) Peak Code Domain Error:

Result display of peak code domain error.

(11) Symbol Constellation:

Result disp of symbol constellation.

(12) Symbol Error Vector Magnitude:

Result disp of symbol error vector magnitude.

(13) Composite Constellation:

Result display of composite constellation.

(14) CDP Power Versus Chip:

Result display of power versus chip.

(15) CDP Power Versus Slot Relative:

Result display of power versus slot, relative.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Measurement Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPOrder

ViStatus rsfsu_confC2kCDPOrder (ViSession instrumentHandle,

ViInt32 order);

Purpose

This function sets the order of the code domain evaluation. The codes are

sorted either in Hadamard order or in BitReverse order.

Note:

(1) This function is used only for compatibility with older versions of

the driver. See new function rsfsu_confCDPOrder.

Remote-control command(s):

[CONFigure:]CDPower:ORDer HADamard | BITReverse

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

order

Variable Type       ViInt32

This control sets the order of the code domain evaluation.

Valid Values:

0 - Hadamard (Default Value)

1 - Bit Reverse

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Order) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPPNOffset

ViStatus rsfsu_confC2kCDPPNOffset (ViSession instrumentHandle,

ViInt32 PNOffset);

Purpose

This function sets the PN offset of the base station in multiples of 64

chips.

Note:

(1) This function is used only for compatibility with older versions of

the driver. See new function rsfsu_confCDPPNOffset.

Remote-control command(s):

[CONFigure:]CDPower:PNOFfset

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

PNOffset

Variable Type       ViInt32

This control sets the PN offset of the base station in multiples of

64 chips.

Valid Range: 0 to 511

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (PN Offset) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confC2kCDPTPM

ViStatus rsfsu_confC2kCDPTPM (ViSession instrumentHandle,

ViBoolean offsetEvaluation);

Purpose

This function supports the specific enabling and disabling of the timing

and phase offset analysis of the channels for the first active channel in

the channel type.

Note:

(1) This function is used only for compatibility with older versions of

the driver. See new function rsfsu_confCDPTPM.

Remote-control command(s):

[CONFigure:]CDPower:TPMeas ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

offsetEvaluation

Variable Type       ViBoolean

This control switches on or off the timing and phase offset

evaluation of the channels for the pilot channel.

Valid Values:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Offset Evaluation) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPAnalysisBase

ViStatus rsfsu_confCDPAnalysisBase (ViSession instrumentHandle,

ViInt32 analysisBase);

Purpose

This function chooses the base of the CDP analysis: At SLOT one slot of

the signal is analyzed only; at FRAME the complete 3GPP frame will be

analyzes.

Note:

(1) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

[SENSe<1|2>:]CDPower:BASE SLOT | FRAMe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

analysisBase

Variable Type       ViInt32

This control chooses the base of the CDP analysis.

Valid Range:

0 - SLOT (Default Value)

1 - FRAME

Notes:

(1) SLOT:

At SLOT one slot of the signal is analyzed only.

(2) FRAME:

At FRAME the complete 3GPP frame will be analyzes.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Analysis Base) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPAntennaType

ViStatus rsfsu_confCDPAntennaType (ViSession instrumentHandle,

ViInt32 antennaMode);

Purpose

This function switches on or off the antenna diversity mode and selects

the antenna of the diversity path in the code domain power measurement

mode.

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K73 or FS-K82.

Remote-control command(s):

[CONFigure:]CDPower:ANTenna OFF | 1 | 2

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

antennaMode

Variable Type       ViInt32

This control sets antenna mode.

Valid Range:

0 - Off (Default Value)

1 - 1

2 - 2

Notes:

(1) OFF:

Antenna diversity mode switched off.

(2) 1, 2:

According to 3 GPP, the pilot symbols are different for each channel

depending on the used antenna. The CPICH symbols vary and the

synchronization channels are transmitted alternately in each second

slot.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Antenna Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPAveraging

ViStatus rsfsu_confCDPAveraging (ViSession instrumentHandle,

ViBoolean averaging);

Purpose

This function is used to enable averaging of the CDP evaluation over all

recorded (half) slots. This function is only available in the CDP

measurement.

Note:

(1) This function is available only in conjunction with option FS-K84 or

FS-K85.

Remote-control command(s):

[SENSe:]CDPower:AVERage ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

averaging

Variable Type       ViBoolean

This control enables averaging of the CDP evaluation over all

recorded (half) slots.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Averaging) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChannelType

ViStatus rsfsu_confCDPChannelType (ViSession instrumentHandle,

ViInt32 channelType);

Purpose

This function is used to select the channel type. The number of results

then changes in most analyses, such as code domain power, symbol EVM, and

bit stream, because either a different spreading factor or a different

number of symbols is available for the analysis.

Note:

(1) This function is available only in conjunction with option FS-K84.

Remote-control command(s):

[SENSe:]CDPower:CTYPe PILot | MAC | PREamble | DATA

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelType

Variable Type       ViInt32

This control select the channel type.

Valid Range:

0 - Pilot

1 - MAC

2 - Preamble

3 - DATA

Default Value: 0 - Pilot

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Type) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChTable

ViStatus rsfsu_confCDPChTable (ViSession instrumentHandle,

ViBoolean channelTable);

Purpose

This function switches the channel table on or off in the code domain

power measurement mode.

On switching on, the measured channel table is stored under the name

RECENT and switched on. After the RECENT channel table is switched on,

another channel table can be selected with the function CDP Channel Table

File (confCDPChTableFile). The RECENT channel table must always be

switched on first with this function and then the required channel table

can be selected with the function CDP Channel Table File

(confCDPChTableFile).

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:CTABle:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTable

Variable Type       ViBoolean

This control switches the channel table on or off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChTableCatalog

ViStatus rsfsu_confCDPChTableCatalog (ViSession instrumentHandle,

ViPInt32 number_ofChannelTables,

ViInt32 bufferSize,

ViChar _VI_FAR channelTablesList[]);

Purpose

This function reads out the names of all channel tables stored on the

harddisk in the code domain power measurement mode.

Notes:

(1) Driver returns an instrument specific error, when catalog list is

empty.

(2) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:CTABle:CATalog?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

number_ofChannelTables

Variable Type       ViInt32 (passed by reference)

This control returns number of channel tables, stored on harddisk.

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the string you specify for the catalog

list parameter.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the value contains more

bytes than you specify with the Buffer Size parameter, the function

copies Buffer Size - 1 bytes into the buffer and returns the buffer

size you must pass to get the entire value. For example, if the value

is "file1,file2" and the Buffer Size is 6, the function places

"file1" into the buffer and returns 12.

channelTablesList

Variable Type       ViChar[]

This control returns array of channel tables, stored on harddisk.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the value contains more

bytes than you specify with the Buffer Size parameter, the function

copies Buffer Size - 1 bytes into the buffer and returns the buffer

size you must pass to get the entire value. For example, if the value

is "file1,file2" and the Buffer Size is 6, the function places

"file1" into the buffer and returns 12.

if you pass 0 for buffer Size, you can use VI_NULL for this

parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Number of Channel Tables) is null pointer.

BFFC0003  Parameter 3 (Buffer Size) out of range.

BFFC0004  Parameter 4 (Channel Tables List) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChTableComment

ViStatus rsfsu_confCDPChTableComment (ViSession instrumentHandle,

ViString channelTableComment);

Purpose

This function defines a comment for the selected channel table in the

code domain power measurement mode.

Prior to this function, the name of the channel table has to be defined

with function CDP Channel Table Name (confCDPChTableName) and the values

of the table have to be defined with function CDP Channel Table Data

(confCDPChTableData).

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:CTABle:COMMent

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableComment

Variable Type       ViString

This control defines a comment for the selected channel table.

Valid Range: up to 128 characters

Default Value: "Comment for table 1"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Comment) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChTableCopy

ViStatus rsfsu_confCDPChTableCopy (ViSession instrumentHandle,

ViString channelTableName);

Purpose

This function copies one channel table onto another one in the code

domain power measurement mode.

The channel table to be copied is selected with function CDP Channel

Table Name (confCDPChTableName).

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:CTABle:COPY

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableName

Variable Type       ViString

This control defines channel table onto which is copied channel table

selected with the function CDP Channel Table Name

(confCDPChTableName).

Valid Range: up to 8 characters

Default Value: "CTAB_2"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChTableData

ViStatus rsfsu_confCDPChTableData (ViSession instrumentHandle,

ViInt32 _VI_FAR channelType[],

ViInt32 _VI_FAR codeClass[],

ViInt32 _VI_FAR codeNumber[],

ViInt32 _VI_FAR modulationType[],

ViInt32 _VI_FAR midambleShift[],

ViInt32 _VI_FAR status[],

ViInt32 _VI_FAR reserved1[],

ViReal64 _VI_FAR reserved2[],

ViInt32 num_ofElements);

Purpose

This function defines a channel table. The whole table is defined in one

operation. The inactive channels (INACtive) do not have to be defined.

Eight values are specified for a line of a table.

Prior to this function, the name of the channel table has to be defined

with function CDP Channel Table Name (confCDPChTableName).

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:CTABle:DATA

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelType

Variable Type       ViInt32[]

Sets the channel type.

Valid Range: option dependent

Default Value: none

Note:

The channel type is numerically coded as follows:

Option FS-K76:

1 = Midamble

2 = DPCH

3 = P-CCPCH

4 = S-CCPCH

5 = FPACH

6 = PRACH

Option FS-K77:

1 = Midamble

2 = DPCH

8 = PRACH

9 = PUSCH

codeClass

Variable Type       ViInt32[]

Sets code class.

Valid Range: 0 to 4

Default Value: none

codeNumber

Variable Type       ViInt32[]

Sets code number.

Valid Range: 0 to 16

Default Value: none

modulationType

Variable Type       ViInt32[]

Modulation type.

Valid Range:

0 = invalid (for midamble)

1 = QPSK

2 = 8PSK

Default Value: none

midambleShift

Variable Type       ViInt32[]

Sets midamble shift.

Valid Range: 1 to 16

Default Value: none

status

Variable Type       ViInt32[]

Can be used to disable a channel.

Valid Range:

0: inactive

1: active

Default Value: none

reserved1

Variable Type       ViInt32[]

Reserved for future use.

reserved2

Variable Type       ViReal64[]

Reserved for future use.

num_ofElements

Variable Type       ViInt32

Sets number of elements for parameter arrays.

Valid Range: >= 1

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Type) out of range.

BFFC0003  Parameter 3 (Code Class) out of range.

BFFC0004  Parameter 4 (Code Number) out of range.

BFFC0005  Parameter 5 (Modulation Type) out of range.

BFFC0006  Parameter 6 (Midamble Shift) out of range.

BFFC0007  Parameter 7 (Status) out of range.

BFFC0008  Parameter 8 (Reserved 1) out of range.

BFFC0009  Parameter 9 (Reserved 2) out of range.

BFFC000A  Parameter 10 (Num of Elements) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChTableDelete

ViStatus rsfsu_confCDPChTableDelete (ViSession instrumentHandle);

Purpose

This function deletes the selected channel table in the code domain power

measurement mode.

The channel table to be deleted is selected with function CDP Channel

Table Name (confCDPChTableName).

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:CTABle:DELete

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChTableFile

ViStatus rsfsu_confCDPChTableFile (ViSession instrumentHandle,

ViString channelTableFile);

Purpose

This function selects a predefined channel table file in the code domain

power measurement mode.

Prior to use this function, the RECENT channel table must be switched on

first with the function CDP Channel Table (confCDPChTable).

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:CTABle:SELect

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableFile

Variable Type       ViString

This control selects a predefined channel table file.

Valid Range: up to 8 characters

Default Value: "RECENT"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table File) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChTableMidambleShift

ViStatus rsfsu_confCDPChTableMidambleShift (ViSession instrumentHandle,

ViInt32 midambleShift);

Purpose

This function defines the maximum number of midamble shifts in the

channel table.

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:CTABle:MSHift 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

midambleShift

Variable Type       ViInt32

This control defines the maximum number of midamble shifts in the

channel table.

Valid Range: 2, 4, 6, 8, 10, 12, 14, 16

Default Value: 16

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Midamble Shift) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChTableName

ViStatus rsfsu_confCDPChTableName (ViSession instrumentHandle,

ViString channelTableName);

Purpose

This function selects an existing channel table or creates the name of a

new channel table in the code domain power measurement mode. It is not

used for the analysis. Refer to the commands CONF:CDP:CTAB:STAT and

CONF:CDP:CTAB:SEL.

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:CTABle:NAME

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableName

Variable Type       ViString

This control selects an existing channel table or creates the name of

a new channel table.

Valid Range: up to 8 characters

Default Value: "NEW_TAB"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPChTableOrder

ViStatus rsfsu_confCDPChTableOrder (ViSession instrumentHandle,

ViInt32 order);

Purpose

This function selects sorting of the channel table in code order or

midamble order.

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:CTABle:ORDer CODE | MIDamble

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

order

Variable Type       ViInt32

This control selects sorting of the channel table in code order or

midamble order.

Valid Values:

0 - Code (Default Value)

1 - Midamble

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Order) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPCodeNum

ViStatus rsfsu_confCDPCodeNum (ViSession instrumentHandle,

ViInt32 codeNumber);

Purpose

This function sets the code number in code domain power measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K76, FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:CODE

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

codeNumber

Variable Type       ViInt32

This control sets the code number.

Valid Range:

Options FS-K72 or FS-K73: 0 to 511

Options FS-K76 or FS-K77: 1 to 16

Option FS-K82: 0 to 127

Option FS-K83: 0 to 63

Option FS-K85: 0 to 15

Option FS-K84:

PILOT/PREAMBLE: 0 to 31

DATA: 0 to 15

MAC: 0 to 63

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Code Number) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPConstellationParameterB

ViStatus rsfsu_confCDPConstellationParameterB

(ViSession instrumentHandle, ViInt32 parameterB);

Purpose

This function sets the constellation parameter B. This parameter can be

adjusted to decide which bit mapping should be used for bit-stream

evaluation.

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K74.

Remote-control command(s):

SENSe1:CDPower:CPB

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

parameterB

Variable Type       ViInt32

This control sets the constellation parameter B.

Valid Range:

0 to 3

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Parameter B) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPEliminateTailChips

ViStatus rsfsu_confCDPEliminateTailChips (ViSession instrumentHandle,

ViBoolean eliminateTailChips);

Purpose

This function selects length of the measurement interval for

calculation of error vector magnitude. According to 3GPP specification

Release 5 the EVM measurement interval is one slot (4096 chips) less 25

us at each end of the burst (3904 chips) if power changes are expected.

If no power changes are expected, the evaluation length is one slot (4096

chips).

Note:

(1) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

[SENSe:]CDPower:ETCHips ON|OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

eliminateTailChips

Variable Type       ViBoolean

Selects length of the measurement interval for calculation of error

vector magnitude

Valid Values:

VI_TRUE  (1) - On

VI_FALSE (0) - Off (Default Value)

Notes:

(1) On:

Changes of power are expected. Therefore an EVM measurement interval

of one slot less 25 �s (3904 chips) is considered.

(2) Off:

Changes of power are not expected. Therefore an EVM measurement

interval of one slot (4096 chips) is considered.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Eliminate Tail Chips) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPFrameAnalyze

ViStatus rsfsu_confCDPFrameAnalyze (ViSession instrumentHandle,

ViInt32 frameValue);

Purpose

This function defines FRAME TO ANALYZE, the data range which want to be

analyzed can be selected. The range of this parameter depends on the

CAPURE LENGTH.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73 or FS-K74.

Remote-control command(s):

[SENSe:]CDPower:FRAMe[:VALue]

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

frameValue

Variable Type       ViInt32

This control defines the frame value.

Valid Range: 0 to CAPTURE_LENGTH - 1

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Frame Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPICT

ViStatus rsfsu_confCDPICT (ViSession instrumentHandle,

ViReal64 tresholdValuedB);

Purpose

This function sets the threshold value in the CDPR diagram from which a

channel is treated as active in the code domain power measurement mode.

The level entered refers to the total signal power.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K76, FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CDPower:ICTReshold

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

tresholdValuedB

Variable Type       ViReal64

This control sets channel treshold value.

Valid Range:

Option FS-K72 or FS-K73: -50.0 dB to 10.0 dB

Option FS-K76 or FS-K77: -100.0 to 0 dB

Option FS-K82 or FS-K83: -100.0 to 0 dB

Default Value: -10.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Treshold Value [dB]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPIQLength

ViStatus rsfsu_confCDPIQLength (ViSession instrumentHandle,

ViInt32 IQLength);

Purpose

This function sets the capture length in multiples of the power control

group.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K74, FS-K76, FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:IQLength

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

IQLength

Variable Type       ViInt32

This control sets the capture length in multiples of the power

control group.

Valid Range: >= 1

Default Value: 3

Note:

The range might vary depending on analyzer model.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (IQ Length) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPLCode

ViStatus rsfsu_confCDPLCode (ViSession instrumentHandle,

ViInt32 scramblingCode);

Purpose

This function defines the scrambling code in the code domain power

measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K76 or FS-K77.

Remote-control command(s):

[SENSe:]CDPower:LCODe

[SENSe:]CDPower:SCODe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

scramblingCode

Variable Type       ViInt32

This control defines the scrambling code.

Valid Range:

Option FS-K73:         >= 0

Option FS-K72, FS-K74: 0 to 24575

Option FS-K76, FS-K77: 0 to 127

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Scrambling Code) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPLCodeMask

ViStatus rsfsu_confCDPLCodeMask (ViSession instrumentHandle,

ViInt32 branch, ViString longCodeMask);

Purpose

This function defines the mask of the long code in hexadecimal format

(optionally for the I or Q branch) of the code domain power measurement

mode.

Note:

(1) This function is available only in conjunction with option FS-K83 or

FS-K85.

Remote-control command(s):

[SENSe:]CDPower:LCODe:MASK

[SENSe:]CDPower:LCODe:I

[SENSe:]CDPower:LCODe:Q

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

branch

Variable Type       ViInt32

This control selects branch for which is defined long code mask.

Applied only for option FS-K85 mode.

Valid Range:

0 - I (Default Value)

1 - Q

longCodeMask

Variable Type       ViString

This control defines the mask of the long code in hexadecimal format.

Valid Range: "0" to "4FFFFFFFFFF" (I or Q branch "3FFFFFFFFFF")

Default Value: "0"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Branch) out of range.

BFFC0003  Parameter 3 (Long Code Mask) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPLCodeMode

ViStatus rsfsu_confCDPLCodeMode (ViSession instrumentHandle,

ViInt32 longCodeMode);

Purpose

This function defines the mode of the long code generation. If the

softkey is unselected (default) the cdma2000 standard long

code generator is used. If the softkey is selected (green) the Agilent

ESG option 101 long code is used. Only signals from that generator are

possible to analyse then.

Note:

(1) This function is available only in conjunction with option FS-K83.

Remote-control command(s):

SENSe:CDPower:LCODe:MODE STANdard|ESG101

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

longCodeMode

Variable Type       ViInt32

This control defines the mode of the long code generation.

Valid Range:

0 - Standard (Default Value)

1 - ESG101

Note(s):

(1) Standard - the cdma2000 standard long code generator is used

(2) ESG101 - the Agilent ESG option 101 long code is used.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Long Code Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPLCodeOffset

ViStatus rsfsu_confCDPLCodeOffset (ViSession instrumentHandle,

ViString longCodeOffset);

Purpose

This function defines the offset of the long code including the PN offset

in chips in hexadecimal format with a 52-bit resolution. This value

corresponds to the GPS timing since 6.1.1980 00:00:00 UTC. This long code

offset is used for offsetting at the next trigger pulse (which cannot

occur until a setup time of 300 ms has elapsed). The command is only

available if a long code mask unequal to 0 is set.The chips offset is

calculated as follows: tSinceStartGPS * 1.2288 MChips/s, where

tSinceStartGPS should be in seconds. For example, the hexadecimal offset

of 258000 h chips is set for the first even second clock trigger.

Note:

(1) This function is available only in conjunction with option FS-K83.

Remote-control command(s):

[SENSe:]CDPower:LCODe:OFFSet

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

longCodeOffset

Variable Type       ViString

This control defines the offset of the long code including the PN

offset in chips in hexadecimal format with a 52-bit resolution.

Valid Range: "0" to "1FFFFFFFFFF8000"

Default Value: "0"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Long Code Offset) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPLType

ViStatus rsfsu_confCDPLType (ViSession instrumentHandle, ViInt32 type);

Purpose

This function defines the scrambling type.

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K73.

Remote-control command(s):

[CONFigure:]CDPower:LCODe:TYPE SHORt | LONG

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

type

Variable Type       ViInt32

Sets the scrambling code type.

Valid Range:

0 - Short

1 - Long (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Type) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPMeas

ViStatus rsfsu_confCDPMeas (ViSession instrumentHandle,

ViInt32 CDPMeasurement);

Purpose

This function selects the measurement of the Application Firmware R&S

FS-K76 (FS-K77), TD-SCDMA base (mobile) station test.

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure<1>:CDPower:MEASurement POWer | ACLR | ESPectrum | OBANdwith |

OBWidth | CDPower | PVTime | CCDF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

CDPMeasurement

Variable Type       ViInt32

This control selects CDP Measurement mode.

Valid Range:

0 - POWer (Default Value)

1 - ACLR

2 - ESPectrum

3 - OBWidth

4 - CDPower

5 - CCDF

6 - PVTime

Notes:

(1) POWer:

Power channel measurement with predefined settings

(2) ACLR:

Adjacent channel power measurements with predefined settings.

(3) ESPectrum:

Check on signal power (Spectrum Emission Mask).

(4) OBWidth:

Measurement of occupied bandwidth.

(5) CDPower:

Code domain analyzer measurement.

(6) CCDF:

Measurement of the Complementary Cumulative Distribution

function (signal statistics measurement).

(7) PVTime:

Measurement of power versus time, here trace readout delivers 512

points.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (CDP Measurement) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPMidambleShift

ViStatus rsfsu_confCDPMidambleShift (ViSession instrumentHandle,

ViInt32 midambleShift);

Purpose

This function defines the maximum number of midamble shifts.

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

[SENSe:]CDPower:MSHift 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

midambleShift

Variable Type       ViInt32

This control defines the maximum number of midamble shifts.

Valid Range: 2, 4, 6, 8, 10, 12, 14, 16

Default Value: 16

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Midamble Shift) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPNormalize

ViStatus rsfsu_confCDPNormalize (ViSession instrumentHandle,

ViBoolean IQOffset);

Purpose

This function switches normalization of the unit circle with the IQ

offset on or off in the code domain power measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K76, FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:NORMalize ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

IQOffset

Variable Type       ViBoolean

This control switches normalization of the unit circle with the IQ

offset on or off.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (IQ Offset) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPOperationMode

ViStatus rsfsu_confCDPOperationMode (ViSession instrumentHandle,

ViInt32 operationMode);

Purpose

This function is used to set the operation mode. This information is used

for the channel search. In TRAFFIC mode all channels (PICH/RRI/DATA/ACK

and DRC) can exist. PICH and RRI are always in the signal. In ACCESS mode

only PICH (always available) and DATA channel can exist.

Note:

(1) This function is available only in conjunction with option FS-K85.

Remote-control command(s):

[SENSe:]CDPower:OPERation ACCess | TRAFfic

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

operationMode

Variable Type       ViInt32

This control defines the operating mode.

Valid Range:

0 - Access

1 - Traffic

Default Value: 1 - Traffic

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Operation Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPOrder

ViStatus rsfsu_confCDPOrder (ViSession instrumentHandle, ViInt32 order);

Purpose

This function sets the order of the code domain evaluation. The codes are

sorted either in Hadamard order or in BitReverse order.

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:ORDer HADamard | BITReverse

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

order

Variable Type       ViInt32

This control sets the order of the code domain evaluation.

Valid Values:

0 - Hadamard (Default Value)

1 - Bit Reverse

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Order) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPOverviewDisplay

ViStatus rsfsu_confCDPOverviewDisplay (ViSession instrumentHandle,

ViBoolean overviewDisplay);

Purpose

This function can be enabled by means of ON when either the Code Domain

Power or the Code Domain Error Power analysis is active.

If is enabled (ON) the I-Branch of the code power is displayed in screen

A and the Q-Branch in screen B. Both results can be read via IEC be using

TRACE:DATA? TRACE1 and TRACE:DATA? TRACE2 respectively.

If an analysis other than Code Domain Power or Code Domain Error Power is

selected when Overview mode is active, you exit Overview mode and the

previous analysis is reset on the other screen.

Note:

(1) This function is available only in conjunction with option FS-K73,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

[SENSe<1|2>:]CDPower:OVERview ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

overviewDisplay

Variable Type       ViBoolean

This control switches to an overview display of a code domain

measurement.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Overview Display) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPPControl

ViStatus rsfsu_confCDPPControl (ViSession instrumentHandle,

ViInt32 mode);

Purpose

This function selects the power control mode.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K74.

Remote-control command(s):

[CONFigure:]CDPower:PCONtrol PILot | SLOT

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

mode

Variable Type       ViInt32

This control sets power control mode.

Valid Range:

0 - Pilot (Default Value)

1 - Slot

Notes:

(1) Pilot:

Power control from the beginning of the pilot symbol.

(2) Slot:

Compressed mode.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPPNOffset

ViStatus rsfsu_confCDPPNOffset (ViSession instrumentHandle,

ViInt32 PNOffset);

Purpose

This function sets the PN offset of the base station in multiples of 64

chips.

Note:

(1) This function is available only in conjunction with option FS-K82 or

FS-K84.

Remote-control command(s):

[CONFigure:]CDPower:PNOFfset

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

PNOffset

Variable Type       ViInt32

This control sets the PN offset of the base station in multiples of

64 chips.

Valid Range: 0 to 511

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (PN Offset) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPPref

ViStatus rsfsu_confCDPPref (ViSession instrumentHandle,

ViInt32 preference);

Purpose

This function switches between the use of total power or CPICH (PICH)

power as reference for the relative CDP measurement values in the code

domain power measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K82, FS-K83 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:PREFerence TOTal| CPICh | PICH

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

preference

Variable Type       ViInt32

This control switches between the use of total power or CPICH (PICH)

power as reference for the relative CDP measurement values.

Valid Range:

0 - Total (Default Value)

1 - CPICH

2 - PICH

Notes:

(1) CPICH is available for options FS-K72 or FS-K73 only.

(2) PICH is available for options FS-K82, FS-K83 or FS-K85 only.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Preference) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPQInvert

ViStatus rsfsu_confCDPQInvert (ViSession instrumentHandle,

ViBoolean qInvert);

Purpose

This function inverts the Q component of the signal in the code domain

power measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K76, FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:QINVert ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

qInvert

Variable Type       ViBoolean

This control inverts the Q component of the signal.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Q Invert) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPRFSlot

ViStatus rsfsu_confCDPRFSlot (ViSession instrumentHandle,

ViInt32 RFSlot);

Purpose

This function defines the expected signal: FULL slot or IDLE slot.

Accordingly the limit lines and the borders for calculating the mean

power are set. The lower and upper limit line are called DOPVTFL/DOPVTFU

for FULL and DOPVTIL/DOPVTIU for IDLE mode. It is possible to change

these lines with the standard limit line editor.

Note:

(1) This function is available only in conjunction with option FS-K84.

Remote-control command(s):

CONFigure:CDPower:RFSLot FULL | IDLE

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

RFSlot

Variable Type       ViInt32

This control defines the expected signal: FULL slot or IDLE slot.

Valid Range:

0 - Full

1 - Idle

Default Value: 0 - Full

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (RF Slot) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPRrcFilter

ViStatus rsfsu_confCDPRrcFilter (ViSession instrumentHandle,

ViBoolean state);

Purpose

This function switches the root raised cosine (RRC) receiver filter ON or

OFF

Note:

(1) This function is available only in conjunction with the option

FS-K73.

Remote-control command(s):

[SENSe:]CDPower:FILTer ON|OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

state

Variable Type       ViBoolean

Switches the YIG filter for image frequency suppression into the

signal path (condition ON) and/or removes it from the signal path

(condition OFF).

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSBand

ViStatus rsfsu_confCDPSBand (ViSession instrumentHandle,

ViInt32 sideBand);

Purpose

This function is for interchanging the left and the right sideband in the

code domain power measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K76, FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:SBANd NORMal | INVerse

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

sideBand

Variable Type       ViInt32

This control selects the left or the right sideband.

Valid Range:

0 - Normal (Default Value)

1 - Inverse

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Side Band) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSetStd

ViStatus rsfsu_confCDPSetStd (ViSession instrumentHandle);

Purpose

This function is for CDP presetting.

Inactive Channel Threshold      -23 dB

Limits                          set to standard

Result display                  3GPPR

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K73.

Remote-control command(s):

[CONFigure:]CDPower:PRESet

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSFactor

ViStatus rsfsu_confCDPSFactor (ViSession instrumentHandle,

ViInt32 spreadingFactor);

Purpose

This function defines the spreading factor in code domain power

measurement mode. The spreading factor is only significant for display

mode PEAK CODE DOMAIN ERROR. The base spreading factor should be set to

64 if channels of code class 6 (spreading factor 64) exist, otherwise

alias power values will be measured in the code domain power analysis.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K82 or FS-K83.

Remote-control command(s):

[CONFigure:]CDPower:SFACtor

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

spreadingFactor

Variable Type       ViInt32

This control defines the spreading factor.

Valid Range FS-K72, FS-K73:

4 - 4

8 - 8

16 - 16

32 - 32

64 - 64

128 - 128

256 - 256

512 - 512

Valid Range FS-K82:

64 - 64

128 - 128

Valid Range FS-K83:

16 - 16

32 - 32

64 - 64

Default Value: 64

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Spreading Factor) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSigMap

ViStatus rsfsu_confCDPSigMap (ViSession instrumentHandle,

ViInt32 signalMapping);

Purpose

This function defines the mapping (selects whether the I or Q branch

should be analyzed) of the CDP signal in the code domain power

measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:MAPPing I | Q

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

signalMapping

Variable Type       ViInt32

This control defines the mapping of the CDP signal. Selects whether

the I or Q branch should be analyzed.

Valid Range:

0 - I (Default Value)

1 - Q

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Signal Mapping) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSigMapMode

ViStatus rsfsu_confCDPSigMapMode (ViSession instrumentHandle,

ViInt32 signalMappingMode);

Purpose

This function defines the mapping mode. With AUTO, internal working takes

place in "Complex" mapping mode for the DATA channel type and "I or Q"

for the PILOT, MAC and PREAMBLE channel types. With this function, each

mapping mode can be determined personally for all channel types, or it is

possible to return to AUTO.

Note:

(1) This function is available only in conjunction with option FS-K84.

Remote-control command(s):

[SENSe:]CDPower:MMODe AUTO | IOQ | COMPlex

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

signalMappingMode

Variable Type       ViInt32

This control defines the mapping mode.

Valid Range:

0 - Auto

1 - I or Q

2 - Complex

Default Value: 0 - Auto

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Signal Mapping Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSlot

ViStatus rsfsu_confCDPSlot (ViSession instrumentHandle,

ViInt32 slotNumber);

Purpose

This function sets the slot number in the code domain power measurement

mode.

Note:

(1) For FS-K85 option selects the half slot (and not the whole slot). To

ensure compatibility with other 3G mobile radio applications, no new

command has been introduced for the half slot (the slot command has

simply been reused).

(2) This function is available only in conjunction with option FS-K72,

FS-K73, FS-K76, FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:SLOT

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

slotNumber

Variable Type       ViInt32

This control sets the slot number.

Valid Range:

FS-K72 or FS-K73: 0 to 14

FS-K76 or FS-K77: 0 to IQLength - 1

FS-K82 or FS-K83: 0 to IQLength - 1

FS-K84 or FS-K85: 0 to IQLength - 1

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Slot Number) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSlotDifference

ViStatus rsfsu_confCDPSlotDifference (ViSession instrumentHandle,

ViBoolean differenceState);

Purpose

This function selects the indication of the slot power difference between

the actual slot and the previous slot.

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K74

(2) To get the results it is necessary to activate the power versus slot

measurement in the result display.

Remote-control command(s):

SENSe1:CDPower:PDIFf

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

differenceState

Variable Type       ViBoolean

Selects the indication of the slot power difference between the

actual slot and the previous slot.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Difference State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSlotSetsCount

ViStatus rsfsu_confCDPSlotSetsCount (ViSession instrumentHandle,

ViInt32 count);

Purpose

This function defines how many SETS of the SLOTS shall be captured

consecutively into the IQ RAM of the R&S FSQ.

Notes:

(1) This function is available only in conjunction with option FS-K76,

FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

(2) FS-K82 and FS-K83 - the SLOTS are 64 PCGs.

(3) FS-K84 - the SLOTS are 32 slots.

(4) FS-K85 - the SLOTS are 64 half slots.

(5) FS-K76 and FS-K77 - the SLOTS are 63 slots.

Remote-control command(s):

SENSe:CDPower:SET:COUNt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

count

Variable Type       ViInt32

This control sets the number of slots.

Valid Range:

1 to 57: for options FS-K84, FS-K85

1 to 37: for option FS-K82

1 to 38: for option FS-K83

1 to 190: for option FS-K76, FS-K77

Default Value: 1

Note:

(1) For the option FS-K85: If the count is greater than 1 the CAPTURE

LENGTH will be implicitly set to 64 half slots.

(2) For the option FS-K84: If the count is greater than 1 the CAPTURE

LENGTH will be implicitly set to 32 slots.

(3) For the option FS-K82: If the count is greater than 1 the CAPTURE

LENGTH will be implicitly set to 64 PCGs.

(3) For the option FS-K83: If the count is greater than 1 the CAPTURE

LENGTH will be implicitly set to 64 PCGs.

(4) For the option FS-K76 and FS-K77: If the count is greater than 1

the CAPTURE LENGTH will be implicitly set to 63 slots.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Count) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPStandard

ViStatus rsfsu_confCDPStandard (ViSession instrumentHandle,

ViInt32 standard);

Purpose

This function switches between the standard 3GPP and TSM.

Note:

(1) This function is available only in conjunction with option FS-K76.

Remote-control command(s):

[SENSe:]CDPower:STANdard GPP | TSM

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

standard

Variable Type       ViInt32

This control switches between the standard 3GPP and TSM.

Valid Range:

0 - 3GPP (Default Value)

1 - TSM

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Standard) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSubframes

ViStatus rsfsu_confCDPSubframes (ViSession instrumentHandle,

ViInt32 number_ofSubframes);

Purpose

This function defines the number of subframes for averaging.

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:PVTime:SFRames

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

number_ofSubframes

Variable Type       ViInt32

This control defines the number of subframes for averaging.

Valid Range: > 0

Default Value: 100

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Number of Subframes) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSwitchingPoint

ViStatus rsfsu_confCDPSwitchingPoint (ViSession instrumentHandle,

ViInt32 switchingPoint);

Purpose

This function sets the switching point between uplink and downlink slots.

Note:

(1) This function is available only in conjunction with option FS-K76 or

FS-K77.

Remote-control command(s):

CONFigure:CDPower:PVTime:SPOint

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

switchingPoint

Variable Type       ViInt32

This control sets the switching point between uplink and downlink

slots.

Valid Range: 1 to 7

Default Value: 3

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Switching Point) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPSyncType

ViStatus rsfsu_confCDPSyncType (ViSession instrumentHandle,

ViInt32 syncType);

Purpose

This function selects the type of synchronization in the code domain

power measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K73.

Remote-control command(s):

[CONFigure:]CDPower:STYPe SCHannel | CPICh

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

syncType

Variable Type       ViInt32

This control selects type of synchronization.

Valid Range:

0 - SCHannel (Default Value)

1 - CPICH

Notes:

(1) CPICH:

CPICh synchronizaton to CPICH (default). The CPICH has to be

contained in signal.

(2) SCHannel:

SCHannel synchronization without CHICH. This type of synchronization

is provided for test model 4 which is defined with or without CPICH.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Sync Type) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCDPTPM

ViStatus rsfsu_confCDPTPM (ViSession instrumentHandle,

ViBoolean offsetEvaluation);

Purpose

This function supports the specific enabling and disabling of the timing

and phase offset analysis of the channels for the first active channel in

the channel type.

Note:

(1) This function is available only in conjunction with option FS-K82,

FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

[CONFigure:]CDPower:TPMeas ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

offsetEvaluation

Variable Type       ViBoolean

This control switches on or off the timing and phase offset

evaluation of the channels for the pilot channel.

Valid Values:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Offset Evaluation) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCoup

ViStatus rsfsu_confCoup (ViSession instrumentHandle, ViInt32 window,

ViInt32 bandwidthOperation, ViInt32 bandwidth,

ViInt32 resolutionMode,

ViInt32 bandwidthPLL_not_used, ViReal64 ratio);

Purpose

This function controls the setting of the analyzer coupled functions for

resolution and video filter.

Note:

(1) This function is reserved for compatibility with FSEx driver. Prefer

to use function rsfsu_confCoupExt.

(2) If compatibility checking is enabled, the function will return an

error if values marked with RESERVED are used in any parameter.

Remote-control command(s):

SENSe<1|2>:BANDwidth:RESolution:AUTO ON | OFF

SENSe<1|2>:BANDwidth:RESolution:RATio

SENSe<1|2>:BANDwidth:RESolution:MODE ANALOG | DIGITAL

SENSe<1|2>:BANDwidth:RESolution:FFT ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

bandwidthOperation

Variable Type       ViInt32

This control selects the bandwidth operation to be processed.

Bandwidth operation refers to the control with similar label. See

their helps for more information.

Valid Range:

0 - Resolution Bandwidth (Default Value)

1 - Video Bandwidth

2 - Sweeptime

3 - PLL Bandwidth (RESERVED)

4 - Coupling Set To Default

bandwidth

Variable Type       ViInt32

This control switchs the resolution modes.

Valid Range:

0 - Auto Off (Default Value)

1 - Auto On

2 - Ratio

3 - Digital Filter Mode (see Notes)

Notes:

(1) Auto On/Off:

This command either automatically couples the resolution (video)

bandwidth of the analyzer to the span or cancels the coupling.

Warning for the AUTO functions: If you first select AUTO ON, and

afterwards you define the value, the FS switches back to AUTO OFF.

This control also sets sweep time mode to AUTO ON or OFF (if

selected).

(2) Ratio:

Allows to set up resolution (video) ratio.

(3) Digital Filter Mode:

This function is only valid for Bandwidth Operation == 0 ->

Resolution Bandwidth. Only in this case, command  Resolution Mode has

to be sent to the FSU.

resolutionMode

Variable Type       ViInt32

This control sets the resolution mode.

Valid Range:

0 - Analog (RESERVED) (Default Value)

1 - Digital (RESERVED)

2 - FFT On

3 - FFT Off

Notes:

(1) RESERVED values does not apply to the instrument.

(2) Analog/Digital (RESERVED):

This command toggles between analog and digital resolution filters

for the 1-kHz bandwidth. Depending on the bandwidth, the resolution

filters are automatically toggled between digital filters (<1kHz) and

analog filters (>1kHz). The 1-kHz bandwidth is present in the

instrument as a digital filter and as an analog filter and can be

toggled using this command. If the analog filter is selected for the

bandwidth 1kHz, the FFT-filtering for bandwidths <= 1kHz is switched

off.

(3) FFT On/Off:

This command toggles the digital filters used for bandwidths <= The

filter bandwidth of 1 kHz is switched to digital filtering for both

ON and OFF. This command is only available in conjunction with option

FFT Filter, FSE-B5.

bandwidthPLL_not_used

Variable Type       ViInt32

This control defines the bandwidth of the main PLL of the analyzer

synthesizer.

Valid Range:

0 - Auto (Default Value)

1 - High

2 - Medium

3 - Low

Note:

(1) This control is reserved for compatibility with FSEx driver  and

does not apply to the instrument.

ratio

Variable Type       ViReal64

This control defines the ratio resolution (video) bandwidth (Hz) span

(Hz).

Valid Ranges:

RBW/VBW 0.01 to 1000.0

span/RBW 0.0001 to 1.0

Default Value: 1.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Bandwidth Operation) out of range.

BFFC0004  Parameter 4 (Bandwidth) out of range.

BFFC0005  Parameter 5 (Resolution Mode) out of range.

BFFC0006  Parameter 6 (Bandwidth PLL (not used)) out of range.

BFFC0007  Parameter 7 (Ratio) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confCoupExt

ViStatus rsfsu_confCoupExt (ViSession instrumentHandle, ViInt32 window,

ViInt32 bandwidthOperation,

ViInt32 bandwidth, ViInt32 filterType,

ViReal64 couplingRatio);

Purpose

This function controls the setting of the analyzer coupled functions for

resolution and video filter.

Remote-control command(s):

SENSe<1|2>:FREQuency:SPAN?

INSTrument<1|2>:NSELect?

SENSe<1|2>:BANDwidth:RESolution:AUTO ON | OFF

SENSe<1|2>:BANDwidth:RESolution:RATio

SENSe<1|2>:BANDwidth:RESolution:TYPE NORMal | FFT | CFILter | RRC | NOISe

| PULSe | P5

SENSe<1|2>:BANDwidth:VIDeo:AUTO ON | OFF

SENSe<1|2>:BANDwidth:VIDeo:RATio

SENSe<1|2>:SWEep:TIME:AUTO ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

bandwidthOperation

Variable Type       ViInt32

This control selects the bandwidth operation to be processed.

Bandwidth operation refers to the control with similar label. See

their helps for more information.

Valid Range:

0 - Resolution Bandwidth (Default Value)

1 - Video Bandwidth

2 - Sweeptime

3 - Coupling Set To Default

bandwidth

Variable Type       ViInt32

This control switchs the resolution modes.

Valid Range:

0 - Auto Off (Default Value)

1 - Auto On

2 - Coupling Ratio

3 - Digital Filter Mode (see Notes)

Notes:

(1) Auto On/Off:

This command either automatically couples the resolution (video)

bandwidth of the analyzer to the span or cancels the coupling.

Warning for the AUTO functions: If you first select AUTO ON, and

afterwards you define the value, the FSQ switches back to AUTO OFF.

This control also sets sweep time mode to AUTO ON or OFF (if

selected). This command is not available for GSM/EDGE measurements

modulation accuracy (MAC), phase/frequency error (PFE) and power

versus time (PVT).

(2) Coupling Ratio:

Allows to set up resolution (video) coupling ratio. The selection

follows the bandwidth selected with Bandwidth Operation.

(3) Digital Filter Mode:

This function is only valid for Bandwidth Operation == 0 ->

Resolution Bandwidth. Only in this case, command  Resolution Mode

will be sent to the instrument.

filterType

Variable Type       ViInt32

This control switches the filter type for the resolution bandwidth

between "normal" analog or FIR filters in 1, 3, 10 steps and the FFT

filtering for bandwidths <100 kHz. The advantage of FFT filtering is

the higher measurement speed compared to digital filters with analog

filter characteristic. However, FFT filters are only suitable for

periodic signals, and they are only available for span > 0 Hz.

Since firmware version 1.10 steep-edged channel filters and filters

with RRC (Root Raised Cosine) characteristic are available. The

possible combinations of filter type and filter bandwidth are listed

in the table "List of available channel filters" in chapter 4,

section "Setting Bandwidths and Sweep

Time -Key BW".

The filter type PULSe selects the 6 dB bandwidths for EMI

measurements rather than the 3 dB bandwidths for spectrum analysis,

which is selected with type NOISe.

The resolution bandwidths are implemented by 5 pole filter (Gaussian

filter approximation) with the set 3 dB bandwidth. For bandwidths up

to 100 kHz, digital bandpass filters are used. Due to reduced group

delay time, this filter type is usefully for gated sweep

measurements.

Valid Range:

0 - Normal (Default Value)

1 - FFT

2 - Cosine Filter

3 - Root Raised Cosine

4 - Noise

5 - Pulse

6 - 5-pole

Note:

(1) When changing the filter type, the next larger filter bandwidth

is selected if the same filter bandwidth is not available for the new

filter type.

(2) This command is not available for GSM MS measurements modulation

accuracy (MAC), phase/frequency error (PFE), power versus time (PVT)

and and carrier power (CPW).

(3) This function is only valid for Bandwidth Operation == 0 ->

Resolution Bandwidth and Bandwidth == 3 -> Digital Filter Mode. Only

in this case, command  Resolution Mode will be sent to the

instrument.

couplingRatio

Variable Type       ViReal64

This control defines the ratio resolution (video) bandwidth (Hz) span

(Hz).

Valid Ranges:

RBW/VBW 0.01 to 1000.0

span/RBW 0.0001 to 1.0

Default Value: 1.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Bandwidth Operation) out of range.

BFFC0004  Parameter 4 (Bandwidth) out of range.

BFFC0005  Parameter 5 (Filter Type) out of range.

BFFC0006  Parameter 6 (Coupling Ratio) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confDefConvLoss

ViStatus rsfsu_confDefConvLoss (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function set default conversion loss value of external mixer.

Notes:

(1) The command is available only if the external mixer (option FSU-B21)

is switched on.

Remote-control command(s):

[SENSe:]MIXer:LOSS DEFault

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

(2) This parameter is used for compatibility purposes only.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confDeltaMarkOpt

ViStatus rsfsu_confDeltaMarkOpt (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 trace, ViBoolean deleteAllMkrs,

ViBoolean deltaMarker);

Purpose

This function selects the trace for a delta marker, deletes all delta

marker and controls the relative/absolute delta marker position.

Note:

Markers 2 to 4 are either Normal or Delta Markers. Marker 1 always serves

as the reference marker for all deltamarkers.

Remote-control command(s):

CALCulate<1|2>:DELTamarker:AOFF

CALCulate<1|2>:DELTamarker<1..4>:TRACe

CALCulate<1|2>:DELTamarker<1..4>:MODE RELative | ABSolute

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

trace

Variable Type       ViInt32

Assigns the selected marker to the indicated test curve.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

deleteAllMkrs

Variable Type       ViBoolean

Deletes all delta markers, if set to VI_TRUE.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On (Delete all delta markers)

deltaMarker

Variable Type       ViBoolean

Toggles between absolute and relative measurement.

Valid Range:

VI_FALSE (0) - Abs

VI_TRUE  (1) - Rel (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Trace) out of range.

BFFC0005  Parameter 5 (Delete All Mkrs) out of range.

BFFC0006  Parameter 6 (Delta Marker) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confDeltaMarkPosition

ViStatus rsfsu_confDeltaMarkPosition (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViReal64 markerPosition);

Purpose

This function sets a delta marker to the specified frequency.

Remote-control command(s):

CALCulate<1|2>:DELTamarker<1..4>:X

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

markerPosition

Variable Type       ViReal64

Sets the selected marker to the specified position. This position may

be absolute or relative to the reference marker (see

rsfsu_confSAMDeltaMarkOpt).

Valid Range: not checked

Default Value: 3.0E9

Notes:

(1) Valid Range depends on selected horizontal display unit (analyzer

mode: freq. domain (Hz), time domain (s)).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Marker Position) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confDeltaMarkState

ViStatus rsfsu_confDeltaMarkState (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViBoolean markerState);

Purpose

This function enables or disables a delta marker.

Remote-control command(s):

CALCulate<1|2>:DELTamarker<1..4>:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

markerState

Variable Type       ViBoolean

Enables or disables the selected marker in the specified window.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Marker State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confDisplayComment

ViStatus rsfsu_confDisplayComment (ViSession instrumentHandle,

ViString commentString);

Purpose

This function displays the comment string on the instruments screen.

Remote-control command(s):

DISPlay:TEXT:DATA

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

commentString

Variable Type       ViString

This control defines a comment which can be displayed on the screen.

Valid Range: not checked

Default Value: "This is the comment string ..."

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Comment String) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confDisplayControl

ViStatus rsfsu_confDisplayControl (ViSession instrumentHandle,

ViBoolean frequency, ViBoolean logo,

ViBoolean time,

ViBoolean displayComment);

Purpose

This function controls the apperance of display elements.

Remote-control command(s):

DISPlay:LOGO ON | OFF

DISPlay:TIME ON | OFF

DISPlay:ANNotation:FREQuency ON | OFF

DISPlay:TEXT:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

frequency

Variable Type       ViBoolean

This control switches the x-axis annotation on or off.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

logo

Variable Type       ViBoolean

This control switches the factory logo on the screen on or off.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

time

Variable Type       ViBoolean

This control switches on or off the screen display of date and time.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

displayComment

Variable Type       ViBoolean

This control switches on or off the screen display of the comment.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Frequency) out of range.

BFFC0003  Parameter 3 (Logo) out of range.

BFFC0004  Parameter 4 (Time) out of range.

BFFC0005  Parameter 5 (Display Comment) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confDisplayMode

ViStatus rsfsu_confDisplayMode (ViSession instrumentHandle,

ViBoolean displayFormat);

Purpose

This function switches the test result display between FULL SCREEN and

SPLIT SCREEN.

Remote-control command(s):

DISPlay:FORMat SINGLE | SPLIT

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

displayFormat

Variable Type       ViBoolean

This control switches the display format.

Valid Range:

VI_FALSE (0) - Single Screen (Default Value)

VI_TRUE  (1) - Split Screen

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Display Format) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confDisplayParams

ViStatus rsfsu_confDisplayParams (ViSession instrumentHandle,

ViInt32 displayParameter,

ViInt32 colorMap,

ViInt32 predefinedColors,

ViReal64 TINT, ViReal64 SATURATION,

ViReal64 BRIGHTNESS);

Purpose

This function configures the apperance of display.

Remote-control command(s):

DISPlay:CMAP<1..26>:HSL

DISPlay:CMAP:DEFault<1|2>

DISPlay:CMAP<1..26>:PDEFined BLACk | BLUE | BROWn | GREen | CYAN | RED |

MAGenta | YELLow | WHITe | DGRAy | LGRAy | LBLUe | LGREen | LCYan | LRED

| LMAGenta

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

displayParameter

Variable Type       ViInt32

This control selects desired display parameter to be configured.

Valid Range:

0 - Brightness, Saturation, Tint (Default Value)

1 - Predefined Colors

2 - Default Colors (settings 1)

3 - Default Colors (settings 2)

Notes:

(1) Brightness, Saturation, Tint:

Sets Brightness, Saturation, Tint for selected color table.

(2) Predefined Colors:

Sets predefined colors for selected color table.

(3) 2 - Default Colors (settings 1 or 2):

Selects default setting 1 or 2 for setting the colours.

colorMap

Variable Type       ViInt32

This control defines the color table of the instrument.

Valid Range:

1 - Background

2 - Grid

3 - Function field + status field + data entry text

4 - Function field LED on

5 - Function field LED warn

6 - Enhancement label text

7 - Status field background

8 - Trace 1 (Default Value)

9 - Trace 2

10 - Trace 3

11 - Marker

12 - Lines

13 - Measurement status + limit check pass

14 - Limit check fail

15 - Table + softkey background

16 - Table + softkey text

17 - Table selected field text

18 - Table selected field background

19 - Table + data entry field opaq titlebar

20 - Data entry field opaq text

21 - Data entry field opaq background

22 - 3D shade bright part

23 - 3D shade dark part

24 - Softkey state on

25 - Softkey state data entry

26 - Logo

Note:

(1) FSU handles colour map different way than FSEx driver. Be careful

when using this control. Below is the table with FSEx old values:

1 - Trace 1

2 - Trace 2

3 - Trace 3

4 - Trace 4

5 - Marker

6 - Grid

7 - Softkey State On

8 - Softkey State Data Entry

9 - Softkey State OFF

10 - Softkey Shade

11 - Text

12 - Title

13 - Background

predefinedColors

Variable Type       ViInt32

This control defines the color table of the instrument using

predefined color values. To each numeric suffix of CMAP is assigned

one or several picture elements which can be modified by varying the

corresponding color setting.

Valid Range:

0 - BLACk (Default Value)

1 - BLUE

2 - BROWn

3 - GREen

4 - CYAN

5 - RED

6 - MAGenta

7 - YELLow

8 - WHITe

9 - DGRAy

10 - LGRAy

11 - LBLUe

12 - LGREen

13 - LCYan

14 - LRED

15 - LMAGenta

TINT

Variable Type       ViReal64

This control sets the hue = TINT.

Valid Range: 0.0 to 100.0 %

Default Value: 30.0 %

Note:

If (TINT > 1) then driver divides this value by 100.0 due to

compatibility with FSEx.

SATURATION

Variable Type       ViReal64

This control sets the sat = SATURATION.

Valid Range: 0.0 to 100.0 %

Default Value: 80.0 %

If (SATURATION > 1) then driver divides this value by 100.0 due to

compatibility with FSEx.

BRIGHTNESS

Variable Type       ViReal64

This control sets the lum = BRIGHTNESS.

Valid Range: 0.0 to 100.0 %

Default Value: 100.0 %

If (BRIGHTNESS > 1) then driver divides this value by 100.0 due to

compatibility with FSEx.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Display Parameter) out of range.

BFFC0003  Parameter 3 (Color Map) out of range.

BFFC0004  Parameter 4 (Predefined Colors) out of range.

BFFC0005  Parameter 5 (TINT) out of range.

BFFC0006  Parameter 6 (SATURATION) out of range.

BFFC0007  Parameter 7 (BRIGHTNESS) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confDisplaySize

ViStatus rsfsu_confDisplaySize (ViSession instrumentHandle,

ViInt32 size);

Purpose

This function selects the size of the displayed graph in channel power

and adjacent channel power measurement.

Remote-control command(s):

DISPlay:SIZE:LARGe | SMALl

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

size

Variable Type       ViInt32

This control selects the size of the graph of CP or ACP measurement.

Valid Values:

0 - Full size (default value)

1 - Half size

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Size) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confDLines

ViStatus rsfsu_confDLines (ViSession instrumentHandle, ViInt32 window,

ViInt32 lineType, ViInt32 lineNumber,

ViInt32 lineOperationSwitch,

ViReal64 linePositionValue);

Purpose

This function sets the display lines in the instrument, i.e.,  frequency

and time lines (depending on the X-axis).

Notes:

(1) The lines for frequency and time are displayed as vertical, solid

lines over the total height of the

diagram. They can be moved in x-direction.

Remote-control command(s):

CALCulate<1..4>:FLINE<1|2>

CALCulate<1..4>:FLINE<1|2>:STATe OFF | ON

CALCulate<1..4>:TLINE<1|2>

CALCulate<1..4>:TLINE<1|2>:STATe OFF | ON

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 2

Default Value: 1

lineType

Variable Type       ViInt32

Selects desired line type.

Valid Range:

0 - Reserved

1 - Frequency Line (Default Value)

2 - Reserved

3 - Reserved

4 - Time Line

Notes:

(1) Frequency Line:

Two vertical frequency lines for indicating frequencies or for

determining frequency search ranges (Frequency Line 1/2). The

frequency lines are only valid for SPAN>0.

(2) Time Line:

Two vertical time lines for indicating times or for determining time

search ranges (Time Line 1/2). The time lines are only valid for

SPAN=0.

lineNumber

Variable Type       ViInt32

This control defines the line number.

Valid Range:

1, 2

Default Value:

1

lineOperationSwitch

Variable Type       ViInt32

Switchs on or off display lines.

Valid Range:

0 - Line Off (Default Value)

1 - Line On

linePositionValue

Variable Type       ViReal64

This control sets the line position value. This value depends on line

type. Parameter unit depends on type of display line and selected

horizontal or vertical scaling.

Valid Range:

For Frequency Line:

0.0 to f(max) Hz

For Time Line:

0.0 to 1000.0 s

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Line Type) out of range.

BFFC0004  Parameter 4 (Line Number) out of range.

BFFC0005  Parameter 5 (Line Operation Switch) out of range.

BFFC0006  Parameter 6 (Line Position Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confEmulation

ViStatus rsfsu_confEmulation (ViSession instrumentHandle,

ViInt32 emulation);

Purpose

This function activates the emulation of various spectrum analysers.

Notes:

Selecting a language different from "SCPI" will set the GPIB address to

18 if it was 20 before. Start / stop frequency, reference level and # of

sweep points will be adapted to the selected instrument model.

Remote-control command(s):

SYSTem:LANGuage "SCPI" | "8560E" | "8561E" | "8562E" | "8563E" | "8564E"

| "8565E" | "8566A" | "8566B" | "8568A" | "8568B" | "8594E"

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

emulation

Variable Type       ViInt32

This control selects the emulated analyzer.

Valid Values:

0 - SCPI (Default Value)

1 - 8566A

2 - 8566B

3 - 8568A

4 - 8568B

5 - 8594E

6 - 8560E

7 - 8561E

8 - 8562E

9 - 8563E

10 - 8564E

11 - 8565E

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Emulation) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confExternGainCorrect

ViStatus rsfsu_confExternGainCorrect (ViSession instrumentHandle,

ViInt32 window, ViReal64 gain);

Purpose

This function sets the extern preamplifier gain value.

Note:

(1) This function works only in conjunction with the option B10

Remote-control command(s):

SENSe<1|2>:CORRection:EGAin:INPut:MAGNitude

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

gain

Variable Type       ViReal64

Sets the gain of an external preamplifier.

Valid Range:

-200.0 to 200.0 dB

Default Value:

0.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Gain) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confExtMix

ViStatus rsfsu_confExtMix (ViSession instrumentHandle, ViInt32 window,

ViBoolean externalMixerState);

Purpose

This function activates or shuts off the external mixer.

Notes:

(1) This function is available only with option FSU-B21.

Remote-control command(s):

[SENSe:]MIXer[:STATe] ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

(2) This parameter is used for compatibility purposes only.

externalMixerState

Variable Type       ViBoolean

Turns on or off the external mixer.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Note:

Turning external mixer off automatically enables internal mixer.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (External Mixer State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confExtMixerLOLevel

ViStatus rsfsu_confExtMixerLOLevel (ViSession instrumentHandle,

ViReal64 mixerLOLevel);

Purpose

This function changes the LO level of the external mixer.

Notes:

(1) This function is available only with option FSU-B21.

Remote-control command(s):

[SENSe1:]MIXer:LOPower

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

mixerLOLevel

Variable Type       ViReal64

Sets the LO level of the external mixer LO port

Valid Range: 13.0 to 17.0 dBm

Default Value: 15.5 dBm

Note:

(1) Values can be changed in 0.1 dB steps.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Mixer LO Level) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confExtMixLossTab

ViStatus rsfsu_confExtMixLossTab (ViSession instrumentHandle,

ViInt32 window, ViString tableName,

ViString mixerType,

ViString serialNumber, ViInt32 band,

ViInt32 harmonic, ViReal64 bias,

ViString comment, ViBoolean ports,

ViInt32 numberOfValues,

ViReal64 _VI_FAR frequencyValues[],

ViReal64 _VI_FAR conversionLossValues[]);

Purpose

This function creates or overwrites conversion loss table of external

mixer.

Notes:

(1) This function is available only with option FSU-B21.

(2) If the external mixer is not turned on, this function will return an

error.

Remote-control command(s):

[SENSe:]CORRection:CVL:SELect

[SENSe:]CORRection:CVL:MIXer

[SENSe:]CORRection:CVL:SNUMber

[SENSe:]CORRection:CVL:BAND A|Q|U|V|E|W|F|D|G|Y|J|USER

[SENSe:]CORRection:CVL:HARMonic

[SENSe:]CORRection:CVL:BIAS

[SENSe:]CORRection:CVL:COMMent

[SENSe:]CORRection:CVL:PORTs 2 | 3

[SENSe:]CORRection:CVL:DATA

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

(2) This parameter is used for compatibility purposes only.

tableName

Variable Type       ViString

Enter name of table (without extention) to edit or create. A new

table will be created if the specified name does not exist.

Valid Range: max. 8 characters

Default Value: ""

mixerType

Variable Type       ViString

Defines the type designation of the mixer in the Conversion Loss

Table.

Valid Range: max. 16 characters

Default Value: ""

serialNumber

Variable Type       ViString

Enter the serial number of the mixer.

Valid Range: max. 16 characters

Default Value: ""

band

Variable Type       ViInt32

Defines waveguide band in the Conversion Loss table.

Valid Range:

0 - A

1 - Q

2 - U (Default Value)

3 - V

4 - E

5 - W

6 - F

7 - D

8 - G

9 - Y

10 - J

11 - User

harmonic

Variable Type       ViInt32

Defines the harmonic of the mixer in the Conversion Loss Table.

Valid Range: >= 2

Default Value: 2

bias

Variable Type       ViReal64

Enters the bias current.

Valid Range: -10.0 mA to 10.0 mA

Default Value: 0.0 mA

comment

Variable Type       ViString

Enter the conevrsion loss table comment.

Valid Range: max. 60 characters

Default Value: ""

ports

Variable Type       ViBoolean

Specifies the mixer type (two-port mixer or three-port mixer).

Valid Range:

VI_FALSE (0) - 2 (Default Value)

VI_TRUE  (1) - 3

numberOfValues

Variable Type       ViInt32

Inputs number of reference value pairs to be send to the  instrument.

Valid Range: 1 to 50

Default Value: 50

frequencyValues

Variable Type       ViReal64[]

Specifies array of frequency values in Hz.

Valid Range: none

Default Value: none

Notes:

(1) The reference values must be entered in order of increasing

frequencies.

(2) A maximum of 50 reference values can be entered.

conversionLossValues

Variable Type       ViReal64[]

Specifies array of level values in dB

Valid Range: none

Default Value: none

Notes:

(1) The reference values must be entered in order of increasing

frequencies.

(2) A maximum of 50 reference values can be entered.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Table Name) out of range.

BFFC0004  Parameter 4 (Mixer Type) out of range.

BFFC0005  Parameter 5 (Serial Number) out of range.

BFFC0006  Parameter 6 (Band) out of range.

BFFC0007  Parameter 7 (Harmonic) out of range.

BFFC0008  Parameter 8 (Bias) out of range.

BFFC0009  Parameter 9 (Comment) out of range.

BFFC000A  Parameter 10 (Ports) out of range.

BFFC000B  Parameter 11 (Number Of Values) out of range.

BFFC000C  Parameter 12 (Frequency Values) out of range.

BFFC000D  Parameter 13 (Conversion Loss Values) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confExtMixLossTabDelete

ViStatus rsfsu_confExtMixLossTabDelete (ViSession instrumentHandle,

ViInt32 window,

ViString tableName);

Purpose

This function deletes a conversion loss table of external mixer.

Notes:

(1) This function is available only with option FSU-B21.

(2) If the external mixer is not turned on, this function will return an

error.

Remote-control command(s):

[SENSe:]CORRection:CVL:SELect

[SENSe:]CORRection:CVL:CLEar

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

(2) This parameter is used for compatibility purposes only.

tableName

Variable Type       ViString

Enter name of table (without extention) to delete.

Valid Range: max. 8 characters

Default Value: ""

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Table Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confExtMixParameters

ViStatus rsfsu_confExtMixParameters (ViSession instrumentHandle,

ViInt32 band,

ViInt32 harmonicOrder,

ViInt32 harmonicType,

ViBoolean ports, ViReal64 biasLow,

ViReal64 biasHigh,

ViReal64 conversionLossLow,

ViReal64 conversionLossHigh,

ViString conversionLossTableLow,

ViString conversionLossTableHigh);

Purpose

This function selects the waveguide band, the harmonic setting and all

the other necessary parameters of external mixer.

Notes:

(1) This function is available only with option FSU-B21.

(2) Prior to use this function switch External Mixer to On.

Remote-control command(s):

[SENSe:]MIXer:HARMonic:BAND A|Q|U|V|E|W|F|D|G|Y|J|USER

[SENSe:]MIXer:HARMonic

[SENSe:]MIXer:HARMonic:TYPE ODD | EVEN | EODD

[SENSe:]MIXer:PORTs 2 | 3

[SENSe:]MIXer:BIAS[:LOW]

[SENSe:]MIXer:BIAS:HIGH

[SENSe:]MIXer:LOSS[:LOW]

[SENSe:]MIXer:LOSS:HIGH

[SENSe:]MIXer:LOSS:TABLe[:LOW]

[SENSe:]MIXer:LOSS:TABLe:HIGH

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

band

Variable Type       ViInt32

Defines waveguide band in the Conversion Loss table.

Valid Range:

0 - A

1 - Q

2 - U (Default Value)

3 - V

4 - E

5 - W

6 - F

7 - D

8 - G

9 - Y

10 - J

11 - User

harmonicOrder

Variable Type       ViInt32

Sets order of the harmonic used for conversion.

Valid Range: 2 to 64

Default Value: 2

harmonicType

Variable Type       ViInt32

Sets the harmonic type for the active band.

Valid Range:

0 - Even (Default Value)

1 - Odd

2 - Even And Odd

ports

Variable Type       ViBoolean

Specifies the mixer type (two-port mixer or three-port mixer).

Valid Range:

VI_FALSE (0) - 2 (Default Value)

VI_TRUE  (1) - 3

biasLow

Variable Type       ViReal64

Sets the bias current for the active band. If two harmonics are used

for the band, it is set for the lower harmonic.

Valid Range: -10.0 mA to 10.0 mA

Default Value: 0.0 mA

biasHigh

Variable Type       ViReal64

Sets the bias current of the mixer for higher harmonics in bands with

two harmonics.

Valid Range: -10.0 mA to 10.0 mA

Default Value: 0.0 mA

conversionLossLow

Variable Type       ViReal64

Sets the conversion loss of the mixer. If two harmonics are used for

the band, it is set for the lower harmonic.

Valid Range: 0.0 dB to 100.0 dB

Default Value: 0.0 dB

conversionLossHigh

Variable Type       ViReal64

Sets the conversion loss of the mixer for higher harmonics in bands

with two harmonics.

Valid Range: 0.0 dB to 100.0 dB

Default Value: 0.0 dB

conversionLossTableLow

Variable Type       ViString

Selects a conversion loss table and sets it for the active band. If

two harmonics are used for the band, it is set for the lower

harmonic.

Valid Range: up to 8 characters

Default Value: VI_NULL

Note:

Pass NULL pointer to skip this parameter.

conversionLossTableHigh

Variable Type       ViString

Selects the conversion loss table for higher harmonics in bands with

two harmonics.

Valid Range: up to 8 characters

Default Value: VI_NULL

Note:

Pass NULL pointer to skip this parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Band) out of range.

BFFC0003  Parameter 3 (Harmonic Order) out of range.

BFFC0004  Parameter 4 (Harmonic Type) out of range.

BFFC0005  Parameter 5 (Ports) out of range.

BFFC0006  Parameter 6 (Bias Low) out of range.

BFFC0007  Parameter 7 (Bias High) out of range.

BFFC0008  Parameter 8 (Conversion Loss Low) out of range.

BFFC0009  Parameter 9 (Conversion Loss High) out of range.

BFFC000A  Parameter 10 (Conversion Loss Table Low) out of range.

BFFC000B  Parameter 11 (Conversion Loss Table Hig) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confExtMixSignal

ViStatus rsfsu_confExtMixSignal (ViSession instrumentHandle,

ViInt32 window, ViInt32 signalID,

ViReal64 autoIDThresholddB);

Purpose

This function controls signal identification mode of external mixer.

Note:

This function is available only if the external mixer (option FSU-B21) is

switched on.

Remote-control command(s):

[SENSe:]MIXer:SIGNal ON | OFF | AUTO

[SENSe:]MIXer:THReshold

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

(2) This parameter is used for compatibility purposes only.

signalID

Variable Type       ViInt32

Selects signal identification mode.

Valid Range:

0 - Off (Default Value)

1 - Signal

2 - Auto

Notes:

(1) Signal:

Activates visual signal identification.

(2) Auto:

Activates automatic signal identification.

autoIDThresholddB

Variable Type       ViReal64

Determines the maximum permissible level difference between

measurement and reference sweep in AUTO ID mode.

Valid Range: 0.1 to 100.0 dB

Default Value: 10.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Signal ID) out of range.

BFFC0004  Parameter 4 (Auto ID Threshold [dB]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confFileDecSep

ViStatus rsfsu_confFileDecSep (ViSession instrumentHandle,

ViInt32 decimalSeparator);

Purpose

This function configures the decimal separator (decimal point or comma)

for floating-point numerals contained in the file.

Remote-control command(s):

FORMat:DEXPort:DSEParator POINt | COMMa

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

decimalSeparator

Variable Type       ViInt32

This control sets the decimal separator for the ASCII data stored to

the file.

Valid Values:

0 - Point (.) (Default Value)

1 - Comma (,)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Decimal Separator) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confFMDemod

ViStatus rsfsu_confFMDemod (ViSession instrumentHandle, ViInt32 window,

ViBoolean state);

Purpose

This function controls FM demodulation output state.

Note:

(1) This function is available only in conjunction with option FS-B27.

Remote-control command(s):

[SENSe<1|2>:]FM ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

This control determines the display screen that will be affected.

Valid Range:

1 - Window A (Default Value)

2 - Window B

state

Variable Type       ViBoolean

Activates the FM demodulator output of the instrument.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confFMDemodFilt

ViStatus rsfsu_confFMDemodFilt (ViSession instrumentHandle,

ViInt32 window, ViInt32 filterType,

ViReal64 filterFrequency);

Purpose

This function selects desired AF bandwidth for FM demodulation output.

Note:

(1) This function is available only in conjunction with option FS-B27.

Remote-control command(s):

[SENSe<1|2>:]FM:FILTer[:LPASs]:AUTO OFF

[SENSe<1|2>:]FM:FILTer[:LPASs]:FREQuency <numeric_value>

[SENSe<1|2>:]FM:FILTer:HPASs:FREQuency 0 Hz | 10 Hz | 100 Hz | 1 kHz

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

This control determines the display screen that will be affected.

Valid Range:

1 - Window A (Default Value)

2 - Window B

filterType

Variable Type       ViInt32

This control selects the filter type.

Valid Range:

0 - AF High Pass Filter (Default Value)

1 - AF Low Pass Filter

Notes:

(1) High Pass AF Filter:

Activates the high-pass filter for FM demodulation output with the

given limit to separate the DC component. 1kHz has to be selected if

a short transient time is required.

(2) Low Pass AF Filter:

Activates the low-pass filter for FM demodulation output.

filterFrequency

Variable Type       ViReal64

This control sets the filter frequency.

Valid Values:

- High Pass filter: 0.0, 10.0, 100.0, 1.0e3 Hz

- Low Pass filter:  30.0e3, 100.0e3, 300.0e3, 1.0e6, 5.0e6 Hz

Default Value: 0.0 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Filter Type) out of range.

BFFC0004  Parameter 4 (Filter Frequency) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confFMDemodLPFiltAuto

ViStatus rsfsu_confFMDemodLPFiltAuto (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function couples the AF bandwidth to the selected IF bandwidth

according to the following table.

IF bandwidth         low pass filter

-                    30 kHz (can only be set manually)

<= 500 kHz           100 kHz

<= 1 MHz             300 kHz

<= 5 MHz             1 MHz

> 5 MHz              5 MHz

Note:

(1) This function is available only in conjunction with option FS-B27.

Remote-control command(s):

[SENSe<1|2>:]FM:FILTer[:LPASs]:AUTO ON

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

This control determines the display screen that will be affected.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confFMDemodRange

ViStatus rsfsu_confFMDemodRange (ViSession instrumentHandle,

ViInt32 window, ViReal64 range);

Purpose

This function selects desired AF bandwidth for FM demodulation output.

Note:

(1) This function is available only in conjunction with option FS-B27.

Remote-control command(s):

[SENSe<1|2>:]FM[:DEViation]:RANGe:AUTO OFF

[SENSe<1|2>:]FM[:DEViation]:RANGe[:UPPer]

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

This control determines the display screen that will be affected.

Valid Range:

1 - Window A (Default Value)

2 - Window B

range

Variable Type       ViReal64

This control sets the desired AF bandwidth.

Valid Values: 50.0e3, 150.0e3, 500.0e3, 15.0e6 and 50.0e6 Hz

Default Value: 50.0e3 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Range) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confFMDemodRangeAuto

ViStatus rsfsu_confFMDemodRangeAuto (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function couples the FM range to the selected IF bandwidth. The FM

range is set to the highest value which is smaller or

equal to the IF bandwidth.

Note:

(1) This function is available only in conjunction with option FS-B27.

Remote-control command(s):

[SENSe<1|2>:]FM[:DEViation]:RANGe:AUTO ON

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

This control determines the display screen that will be affected.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confFreqAxisMode

ViStatus rsfsu_confFreqAxisMode (ViSession instrumentHandle,

ViInt32 window, ViBoolean axisMode);

Purpose

This function switches between default linear frequency axis and

logarithmic frequency axis. The logarithmic frequency axis is only

available in spectrum mode and it is not available in zero span mode, in

external mixer mode, with frequency offset or if the ratio stop frequency

/ start frequency is below 1.4.

Remote-control command(s):

DISPlay:WINDow<1|2>:TRACe:X:SPACing LINear | LOGarithmic

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

axisMode

Variable Type       ViBoolean

Switches between default linear frequency axis and logarithmic

frequency axis.

Valid Range:

VI_FALSE (0) - Lin (Default Value)

VI_TRUE  (1) - Log

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Axis Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confFreqOffset

ViStatus rsfsu_confFreqOffset (ViSession instrumentHandle,

ViInt32 window,

ViReal64 frequencyOffsetHz);

Purpose

This function configures the frequency offset of the spectrum analyzer.

Remote-control command(s):

SENSe<1|2>:FREQuency:OFFSet

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

frequencyOffsetHz

Variable Type       ViReal64

Sets the frequency offset which is added to the frequency axis of the

analyzer.

Valid Range: -100.0 GHz to 100.0 GHz

Default Value: 0.0 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Frequency Offset [Hz]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confGetPrinterList

ViStatus rsfsu_confGetPrinterList (ViSession instrumentHandle,

ViChar _VI_FAR printerList[]);

Purpose

This function queries the list of printers under Windows NT.

Remote-control command(s):

SYSTem:COMMunicate:PRINter:ENUMerate:FIRSt?

SYSTem:COMMunicate:PRINter:ENUMerate:NEXT?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

printerList

Variable Type       ViChar[]

This control returns list of printers available under  Windows NT

system. Names are separated using commas (,).

Note:

If no printer is configured an empty string is output.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Printer List) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confGridMode

ViStatus rsfsu_confGridMode (ViSession instrumentHandle, ViInt32 window,

ViBoolean gridMode);

Purpose

This function defines the scale of the Y-axis (absolute or relative) in

the selected measurement window.

Remote-control command(s):

DISPlay:WINDow<1|2>:TRACe:Y:MODE ABSolute | RELative

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

gridMode

Variable Type       ViBoolean

Toggles between absolute and relative modes of grid.

Valid Range:

VI_FALSE (0) - Abs (Default Value)

VI_TRUE  (1) - Rel

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Grid Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confGSMBurstMeasFilter

ViStatus rsfsu_confGSMBurstMeasFilter (ViSession instrumentHandle,

ViInt32 measFilter);

Purpose

This function sets changes between the default low pass filter with 600k

bandwidth and the gaussian 500 kHz bandwidth filter. The later is useful

for steep edges.

Remote-control command(s):

CONFigure:BURSt:PTEMplate:FILTer G500 | B600

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measFilter

Variable Type       ViInt32

Sets changes between the default low pass filter with 600k bandwidth

and the gaussian 500 kHz bandwidth filter.

Valid Range:

0 - 500 kHz

1 - 600 kHz

Default Value: 0 - 500 kHz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Meas Filter) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confGSMBurstSearch

ViStatus rsfsu_confGSMBurstSearch (ViSession instrumentHandle,

ViBoolean burstSearch);

Purpose

This function sets the IQ burst search to ON or OFF.

Remote-control command(s):

CONFigure:BSEarch ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

burstSearch

Variable Type       ViBoolean

Sets the IQ burst search to ON or OFF.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Burst Search) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confGSMBurstSearchThreshold

ViStatus rsfsu_confGSMBurstSearchThreshold (ViSession instrumentHandle,

ViReal64 burstSearchThreshold);

Purpose

This function sets the burst search threshold in dB.

Remote-control command(s):

CONFigure:BSTHreshold

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

burstSearchThreshold

Variable Type       ViReal64

Sets the burst search threshold in dB.

Valid Range: -100.0 dB to 0.0 dB

Default Value: -35.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Burst Search Threshold) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confGSMBurstTimeMeasHighResolution

ViStatus rsfsu_confGSMBurstTimeMeasHighResolution

(ViSession instrumentHandle, ViBoolean highResolutionMode);

Purpose

This function sets the Time Meas High Resolution mode to ON or OFF.

Remote-control command(s):

CONFigure:BURSt:PTEMplate:TMHRes ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

highResolutionMode

Variable Type       ViBoolean

Sets the Time Meas High Resolution mode to ON or OFF.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (High Resolution Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confGSMBurstZoomTransitionNumber

ViStatus rsfsu_confGSMBurstZoomTransitionNumber

(ViSession instrumentHandle, ViInt32 transitionNumber);

Purpose

This function sets transition number to zoom in.

Note:

(1) This function is available only in conjunction with option FS-K5.

Remote-control command(s):

CONFigure:BURSt:PTEMplate:FRZoom <numeric_value>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

transitionNumber

Variable Type       ViInt32

This control sets transition number to zoom in.

Valid Range:

1 to (active slots-1)

Default Value:

1

Note:

(1) No range is checked.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Transition Number) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confGSMMultiCarrierModeState

ViStatus rsfsu_confGSMMultiCarrierModeState (ViSession instrumentHandle,

ViBoolean multiCarrierMode);

Purpose

This function sets the multi carrier mode to ON or OFF.

Note:

(1) This function is available only in conjunction with option FS-K5.

Remote-control command(s):

CONFigure:MCARrier ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

multiCarrierMode

Variable Type       ViBoolean

Sets the multi carrier mode to ON or OFF.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Multi Carrier Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confGSMSamplesPerSymbol

ViStatus rsfsu_confGSMSamplesPerSymbol (ViSession instrumentHandle,

ViInt32 samplesPerSymbol);

Purpose

This function determines the number of samples per symbol (points per

symbol).

Note:

(1) This function is available only in GSM/EDGE mode (option FS-K5).

Remote-control command(s):

CONFigure:PRATe 4 | 8

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

samplesPerSymbol

Variable Type       ViInt32

This control selects the number of samples per symbol.

Valid Range:

4, 8

Default Value:

4

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Samples Per Symbol) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confGSMSyncSearch

ViStatus rsfsu_confGSMSyncSearch (ViSession instrumentHandle,

ViBoolean syncSearch);

Purpose

This function sets the IQ sync search to ON or OFF.

Remote-control command(s):

CONFigure:SSEarch ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

syncSearch

Variable Type       ViBoolean

Sets the IQ sync search to ON or OFF.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Sync Search) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confHCComment

ViStatus rsfsu_confHCComment (ViSession instrumentHandle,

ViString screenTitle,

ViString _1stScreenComment,

ViString _2ndScreenComment);

Purpose

Sets the comment line for screen and comment text for printout to trace 1

or 2.

Remote-control command(s):

DISPlay:TEXT

HCOPy:ITEM:WINDow<1|2>:TEXT

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

screenTitle

Variable Type       ViString

This control defines the title of the screen with a maximum of 60

characters.

Valid Range: 60 characters

Default Value: "abcdef"

_1stScreenComment

Variable Type       ViString

This control defines the comment text for printout to trace 1 with a

maximum of 100 characters.

Valid Range: 100 characters

Default Value: "1st comment ..."

_2ndScreenComment

Variable Type       ViString

This control defines the comment text for printout to trace 2 with a

maximum of 100 characters.

Valid Range: 100 characters

Default Value: "2nd comment ..."

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Screen Title) out of range.

BFFC0003  Parameter 3 (1st Screen Comment) out of range.

BFFC0004  Parameter 4 (2nd Screen Comment) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confHCOpt

ViStatus rsfsu_confHCOpt (ViSession instrumentHandle, ViInt32 window,

ViInt32 output, ViInt32 quadrant_not_used,

ViBoolean color,

ViBoolean changeTraceColor_not_used);

Purpose

This function configures the object to print, how to print and the device

used to print.

Remote-control command(s):

HCOPy:DEVice:COLor ON | OFF

HCOPy:ITEM:ALL

HCOPy:ITEM:WINDow<1|2>:TRACe:STATe ON

HCOPy:ITEM:WINDow:TABLe:STATe ON

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window.

Valid Range:

1 - Window A (Default Value)

2 - Window B

output

Variable Type       ViInt32

This control selects the current hardcopy output.

Valid Range:

0 - Screen (Default Value)

1 - Trace

2 - Table

quadrant_not_used

Variable Type       ViInt32

This control selects the quadrant that is allocated to the screen

output.

Valid Range:

0 - Full Page

1 - Quadrant 1 (Default Value)

2 - Quadrant 2

3 - Quadrant 3

4 - Quadrant 4

Notes:

(1) Quadrant 1 = Top Right, Quadrant 2 = Top Left, Quadrant 3 =

Bottom Left, Quadrant 4 = Bottom Right.

(2) This control is reserved for compatibility with FSEx driver  and

does not apply to the instrument.

color

Variable Type       ViBoolean

This control selects between the color and monochrome hardcopy of the

screen.

Valid Range:

VI_TRUE  (1) -  On (Default Value)

VI_FALSE (0) - Off

changeTraceColor_not_used

Variable Type       ViBoolean

This control automatically changes the color of the currently

displayed trace after printout.

Valid Range:

VI_TRUE (1)  - ON

VI_FALSE (0) - OFF (Default)

Note:

(1) This control is reserved for compatibility with FSEx driver  and

does not apply to the instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Output) out of range.

BFFC0004  Parameter 4 (Quadrant (not used)) out of range.

BFFC0005  Parameter 5 (Color) out of range.

BFFC0006  Parameter 6 (Change Trace Color (not u) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_configHardcopyColors

ViStatus rsfsu_configHardcopyColors (ViSession instrumentHandle,

ViInt32 displayParameter,

ViInt32 colorMap,

ViInt32 predefinedColors,

ViReal64 TINT, ViReal64 SATURATION,

ViReal64 BRIGHTNESS);

Purpose

This function allows selection of colors for the printouts of screen

contents. The command tree is analoguous to the Config Display.

Remote-control command(s):

HCOPy:CMAP<1..26>:HSL

HCOPy:CMAP:DEFault<1|2>

HCOPy:CMAP<1..26>:PDEFined BLACk | BLUE | BROWn | GREen | CYAN | RED |

MAGenta | YELLow | WHITe | DGRAy | LGRAy | LBLUe | LGREen | LCYan | LRED

| LMAGenta

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

displayParameter

Variable Type       ViInt32

This control selects desired hardcopy parameter to be configured.

Valid Range:

0 - Brightness, Saturation, Tint (Default Value)

1 - Predefined Colors

2 - Default Colors (settings 1)

3 - Default Colors (settings 2)

Notes:

(1) Brightness, Saturation, Tint:

Defines the color table of the instrument.

(2) Predefined Colors:

Defines the color table of the instrument using predefined color

values.

(3) Default Colors (settings 1 or 2):

Resets the screen colors of all HCOPy items to their default

settings.

colorMap

Variable Type       ViInt32

This control defines the color table of the instrument.

Valid Range:

1 - Background

2 - Grid

3 - Function field + status field + data entry text

4 - Function field LED on

5 - Function field LED warn

6 - Enhancement label text

7 - Status field background

8 - Trace 1 (Default Value)

9 - Trace 2

10 - Trace 3

11 - Marker

12 - Lines

13 - Measurement status + limit check pass

14 - Limit check fail

15 - Table + softkey background

16 - Table + softkey text

17 - Table selected field text

18 - Table selected field background

19 - Table + data entry field opaq titlebar

20 - Data entry field opaq text

21 - Data entry field opaq background

22 - 3D shade bright part

23 - 3D shade dark part

24 - Softkey state on

25 - Softkey state data entry

26 - Logo

Note:

(1) FSU handles colour map different way than FSEx driver. Be paitent

when using this control. Below is the table with FSEx old values:

1 - Trace 1

2 - Trace 2

3 - Trace 3

4 - Trace 4

5 - Marker

6 - Grid

7 - Softkey State On

8 - Softkey State Data Entry

9 - Softkey State OFF

10 - Softkey Shade

11 - Text

12 - Title

13 - Background

(2) For default setting (1 or 2) has this control no effect.

predefinedColors

Variable Type       ViInt32

This control defines the color table of the instrument using

predefined color values. To each numeric suffix of CMAP is assigned

one or several picture elements which can be modified by varying the

corresponding color setting.

Valid Range:

0 - BLACk (Default Value)

1 - BLUE

2 - BROWn

3 - GREen

4 - CYAN

5 - RED

6 - MAGenta

7 - YELLow

8 - WHITe

9 - DGRAy

10 - LGRAy

11 - LBLUe

12 - LGREen

13 - LCYan

14 - LRED

15 - LMAGenta

TINT

Variable Type       ViReal64

This control sets the hue = TINT.

Valid Range: 0.0 to 100.0 %

Default Value: 30.0 %

Note:

(1) If (TINT > 1) then driver divides this value by 100.0 due to

compatibility with FSEx.

SATURATION

Variable Type       ViReal64

This control sets the sat = SATURATION.

Valid Range: 0.0 to 100.0 %

Default Value: 80.0 %

(1) If (SATURATION > 1) then driver divides this value by 100.0 due

to compatibility with FSEx.

BRIGHTNESS

Variable Type       ViReal64

This control sets the lum = BRIGHTNESS.

Valid Range: 0.0 to 100.0 %

Default Value: 100.0 %

(1) If (BRIGHTNESS > 1) then driver divides this value by 100.0 due

to compatibility with FSEx.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Display Parameter) out of range.

BFFC0003  Parameter 3 (Color Map) out of range.

BFFC0004  Parameter 4 (Predefined Colors) out of range.

BFFC0005  Parameter 5 (TINT) out of range.

BFFC0006  Parameter 6 (SATURATION) out of range.

BFFC0007  Parameter 7 (BRIGHTNESS) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_configureTraceIQAver

ViStatus rsfsu_configureTraceIQAver (ViSession instrumentHandle,

ViBoolean averageState,

ViInt32 averageCount);

Purpose

This function configures the IQ data measurement averaging.

Remote-control command(s):

TRACe:IQ:AVERage:STATe ON | OFF

TRACe:IQ:AVERage:COUNt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

averageState

Variable Type       ViBoolean

This command switches the averaging function for the recorded I/Q

data on or off. Preconditions are that the I/Q data recording

function has been switched on previously and that the sampling rate

has been set to 32 MHz.

Valid Ranges:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Note:

(1) Averaging is not supported for sampling rates <> 32 MHz or for

trigger offsets < 0.

averageCount

Variable Type       ViInt32

This command defines the number of averages to perform on the

recorded I/Q data.

Valid Ranges:

0 to 32767

Default Value:

0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Average State) out of range.

BFFC0003  Parameter 3 (Average Count) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLevelRange

ViStatus rsfsu_confLevelRange (ViSession instrumentHandle,

ViInt32 window, ViInt32 levelRange,

ViReal64 logManualValuedB);

Purpose

This function configures the Y axis range of the spectrum analyzer.

Notes:

(1) If the display is in Full Screen mode, the Active Screen must be

selected by the 'Window' control or an error will be generated. In Split

Screen mode either window may be selected, if the screens are Uncoupled

the selected window will be effected, otherwise both windows will be set.

Remote-control command(s):

DISPlay:WINDow<1|2>:TRACe:Y:SPACing LOGarithmic | LINear | LDB

DISPlay:WINDow<1|2>:TRACe:Y:SCALe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

levelRange

Variable Type       ViInt32

Selects the type of Y spacing for the selected display.

0 - Logarithmic Manual (Default Value)

1 - Linear / pct

2 - Linear / dB

Notes:

(1) If set to Logarithmic Manual, the display range is set by means

of control Log Manual Value.

logManualValuedB

Variable Type       ViReal64

This control selects the y scale in decibels.

Valid Range: 1.0 dB to 200.0 dB in steps of 10dB

Default Value: 100.0 dB

Notes:

(1) This control is active only if Level Range is set to Logarithmic

Manual.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Level Range) out of range.

BFFC0004  Parameter 4 (Log Manual Value [dB]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineData

ViStatus rsfsu_confLimitLineData (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViInt32 limitLineData,

ViInt32 no_ofValues,

ViReal64 _VI_FAR values[]);

Purpose

This function defines the X-axis values (frequencies) and Y-axis values

(levels) of the upper or lower limit lines.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:CONTrol[:DATA]

CALCulate<1|2>:LIMit<1 to 8>:LOWer[:DATA]

CALCulate<1|2>:LIMit<1 to 8>:UPPer[:DATA]

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

limitLineData

Variable Type       ViInt32

Selects limit line data to be defined.

Valid Range:

0 - X-Axis

1 - Y-Axis (Lower)

2 - Y-Axis (Upper)

Default Value: 0

no_ofValues

Variable Type       ViInt32

This control defines number of X-axis or Y-axis values.

Valid Range: > 0

Default Value: 0

values

Variable Type       ViReal64[]

This control defines the values for the selected lower limit line.

Valid Range: not range checked

Default Value: none

Notes:

(1) The number of values for the X-axis and for the corresponding

Y-axis (upper or lower) limit line has to be identical. Otherwise

default values are entered for missing values or unnecessary values

are deleted.

(2) If the measured values are smaller than the LOWer limit line, the

limit check signals errors.

(3) The unit must be identical with the unit selected by

CALC:LIM:UNIT. If no unit is indicated, the unit defined with

CALC:LIM:UNIT is automatically used.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Limit Line Data) out of range.

BFFC0005  Parameter 5 (No of Values) out of range.

BFFC0006  Parameter 6 (Values) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineDomain

ViStatus rsfsu_confLimitLineDomain (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViBoolean domain);

Purpose

This function defines the frequency or time domain for the x axis values.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:CONTrol:DOMain FREQuency | TIME

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

domain

Variable Type       ViBoolean

Defines the frequency or time domain for the x axis values.

Valid Range:

VI_FALSE (0) - Freq (Default Value)

VI_TRUE  (1) - Time

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Domain) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineMargin

ViStatus rsfsu_confLimitLineMargin (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViBoolean limitLine,

ViReal64 margin);

Purpose

This function defines a margin to a lower (upper) limit line, at which

out-of-limit values are signalled (if the limit check is active), but not

handled as a violation of the limit value. The margin is independent of

the measurement window.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:LOWer:MARGin

CALCulate<1|2>:LIMit<1 to 8>:UPPer:MARGin

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

limitLine

Variable Type       ViBoolean

This control selects limit line to be configured.

Valid Range:

VI_FALSE (0) - Lower (Default Value)

VI_TRUE  (1) - Upper

margin

Variable Type       ViReal64

Defines a margin to a lower (upper) limit line.

Valid Range: none

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Limit Line) out of range.

BFFC0005  Parameter 5 (Margin) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineMeasType

ViStatus rsfsu_confLimitLineMeasType (ViSession instrumentHandle,

ViInt32 window,

ViInt32 limitLineNo,

ViInt32 measType);

Purpose

This function assigns a limit line to a particular measurement type.

Measurement types include noise figure, noise temperature or gain.

Note:

This function is available only in conjunction with option FS-K30.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 6>:TRACe NFIGure | TEFFective | GAIN

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

measType

Variable Type       ViInt32

This parameter assigns a limit line to a particular measurement type.

Valid Range:

0 - Noise Figure

1 - Noise Temperature

2 - Gain

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Meas Type) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineMode

ViStatus rsfsu_confLimitLineMode (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViInt32 limitLine, ViBoolean mode);

Purpose

This function selects the relative or absolute scaling for the X(Y)-axis

of the selected limit line. The definition is independent of the

measurement window.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:CONTrol:MODE RELative | ABSolute

CALCulate<1|2>:LIMit<1 to 8>:LOWer:MODE RELative | ABSolute

CALCulate<1|2>:LIMit<1 to 8>:UPPer:MODE RELative | ABSolute

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

limitLine

Variable Type       ViInt32

Selects limit line.

Valid Range:

0 - X-Axis

1 - Y-Axis (Lower)

2 - Y-Axis (Upper)

Default Value: 0

mode

Variable Type       ViBoolean

Selects the relative or absolute scaling for the X(Y)-axis of the

selected (control, upper, lower) limit line.

Valid Range:

VI_FALSE (0) - Relative

VI_TRUE  (1) - Absolute (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Limit Line) out of range.

BFFC0005  Parameter 5 (Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineOffset

ViStatus rsfsu_confLimitLineOffset (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViInt32 limitLine,

ViReal64 offsetValue);

Purpose

This function defines an offset for the X(Y) axis value of the selected

relative limit line in the frequency or time domain.

An offset for the Y axis of the selected relative lower (upper) limit

line in contrast to CALC:LIM:LOW:SHIFt (CALC:LIM:UPPer:SHIFt), is not

shifted by modifying the individual Y values but by means of an additive

offset.

The offset is independent of the measurement window.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:CONTrol:OFFSet

CALCulate<1|2>:LIMit<1 to 8>:LOWer:OFFSet

CALCulate<1|2>:LIMit<1 to 8>:UPPer:OFFSet

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

limitLine

Variable Type       ViInt32

Selects limit line to apply offset value.

Valid Range:

0 - X axis (Default Value)

1 - Y axis (Lower)

2 - Y axis (Upper)

offsetValue

Variable Type       ViReal64

Defines an offset for the X(Y) axis value. The unit of values depends

on the frequency or time domain of the X axis.

Valid Range: none

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Limit Line) out of range.

BFFC0005  Parameter 5 (Offset Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineParam

ViStatus rsfsu_confLimitLineParam (ViSession instrumentHandle,

ViInt32 window, ViInt32 traceNumber,

ViInt32 limitLineNo,

ViString limitLineName,

ViString comment,

ViBoolean spacing_not_used,

ViBoolean modeX, ViBoolean modeY,

ViBoolean ctrlUnit_not_used,

ViBoolean domain,

ViBoolean lineParameter,

ViInt32 units, ViInt32 no_ofValues,

ViReal64 _VI_FAR xaxisValues[],

ViReal64 _VI_FAR yaxisValues[],

ViReal64 margin);

Purpose

This function sets the X-axis (CONTrol) and Y-axis (upper and/or lower)

limit line parameters.

Note:

(1) This function is reserved for compatibility with FSEx driver. Prefer

to use function rsfsu_confLimitLineParamExt.

Remote-control command(s):

CALCulate<1|2>:LIMit<1..8>:NAME

CALCulate<1|2>:LIMit<1..8>:DELete

CALCulate<1|2>:LIMit<1..8>:CONTrol:DOMain FREQuency | TIME

CALCulate<1|2>:LIMit<1..8>:CONTrol:MODE RELative | ABSolute

CALCulate<1|2>:LIMit<1..8>:CONTrol:UNIT SYM | S

CALCulate<1|2>:LIMit<1..8>:CONTrol:DATA

CALCulate<1|2>:LIMit<1..8>:COMMent

CALCulate<1|2>:LIMit<1..8>:UPPer:MODE RELative | ABSolute

CALCulate<1|2>:LIMit<1..8>:UPPer:SPACing LINear | LOGarithmic

CALCulate<1|2>:LIMit<1..8>:UPPer:MARGin

CALCulate<1|2>:LIMit<1..8>:UPPer:DATA

CALCulate<1|2>:LIMit<1..8>:LOWer:MODE RELative | ABSolute

CALCulate<1|2>:LIMit<1..8>:LOWer:MARGin

CALCulate<1|2>:LIMit<1..8>:LOWer:DATA

CALCulate<1|2>:LIMit<1..8>:TRACe

CALCulate<1|2>:LIMit<1..8>:UPPer:STATE ON

CALCulate<1|2>:LIMit<1..8>:LOWer:STATE ON

CALCulate<1|2>:LIMit<1..8>:UNIT DBM | DBPW | WATT | DBUV | VOLT | DBUA |

AMPere | DB | DBUV_M | DBUA_M  | DEG | RAD | S | HZ | PCT | UNITLESS |

DBMV

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

traceNumber

Variable Type       ViInt32

This control assigns a trace to a limit line.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

limitLineName

Variable Type       ViString

This control assigns a name to a limit numbered 1 to 8. If doesn't

exist previously, a limit line with this name is created.

Valid Range: up to 8 characters

Default Value: "Limit1"

Note:

(1) For Copy operation, this is the target location.

comment

Variable Type       ViString

This control defines a comment for the limit line selected.

Valid Range: not tested

Default Value: "Comment ..."

spacing_not_used

Variable Type       ViBoolean

This control makes a selection between linear and logarithmic

interpolation for the lower limit line.

Valid Range:

VI_FALSE (0) - Logarithmic

VI_TRUE  (1) - Linear (Default Value)

Note:

(1) This control is reserved for compatibility with FSEx driver  and

does not apply to the instrument.

modeX

Variable Type       ViBoolean

This command selects the relative or absolute scaling for the X-axis

of the selected control (upper, lower) limit line.

Valid Range:

VI_FALSE (0) - Relative (Default Value)

VI_TRUE  (1) - Absolute

modeY

Variable Type       ViBoolean

This command selects the relative or absolute scaling for the Y-axis

of the selected control (upper, lower) limit line.

Valid Range:

VI_FALSE (0) - Relative (Default Value)

VI_TRUE  (1) - Absolute

Notes:

(1) Upon selection of RELative, the unit is switched to DB.

ctrlUnit_not_used

Variable Type       ViBoolean

This control defines the unit of the x-axis scaling of limit lines.

Valid Range:

VI_FALSE (0) - S (Default Value)

VI_TRUE  (1) - SYM

Note:

(1) SYM is available only for vector analysis mode.

(2) This control is reserved for compatibility with FSEx driver  and

does not apply to the instrument.

domain

Variable Type       ViBoolean

This control defines the X-axis in the frequency or time domain.

Valid Range:

VI_FALSE (0) - Freq (Default Value)

VI_TRUE  (1) - Time

lineParameter

Variable Type       ViBoolean

This control selects which set of parameters will be actually

configured.

Valid Range:

VI_FALSE (0) - Lower (Default Value)

VI_TRUE  (1) - Upper

units

Variable Type       ViInt32

This control selects  units for the limit line operations.

Valid Range:

0 - DBM  (Default)

1 - DBPW

2 - WATT

3 - DBUV

4 - VOLT

5 - DBUA

6 - AMP

7 - DB

8 - DBUV MHZ

9 - DBUA MHZ

10 - DEG

11 - RAD

12 - S

13 - HZ

14 - PCT

15 - UNITLESS

Notes:

(1) DBUV_MHZ and DBUA_MHZ denote the units DBUV/MHZ or DBUA/MHZ. Upon

selection of the unit DB the limit line is automatically switched to

the relative mode. For units different from DB the limit line is

automatically switched to the absolute mode. The units DEG, RAD, S,

HZ, PCT are available in the signal analysis mode only.

(2) Control Units are sent only in Time Domain mode (SPAN=0).

no_ofValues

Variable Type       ViInt32

This control defines number of X-axis or Y-axis values.

Valid Range: 0 to 50

Default Value: 0

xaxisValues

Variable Type       ViReal64[]

This command defines the X-axis values (frequencies or times) of the

control limit lines.

Valid Range: not range checked

Default Value: none

Notes:

(1) The number of values for the CONTrol axis and the corresponding

UPPer and/or LOWer limit lines have to be identical.

yaxisValues

Variable Type       ViReal64[]

This command defines the Y-axis values (levels or linear units such

as volts) of the upper or lower limit lines.

Valid Range: not range checked

Default Value: none

Notes:

(1) The number of values for the CONTrol axis and the corresponding

UPPer and/or LOWer limit lines have to be identical.

margin

Variable Type       ViReal64

This command defines the margin of the selected upper (lower) limit

line.

Valid Range: not range checked

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Trace Number) out of range.

BFFC0004  Parameter 4 (Limit Line No.) out of range.

BFFC0005  Parameter 5 (Limit Line Name) out of range.

BFFC0006  Parameter 6 (Comment) out of range.

BFFC0007  Parameter 7 (Spacing (not used)) out of range.

BFFC0008  Parameter 8 (Mode X) out of range.

BFFC0009  Parameter 9 (Mode Y) out of range.

BFFC000A  Parameter 10 (Ctrl Unit (not used)) out of range.

BFFC000B  Parameter 11 (Domain) out of range.

BFFC000C  Parameter 12 (Line Parameter) out of range.

BFFC000D  Parameter 13 (Units) out of range.

BFFC000E  Parameter 14 (No of Values) out of range.

BFFC000F  Parameter 15 (X-axis Values) out of range.

BFFC0010  Parameter 16 (Y-axis Values) out of range.

BFFC0011  Parameter 17 (Margin) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineParamExt

ViStatus rsfsu_confLimitLineParamExt (ViSession instrumentHandle,

ViInt32 window,

ViInt32 traceNumber,

ViInt32 limitLineNo,

ViString limitLineName,

ViString comment, ViBoolean modeX,

ViBoolean modeY,

ViBoolean ctrlUnit_RESERVED,

ViBoolean domain,

ViBoolean lineParameter,

ViInt32 units,

ViInt32 no_ofValues,

ViReal64 _VI_FAR xaxisValues[],

ViReal64 _VI_FAR yaxisValues[],

ViReal64 thresholdValue,

ViReal64 margin);

Purpose

This function sets the X-axis (CONTrol) and Y-axis (upper and/or lower)

limit line parameters.

Remote-control command(s):

CALCulate<1|2>:LIMit<1|8>:NAME

CALCulate<1|2>:LIMit<1|8>:DELete

CALCulate<1|2>:LIMit<1|8>:CONTrol:DOMain FREQuency | TIME

CALCulate<1|2>:LIMit<1|8>:CONTrol:MODE RELative | ABSolute

CALCulate<1|2>:LIMit<1|8>:CONTrol:DATA

CALCulate<1|2>:LIMit<1|8>:COMMent

CALCulate<1|2>:LIMit<1|8>:TRACe 1 | 2 | 3

CALCulate<1|2>:LIMit<1|8>:UNIT DBM | DBPW | WATT | DBUV | VOLT | DBUA |

AMP | DB | DBUV_M | DBUA_M | DEG | RAD | S | HZ | PCT | UNITLESS | DBMV

CALCulate<1|2>:LIMit<1|8>UPPer:MODE RELative | ABSolute

CALCulate<1|2>:LIMit<1|8>UPPer:STATE ON

CALCulate<1|2>:LIMit<1|8>UPPer:THReshold

CALCulate<1|2>:LIMit<1|8>UPPer:MARGin

CALCulate<1|2>:LIMit<1|8>UPPer:DATA

CALCulate<1|2>:LIMit<1|8>LOWer:MODE RELative | ABSolute

CALCulate<1|2>:LIMit<1|8>LOWer:STATE ON

CALCulate<1|2>:LIMit<1|8>LOWer:THReshold

CALCulate<1|2>:LIMit<1|8>LOWer:MARGin

CALCulate<1|2>:LIMit<1|8>LOWer:DATA

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

traceNumber

Variable Type       ViInt32

This control assigns a trace to a limit line.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

limitLineName

Variable Type       ViString

This control assigns a name to a limit numbered 1 to 8. If doesn't

exist previously, a limit line with this name is created.

Valid Range: up to 8 characters

Default Value: "Limit1"

Note:

(1) For Copy operation, this is the target location.

comment

Variable Type       ViString

This control defines a comment for the limit line selected.

Valid Range: not tested

Default Value: "Comment ..."

modeX

Variable Type       ViBoolean

This command selects the relative or absolute scaling for the X-axis

of the selected control (upper, lower) limit line.

Valid Range:

VI_FALSE (0) - Relative (Default Value)

VI_TRUE  (1) - Absolute

modeY

Variable Type       ViBoolean

This command selects the relative or absolute scaling for the Y-axis

of the selected control (upper, lower) limit line.

Valid Range:

VI_FALSE (0) - Relative (Default Value)

VI_TRUE  (1) - Absolute

Notes:

(1) Upon selection of RELative, the unit is switched to DB.

ctrlUnit_RESERVED

Variable Type       ViBoolean

This control defines the unit of the x-axis scaling of limit lines.

Valid Range:

VI_FALSE (0) - S (Default Value)

VI_TRUE  (1) - SYM

Note:

(1) SYM is available only for vector analysis mode.

(2) This control is reserved for future use and does not apply to the

instrument.

domain

Variable Type       ViBoolean

This control defines the X-axis in the frequency or time domain.

Valid Range:

VI_FALSE (0) - Freq (Default Value)

VI_TRUE  (1) - Time

lineParameter

Variable Type       ViBoolean

This control selects which set of parameters will be actually

configured.

Valid Range:

VI_FALSE (0) - Lower (Default Value)

VI_TRUE  (1) - Upper

units

Variable Type       ViInt32

This control selects  units for the limit line operations.

Valid Range:

0 - DBM

1 - DBPW

2 - WATT

3 - DBUV

4 - VOLT

5 - DBUA

6 - AMP

7 - DB (Default Value)

8 - DBUV MHZ

9 - DBUA MHZ

10 - DEG

11 - RAD

12 - S

13 - HZ

14 - PCT

15 - UNITLESS

16 - DBMV

Notes:

(1) DBUV_MHZ and DBUA_MHZ denote the units DBUV/MHZ or DBUA/MHZ. Upon

selection of the unit DB the limit line is automatically switched to

the relative mode. For units different from DB the limit line is

automatically switched to the absolute mode. The units DEG, RAD, S,

HZ, PCT are available in the signal analysis mode only.

(2) Control Units are sent only in Time Domain mode (SPAN=0).

no_ofValues

Variable Type       ViInt32

This control defines number of X-axis or Y-axis values.

Valid Range: 0 to 50

Default Value: 0

xaxisValues

Variable Type       ViReal64[]

This command defines the X-axis values (frequencies or times) of the

control limit lines.

Valid Range: not range checked

Default Value: none

Notes:

(1) The number of values for the CONTrol axis and the corresponding

UPPer and/or LOWer limit lines have to be identical.

yaxisValues

Variable Type       ViReal64[]

This command defines the Y-axis values (levels or linear units such

as volts) of the upper or lower limit lines.

Valid Range: not range checked

Default Value: none

Notes:

(1) The number of values for the CONTrol axis and the corresponding

UPPer and/or LOWer limit lines have to be identical.

thresholdValue

Variable Type       ViReal64

This control defines an absolute threshold value for limit lines with

relative Y-axis scaling independently of the measurement window. The

absolute threshold value is used in the limit check

as soon as it exceeds the relative limit value.

Valid Range: not range checked

Default Value: 0.0

Note:

The unit should correspond to the unit selected with control Units

(except dB which is not permissible). If no unit is indicated, the

unit defined with control Units is automatically used (exception: dBm

instead of dB).

margin

Variable Type       ViReal64

This command defines the margin of the selected upper (lower) limit

line.

Valid Range: not range checked

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Trace Number) out of range.

BFFC0004  Parameter 4 (Limit Line No.) out of range.

BFFC0005  Parameter 5 (Limit Line Name) out of range.

BFFC0006  Parameter 6 (Comment) out of range.

BFFC0007  Parameter 7 (Mode X) out of range.

BFFC0008  Parameter 8 (Mode Y) out of range.

BFFC0009  Parameter 9 (Ctrl Unit (RESERVED)) out of range.

BFFC000A  Parameter 10 (Domain) out of range.

BFFC000B  Parameter 11 (Line Parameter) out of range.

BFFC000C  Parameter 12 (Units) out of range.

BFFC000D  Parameter 13 (No of Values) out of range.

BFFC000E  Parameter 14 (X-axis Values) out of range.

BFFC000F  Parameter 15 (Y-axis Values) out of range.

BFFC0010  Parameter 16 (Threshold Value) out of range.

BFFC0011  Parameter 17 (Margin) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineShift

ViStatus rsfsu_confLimitLineShift (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViInt32 limitLineShift,

ViReal64 shiftValue);

Purpose

This function moves a limit line by the indicated value in x (y)

direction. The line is shifted by modifying the individual x (y) values.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:CONTrol:SHIFt

CALCulate<1|2>:LIMit<1 to 8>:LOWer:SHIFt

CALCulate<1|2>:LIMit<1 to 8>:UPPer:SHIFt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

limitLineShift

Variable Type       ViInt32

Selects direction where to moves a limit line by the indicated value.

Valid Range:

0 - X-Direction (Default Value)

1 - Y-Direction (Lower)

2 - Y-Direction (Upper)

shiftValue

Variable Type       ViReal64

Moves a limit line by the indicated value in selected direction.

Valid Range: none

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Limit Line Shift) out of range.

BFFC0005  Parameter 5 (Shift Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineState

ViStatus rsfsu_confLimitLineState (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViInt32 limitLineOperation,

ViString limitLineParameter);

Purpose

This function operates with limit lines. The limit line subsystem

comprises the limit lines and the corresponding limit checks. This

function allows to switch limit line operations.

Remote-control command(s):

CALCulate<1|2>:LIMit<1..8>:NAME

CALCulate<1|2>:LIMit<1..8>:COPY

CALCulate<1|2>:LIMit<1..8>:DELete

CALCulate<1|2>:LIMit<1..8>:STATe OFF | ON

CALCulate<1|2>:LIMit<1..8>:COMMent

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

limitLineOperation

Variable Type       ViInt32

This control turns limit checking On, Off or manipulates with limit

lines.

Valid Range:

0 - Off (Default Value)

1 - On

2 - Copy

3 - Delete

4 - Comment

Notes:

(1) Off:

Turns limit line checking to Off.

(2) On:

Turns limit line checking to On.

(3) Copy:

For Copy operation, limit line no. is source location and limit line

parameter is target location.

(4) Delete:

Deletes selected limit line (limit line name is not significant).

(5) Comment:

Defines a comment for the limit line selected.

limitLineParameter

Variable Type       ViString

This control assigns a name to a limit line numbered 1 to 8. If

doesn't exist previously, a limit line with this name is created.

Valid Range: > 0

Default Value: "REM1"

Notes:

(1) For Copy operation, this is the target location.

(2) For Comment operation this parameter defines a comment for the

limit line selected.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Limit Line Operation) out of range.

BFFC0005  Parameter 5 (Limit Line Parameter) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineSwitch

ViStatus rsfsu_confLimitLineSwitch (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViBoolean limitLine,

ViBoolean state);

Purpose

This function switches on or off the upper or lower limit lines.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:LOWer:STATe ON | OFF

CALCulate<1|2>:LIMit<1 to 8>:UPPer:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

limitLine

Variable Type       ViBoolean

This control selects limit line to be configured.

Valid Range:

VI_FALSE (0) - Lower (Default Value)

VI_TRUE  (1) - Upper

state

Variable Type       ViBoolean

Switches on or off the indicated limit line.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Limit Line) out of range.

BFFC0005  Parameter 5 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineThreshold

ViStatus rsfsu_confLimitLineThreshold (ViSession instrumentHandle,

ViInt32 window,

ViInt32 limitLineNo,

ViBoolean limitLine,

ViReal64 threshold);

Purpose

This function defines an absolute threshold value for limit lines with

relative Y axis scaling independently of the measurement window. The

absolute threshold value is used in the limit check as soon as it exceeds

the relative limit value.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:LOWer:THReshold

CALCulate<1|2>:LIMit<1 to 8>:UPPer:THReshold

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

limitLine

Variable Type       ViBoolean

This control selects limit line to be configured.

Valid Range:

VI_FALSE (0) - Lower (Default Value)

VI_TRUE  (1) - Upper

threshold

Variable Type       ViReal64

Defines an absolute threshold value.

The unit must correspond to the unit selected with CALC:LIM:UNIT

(except dB which is not allowed). If no unit is indicated, the unit

defined with CALC:LIM:UNIT is automatically used (exception: dBm

instead of dB).

The units DEG, RAD, S, HZ, PCT are not available in the SPECTRUM

mode.

Valid Range: none

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Limit Line) out of range.

BFFC0005  Parameter 5 (Threshold) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineTrace

ViStatus rsfsu_confLimitLineTrace (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViInt32 traceNumber);

Purpose

This function assigns a limit line to a trace in the indicated

measurement window.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:TRACe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

traceNumber

Variable Type       ViInt32

This control assigns a trace to a limit line.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Trace Number) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confLimitLineUnits

ViStatus rsfsu_confLimitLineUnits (ViSession instrumentHandle,

ViInt32 window, ViInt32 limitLineNo,

ViInt32 units);

Purpose

This function defines the unit of the selected limit line. The definition

is valid independently of the measurement window.

Remote-control command(s):

CALCulate<1|2>:LIMit<1 to 8>:UNIT DBM | DBPW | WATT | DBUV | DBMV | VOLT

|DBUA | AMPere | DB | DEG | RAD | S | HZ | PCT | UNITLESS | DBUA_M |

DBUV_M

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

limitLineNo

Variable Type       ViInt32

Selects the limit line number (which is associated with limit line

name).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

5 - 5

6 - 6

7 - 7

8 - 8

units

Variable Type       ViInt32

This control selects  units for the limit line operations.

Valid Range:

0 - DBM

1 - DBPW

2 - WATT

3 - DBUV

4 - VOLT

5 - DBUA

6 - AMP

7 - DB

8 - DBUV MHZ

9 - DBUA MHZ

10 - DEG

11 - RAD

12 - S

13 - HZ

14 - PCT

15 - UNITLESS

16 - DBMV

Default Value: 7

Notes:

(1) Upon selection of the unit DB the limit line is automatically

switched to the relative mode. For units different from DB the limit

line is automatically switched to absolute mode.

(2) The units DEG, RAD, S, HZ, PCT are not available in SPECTRUM

mode.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Limit Line No.) out of range.

BFFC0004  Parameter 4 (Units) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confListPwrSeq

ViStatus rsfsu_confListPwrSeq (ViSession instrumentHandle,

ViInt32 window, ViInt32 no_ofListItems,

ViReal64 _VI_FAR analyzerFrequencyHz[],

ViReal64 _VI_FAR referenceLeveldBm[],

ViReal64 _VI_FAR RFInputAttenuationdB[],

ViReal64 _VI_FAR RFInputElectronicAttndB[],

ViInt32 _VI_FAR filterType[],

ViReal64 _VI_FAR resolutionBandwidthHz[],

ViReal64 _VI_FAR videoBandwidthHz[],

ViReal64 _VI_FAR measTimes[],

ViReal64 _VI_FAR triggerLevel[]);

Purpose

This function configures the list of settings (max. 100 entries for an

individual frequency point) for the multiple power measurement and starts

a measurement sequence. To reduce the setting time, all indicated

parameters are set up simultaneously at each test point.

Notes:

(1) Each one test point is defined on the same position in all parameter

arrays.

(2) The measurement is performed in the time domain and therefore the

span is set to 0 Hz. If the time domain is quit, the function is

automatically switched off. The measurement is not compatible with other

measurements, especially as far as marker, adjacent channel power

measurement or statistics are concerned. The corresponding commands thus

automatically deactivate the function. The function is only available in

REMOTE operation. It is deactivated when switching the instrument back to

LOCAL.

Remote-control command(s):

SENSe<1|2>:LIST:POWer:SEQuence

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

no_ofListItems

Variable Type       ViInt32

Set number of list items. This value must be equal to array size of

each list parameter.

Valid Range: 1 to 100

Default Value: 1

analyzerFrequencyHz

Variable Type       ViReal64[]

Array of receive frequency for the signal to be measured (= center

frequency in manual operation).

Valid Range: 0.0 Hz to max. analyzer frequency, depending on the

instrument model.

Default Value: none

referenceLeveldBm

Variable Type       ViReal64[]

Array of reference level values for each list item.

Valid Range: -70.0 dBm to 30.0 dBm in 10 dB steps

-75.0 dBm to 30.0 dBm in 5 dB steps (Electronic

Attenuator option FS-B25)

Default Value: none

RFInputAttenuationdB

Variable Type       ViReal64[]

Array of RF input attenuation values for each list item.

Valid Range: 0.0 dB to 70.0 dB in 10 dB steps

0.0 dB to 75.0 dB in 5 dB steps (Electronic

Attenuator Option FS-B25)

Default Value: none

RFInputElectronicAttndB

Variable Type       ViReal64[]

Array of RF input attenuation values of electronic attenuator for

each list item.

Valid Range: 0.0 dB to 30.0 dB in 10 dB steps

Default Value: none

Note:

If value is < 0.0 (negative value) then electronic attenuator is not

in the signal path (OFF). If option FS-B25 is missing, -1 (OFF) is to

be given.

filterType

Variable Type       ViInt32[]

Array of filter types for each list item.

Valid Range:

0 - Normal

1 - CFilter

2 - RRC

Default Value: none

Notes:

(1) Normal:

Normal resolution filter.

(2) CFilter:

Channel filter. These are especially steep-edged filters, which are

used for example in Fast ACP measurement to ensure the band-limiting

of a transmission channel in the time domain.

(3) RRC:

Root Raised Cosine filter. This special filter form is used to

determine the channel power for some mobile radio standards.

resolutionBandwidthHz

Variable Type       ViReal64[]

Array of resolution bandwidth values for each list item.

Valid Range: 10.0 Hz to 10.0e6 Hz in 1, 3, 10 steps

Default Value: 0.0 Hz

Note:

Valid range also depends on the selected filter type.

videoBandwidthHz

Variable Type       ViReal64[]

Array of video bandwidth values for each list item.

Valid Range: 1.0 Hz to 10.0e6 Hz in 1, 3, 10 steps.

Default Value: none

Note:

The value is ignored for <filter type> = CFilter or RRC.

measTimes

Variable Type       ViReal64[]

Array of measurement time values for each list item.

Valid Range: 1.0e-6 s to 30.0 s

Default Value: none

triggerLevel

Variable Type       ViReal64[]

Array of trigger level values for each list item.

This value is Reserved. Must be set to 0 for each list item.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (No of List Items) out of range.

BFFC0004  Parameter 4 (Analyzer Frequency [Hz]) out of range.

BFFC0005  Parameter 5 (Reference Level [dBm]) out of range.

BFFC0006  Parameter 6 (RF Input Attenuation [dB]) out of range.

BFFC0007  Parameter 7 (RF Input Electronic Attn ) out of range.

BFFC0008  Parameter 8 (Filter Type) out of range.

BFFC0009  Parameter 9 (Resolution Bandwidth [Hz]) out of range.

BFFC000A  Parameter 10 (Video Bandwidth [Hz]) out of range.

BFFC000B  Parameter 11 (Meas Time [s]) out of range.

BFFC000C  Parameter 12 (Trigger Level) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confListPwrSet

ViStatus rsfsu_confListPwrSet (ViSession instrumentHandle,

ViInt32 window, ViBoolean peakMeas,

ViBoolean RMSMeas, ViBoolean AVGMeas,

ViInt32 triggerMode,

ViBoolean triggerSlope,

ViReal64 triggerOffset,

ViReal64 gateLength);

Purpose

This function defines the constant settings for the list during multiple

power measurement. Parameters Peak Meas, RMS Meas and AVG Meas define,

which measurements are to be performed at the same time at the frequency

point. Correspondingly, one, two or three results per frequency point are

returned for the function Get List Power Sequence (actListPwrSeq). If all

three parameters are set to OFF, the command generates an execution

error.

Remote-control command(s):

SENSe<1|2>:LIST:POWer:SET

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

peakMeas

Variable Type       ViBoolean

Activates (deactivates) the measurement of the peak power (peak

detector).

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

RMSMeas

Variable Type       ViBoolean

Activates (deactivates) the measurement of the RMS power (RMS

detector).

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

AVGMeas

Variable Type       ViBoolean

Activates (deactivates) the measurement of the average power (average

detector).

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

triggerMode

Variable Type       ViInt32

Selection of the trigger source used for the list measurement.

Valid Range:

0 - Immediate (Default Value)

1 - Video

2 - External

3 - IF Power

triggerSlope

Variable Type       ViBoolean

This control selects the slope of the trigger signal.

Valid Range:

VI_FALSE (0) - Negative

VI_TRUE  (1) - Positive (Default Value)

triggerOffset

Variable Type       ViReal64

This control sets the offset between the detection of the trigger

signal and the start of the measurement at the next frequency point.

Valid Range: 0.0 s (0.125 s) to 100s

Notes:

(1) The value 0 s deactivates the use of GATED TRIGGER; other

values activate the GATED TRIGGER function.

(2) Values <> 0 s are only possible if Trigger Mode is different from

IMMediate. Otherwise, an execution error is triggered.

gateLength

Variable Type       ViReal64

Gate length with Gated Sweep.

Valid Range: 0.0 s (0.125 s) to 100.0 s

Notes:

(1) The value 0 s deactivates the use of GATED TRIGGER; other

values activate the GATED TRIGGER function.

(2) Values <> 0 s are only possible if Trigger Mode is different from

IMMediate. Otherwise, an execution error is triggered.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Peak Meas) out of range.

BFFC0004  Parameter 4 (RMS Meas) out of range.

BFFC0005  Parameter 5 (AVG Meas) out of range.

BFFC0006  Parameter 6 (Trigger Mode) out of range.

BFFC0007  Parameter 7 (Trigger Slope) out of range.

BFFC0008  Parameter 8 (Trigger Offset) out of range.

BFFC0009  Parameter 9 (Gate Length) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confListPwrState

ViStatus rsfsu_confListPwrState (ViSession instrumentHandle,

ViInt32 window,

ViBoolean listMeasurement);

Purpose

This function deactivates the list measurement.

Remote-control command(s):

SENSe<1|2>:LIST:POWer:STATe OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

listMeasurement

Variable Type       ViBoolean

This control deactivates the list measurement.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Note(s):

(1) This control is not used.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (List Measurement) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confMarkOpt

ViStatus rsfsu_confMarkOpt (ViSession instrumentHandle, ViInt32 window,

ViInt32 markerNumber, ViInt32 trace,

ViBoolean deleteAllMkrs);

Purpose

This function selects the trace for a marker or deletes all markers.

Remote-control command(s):

CALCulate<1|2>:MARKer:AOFF

CALCulate<1|2>:MARKer<1..4>:TRACe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

trace

Variable Type       ViInt32

Assigns the selected marker to the indicated test curve.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

deleteAllMkrs

Variable Type       ViBoolean

Deletes all markers, if set to VI_TRUE.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On (Delete all markers)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Trace) out of range.

BFFC0005  Parameter 5 (Delete All Mkrs) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confMarkPosition

ViStatus rsfsu_confMarkPosition (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViReal64 markerPosition);

Purpose

This function enables a marker and sets the marker to the specified

position (frequency or time in sec).

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:X

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

markerPosition

Variable Type       ViReal64

Sets the selected marker to the specified position.

Valid Range: not checked

Default Value: 3.0E9

Notes:

(1) Valid Range depends on selected horizontal display unit (analyzer

mode: freq. domain (Hz), time domain (s)).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Marker Position) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confMarkState

ViStatus rsfsu_confMarkState (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViBoolean markerState);

Purpose

This function enables or disables a marker. When the marker is switched

on, a peak search is done automatically.

Remote-control command(s):

CALCulate<1|2|3>:MARKer<1..4>:STATe OFF | ON

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 3

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

markerState

Variable Type       ViBoolean

Enables or disables the selected marker.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Marker State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confMSModType

ViStatus rsfsu_confMSModType (ViSession instrumentHandle,

ViBoolean modulationType);

Purpose

This function selects the modulation type (GMSK or EDGE (8PSK)).

Note:

(1) This function is available only in conjunction with option FS-K5.

Remote-control command(s):

CONFigure:MTYPe GMSK | EDGE

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

modulationType

Variable Type       ViBoolean

This control selects the modulation type (GMSK or EDGE (8PSK)).

Valid Range:

VI_FALSE (0) - GMSK (Default Value)

VI_TRUE  (1) - EDGE

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Modulation Type) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confMSUserMid

ViStatus rsfsu_confMSUserMid (ViSession instrumentHandle,

ViString trainingSequence);

Purpose

This function defines the bit pattern of the user midamble (training

sequence). For each modulation type, a separate user midamble can be

defined. If you want to switch on the user midamble, use MS Set Channel

function.

Note:

(1) This function is available only in conjunction with option FS-K5.

Remote-control command(s):

CONFigure:CHANnel:TSC:USER

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

trainingSequence

Variable Type       ViString

This control defines training sequence.

Valid Range: see notes below

Default Value: "00000000000000000000000000"

Notes:

(1) For modulation type GMSK:

The first 26 characters are evaluated. �0� and �1� represent the GMSK

symbols -1 and 1. If the number of characters is insufficient, the

string is filled with '0' up to the 26 th character. Characters other

than �0� are processed as �1�.

(2) For modulation type 8PSK (EDGE):

The first 78 characters are evaluated. The character patterns are

assigned to the 8PSK (EDGE) symbols as follows:

Pattern 8PSK Symbol

111 0

011 1

010 2

000 3

001 4

101 5

100 6

110 7

If the number of characters is insufficient, the string is filled

with '0' up to the 78 th character . Characters other than �0� are

processed as �1�.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Training Sequence) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoise2ndStageCorrection

ViStatus rsfsu_confNoise2ndStageCorrection (ViSession instrumentHandle);

Purpose

This function configures R&S FS-K30 for a second stage correction

measurement. After this function has been executed the second stage

correction measurement will be the measurement started when the user

issues the INITiate command.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

CONFigure:CORRection

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoise2ndStageCorrectionState

ViStatus rsfsu_confNoise2ndStageCorrectionState

(ViSession instrumentHandle, ViBoolean correction_state);

Purpose

This function can be used to control whether the R&S FS-K30 option will

factor in the results obtained from second stage correction when

performing a noise and gain measurement.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CORRection[:STATe] ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

correction_state

Variable Type       ViBoolean

Selects whether the R&S FS-K30 option will factor in the results

obtained from second stage correction when performing a noise and

gain measurement.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Correction state) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseAverage

ViStatus rsfsu_confNoiseAverage (ViSession instrumentHandle,

ViInt32 average);

Purpose

This function specifies the number of sweeps over which the measurement

results are to be averaged. The higher the number of sweeps averaged the

more accurate the results however the more sweeps required the longer the

measurement takes to perform.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]SWEep:COUNt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

average

Variable Type       ViInt32

Number of sweeps over which the measurement results are to be

averaged.

Valid Range:

1 to 32767

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Average) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseDUTRange

ViStatus rsfsu_confNoiseDUTRange (ViSession instrumentHandle,

ViReal64 DUTRange);

Purpose

This function allows the maximum gain of the DUT to be specified.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

SYSTem:CONFigure:DUT:GAIN

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

DUTRange

Variable Type       ViReal64

Maximum gain of the DUT.

Valid Range:

10 to 999.99

Default Value: 30.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (DUT Range) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseDUTSettlingTime

ViStatus rsfsu_confNoiseDUTSettlingTime (ViSession instrumentHandle,

ViReal64 DUTSettlingTime);

Purpose

This function allows the DUT settling time to be modified. It represents

the time to wait for the DUT to settle after a noise source has been

turned on or off.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

SYSTem:CONFigure:DUT:STIMe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

DUTSettlingTime

Variable Type       ViReal64

DUT settling time value.

Valid Range:

0 to 20 s

Default Value: 0.05 s

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (DUT Settling Time) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseDUTType

ViStatus rsfsu_confNoiseDUTType (ViSession instrumentHandle,

ViInt32 DUTType);

Purpose

This function allows the type of DUT to be defined. This setting will

determine the method that will be used to create the frequency

measurement list. This function in addition to SENS:CONF:MODE:SYST:LOSC

will determine how the DUT is tested. Fixed Local Oscillator or fixed

Intermediate Frequencies will be taken from the settings supplied by

SENS:CONF:MODE:SYST:LOSC:FREQ and SENS:CONF:MODE:SYST:IF:FREQ

respectively, where appropriate.

The following calculation methods will be used:

Using a variable Local Oscillator frequency as set by the command

SENS:CONF:MODE:SYST:LOSC VAR: i.e. the DUT has a fixed IF frequency.

Using a fixed Local Oscillator frequency as set by the command

SENS:CONF:MODE:SYST:LOSC FIX: i.e. the DUT has a variable IF frequency.

This function will not have any immediate effect if the current frequency

measurement mode is fixed, as is set by the command SENS:FREQ:MODE FIX.

RF: All the above settings for frequency list calculation will

calculate the RF frequency for each step. The list of RF

frequencies generated will be in the following sequence:

Start Frequency, Start Frequency + Step Frequency,

Start Frequency + Step Frequency*2, Stop Frequency

IF: Intermediate frequency

LO: Local Oscillator frequency

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CONFigure:MODE:DUT AMPLifier | DOWNc | UPConv

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

DUTType

Variable Type       ViInt32

Selects type of DUT.

Valid Range:

0 - Amplifier

1 - Down Converter

2 - Up Converter

Default Value:

Notes:

(1) Amplifier:

The DUT is a amplifier and not a frequency converting device.

(2) Down Converter:

The DUT converts the input frequency to a lower output frequency.

(3) Up Converter:

The DUT converts the input frequency to a higher output frequency.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (DUT Type) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseENRSettings

ViStatus rsfsu_confNoiseENRSettings (ViSession instrumentHandle,

ViInt32 selection,

ViReal64 ENRConstant,

ViReal64 roomTemperature);

Purpose

This function is used to specify whether a single ENR spot value applies

for all measured frequencies or whether an ENR list table is to be used.

The room temperature of the operating environment is also to be

specified.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CORRection:ENR:MODE TABLe | SPOT

[SENSe:]CORRection:ENR:SPOT

[SENSe:]CORRection:TEMPerature

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

selection

Variable Type       ViInt32

The ENR mode selection is used to specify whether a single ENR spot

value applies for all measured frequencies or whether an ENR list

table is to be used (ENR values specified at specific input

frequencies).

Valid Range:

0 - Constant

1 - Table

Default Value: 0

ENRConstant

Variable Type       ViReal64

The ENR constant allows the ENR value applicable to the Noise source

for all input frequencies to be specified. This value provides a

simple mechanism to enter an ENR value that is applicable throughout

a range of measurement frequencies.

Valid Range:

-999.99 to 999.99

Default Value: 15.0 dB

roomTemperature

Variable Type       ViReal64

This value allows the room temperature of the operating environment

to be specified. This value will be taken into account when

calculating noise results.

Valid Range:

10 to 1000 K

Default Value: 293.0 K

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Selection) out of range.

BFFC0003  Parameter 3 (ENR Constant) out of range.

BFFC0004  Parameter 4 (Room Temperature) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseENRTable

ViStatus rsfsu_confNoiseENRTable (ViSession instrumentHandle,

ViInt32 no_ofValues,

ViReal64 _VI_FAR frequency[],

ViReal64 _VI_FAR ENRFigure[]);

Purpose

This function allows an ENR lookup table list to be specified that will

be used by the K30 option to determine the correct ENR (excess noise

ratio) figure to use for the input frequency that requires measuring. The

list arguments specified will completely overwrite all current ENR

frequency list entries regardless of how many entries are present and how

many entries are being supplied for the new list. This function will

however not be effective if the remote control command SENS:CORR:ENR:MODE

SPOT was previously entered.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CORRection:ENR[:MEASurement]:TABLe:DATA

<frequency_value>,<ENR_figure_value>,...

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

no_ofValues

Variable Type       ViInt32

Defines number of values in ENR table up to a maximum of argument 100

pairs.

Valid Range:

1 to 100

Default Value: 1

frequency

Variable Type       ViReal64[]

The frequency ENR pair of arguments. The frequency is specified in

Hz.

Valid Range: none

Default Value: none

ENRFigure

Variable Type       ViReal64[]

The ENR figure pair of arguments. The ENR figure must be a value in

dB.

Valid Range: none

Default Value: none

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (No of Values) out of range.

BFFC0003  Parameter 3 (Frequency) out of range.

BFFC0004  Parameter 4 (ENR Figure) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseFixedIFFrequency

ViStatus rsfsu_confNoiseFixedIFFrequency (ViSession instrumentHandle,

ViReal64 fixedIF);

Purpose

This function can be used to specify a fixed intermediate frequency for a

new frequency measurement list that requires computing. This value will

form the basis of one of the input criteria for computing a new frequency

list.

This function will not have any immediate effect if a direct frequency

list is being used (which is set by the SENS:CONF:MODE:DUT AMPL command)

or a variable intermediate frequency has been specified (which is set by

the SENS:CONF:MODE:SYST:LOSC FIX command).

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CONFigure:MODESYSTem:IF:FREQuency

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

fixedIF

Variable Type       ViReal64

Fixed intermediate frequency value.

Valid Range:

FSQ 3  - 0.0 Hz to 3.6 GHz

FSQ 8  - 0.0 Hz to 8.0 GHz

FSQ 26 - 0.0 Hz to 26.5 GHz

Default Value: 0.0 Hz

Note:

(1) This control is not range checked.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Fixed IF) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseFrequency

ViStatus rsfsu_confNoiseFrequency (ViSession instrumentHandle,

ViReal64 frequency);

Purpose

This function is used to specify a single fixed frequency to measure

noise and gain continuously. The value specified with this function is

only used when a single frequency measurement is initiated.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]FREQuency[:CW|:FIXed]

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

frequency

Variable Type       ViReal64

Single fixed frequency to measure noise and gain continuously.

Valid Range:

FSQ 3  - 0.0 Hz to 3.6 GHz

FSQ 8  - 0.0 Hz to 8.0 GHz

FSQ 26 - 0.0 Hz to 26.5 GHz

Default Value: 550.0e6 Hz

Note:

(1) This control is not range checked.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Frequency) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseFrequencyMeasurement

ViStatus rsfsu_confNoiseFrequencyMeasurement (ViSession instrumentHandle,

ViInt32 measurementMode);

Purpose

This function selects a single (continuous) frequency list measurement or

single frequency measurement on the defined frequency. After this

function has been executed a selected measurement will be the measurement

started when the user issues the INITiate command.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

CONFigure:SINGle

CONFigure:LIST:SINGle

CONFigure:LIST:CONTinuous

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measurementMode

Variable Type       ViInt32

Selects type of frequency measurement.

Valid Range:

0 - Single Frequency

1 - Single Frequency List

2 - Continuous Frequency List

Default Value:

Notes:

(1) Single Frequency:

Configures R&S FS-K30 for a single frequency measurement.

(2) Single Frequency List:

Configures R&S FS-K30 for a single frequency list measurement.

(3) Continuous Frequency List:

Configures R&S FS-K30 for a continuous frequency list measurement.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Measurement Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseFrequencyTable

ViStatus rsfsu_confNoiseFrequencyTable (ViSession instrumentHandle,

ViInt32 no_ofValues,

ViReal64 _VI_FAR frequencyList[]);

Purpose

This function allows a new frequency list to be specified which will be

used to perform noise and gain measurements. Each list entry will pertain

to three separate frequency entities, a RF frequency, a local oscillator

frequency and an intermediate frequency. The list arguments specified

will completely overwrite all the current frequency list entries

regardless of how many entries are present and how many entries are being

supplied for the new list. The new list supplied will remain as the

active list until such time that a new list is automatically created. If

the frequency list mode is set to DIRECT then the LO and IF frequencies

specified in this function are ignored.

The values specified in this function are not used if a single frequency

measurement is initiated.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]FREQuency:LIST:DATA <numeric value>,<numeric value>,<numeric

value>...

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

no_ofValues

Variable Type       ViInt32

Defines number of values in frequency list. Each numeric value is

equal to a user defined set of three frequency measurements arguments

up to a maximum of 100 entry triplets.

Valid Range: 3 - 300

Default Value: 3

frequencyList

Variable Type       ViReal64[]

Defines frequency list. Each numeric value is equal to a user defined

set of three frequency measurements arguments up to a maximum of 100

entry triplets in the following order:

Fixed frequency

Local Oscillator frequency

Intermediate frequency

The allowable frequency range is option dependent, specified in Hz

and will be accurate to 2 decimal places.

Default Value: none

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (No of Values) out of range.

BFFC0003  Parameter 3 (Frequency List) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseGainTraceSettings

ViStatus rsfsu_confNoiseGainTraceSettings (ViSession instrumentHandle,

ViBoolean yaxis,

ViBoolean automaticScaling,

ViReal64 minYAxis,

ViReal64 maxYAxis,

ViBoolean symbols);

Purpose

This function settings are the specific settings associated with the

graphical display of Gain results.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

DISPlay[:WINDow<1>]:TRACe<1|2>[:STATe] ON | OFF

DISPlay[:WINDow<1>]:TABLe ON | OFF

DISPlay[:WINDow<1>]:TRACe<1|2>:Y[:SCALe]:AUTO ON | OFF

DISPlay[:WINDow<1>]:TRACe<1|2>:SYMBols ON | OFF

DISPlay[:WINDow<1>]:TRACe<1|2>:Y[:SCALe]:BOTTom

DISPlay[:WINDow<1>]:TRACe<1|2>:Y[:SCALe]:TOP

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

yaxis

Variable Type       ViBoolean

The Y-Axis setting allows the graphical display of gain results to be

turned on and off. When the Y-Axis setting is set to On then gain

results shall be displayed graphically.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

automaticScaling

Variable Type       ViBoolean

Switches on or off automatic scaling of the Y-axis for the specified

trace display. Automatic scaling set the Y-axis to automatically

scale to best fit the measurement results. Note that the specified

trace must be active for this function to have an immediate effect.

This can be achieved using the command DISP:WIND:TRAC<1|2>:STAT ON.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

minYAxis

Variable Type       ViReal64

Sets the minimum (bottom) Y-axis display value for the specified

trace display. Note that this parameter has no affect if automatic

scaling of the Y-axis is enabled or the specified trace is not

currently active.

Valid Range: none

Default Value: 0.0 dB

maxYAxis

Variable Type       ViReal64

Sets the maximum (top) Y-axis display value for the specified trace

display. Note that this parameter has no affect if automatic scaling

of the Y-axis is enabled or the specified trace is not currently

active.

Valid Range: none

Default Value: 0.0 dB

symbols

Variable Type       ViBoolean

Switches on or off the display of the symbols to mark the measurement

points for the specified trace (Gain).

This parameter will only have a noticeable effect if the required

trace is active as set by the command DISP:WIND:TRAC<1|2>:STAT ON.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Y-axis) out of range.

BFFC0003  Parameter 3 (Automatic Scaling) out of range.

BFFC0004  Parameter 4 (Min Y-Axis) out of range.

BFFC0005  Parameter 5 (Max Y-Axis) out of range.

BFFC0006  Parameter 6 (Symbols) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseGeneratorAuto

ViStatus rsfsu_confNoiseGeneratorAuto (ViSession instrumentHandle,

ViBoolean automaticControl);

Purpose

This function is used to specify whether the setup of the external

generator is to be automatically controlled by the R&S FS-K30 option (via

GPIB) or manually by the user.

Note:

(1) This function is available only in conjunction with option FS-K30 and

FS-B10.

Remote-control command(s):

SYSTem:CONFigure:GENerator:CONTrol:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

automaticControl

Variable Type       ViBoolean

Specify whether the setup of the external generator is to be

automatically controlled or not.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Automatic Control) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseGeneratorFrequency

ViStatus rsfsu_confNoiseGeneratorFrequency (ViSession instrumentHandle,

ViReal64 offset1,

ViReal64 numerator,

ViReal64 denominator,

ViReal64 offset2);

Purpose

This function is used to set frequency of external generator.

When an external signal generator (or VCO) is used to create a Local

Oscillator (LO) frequency, the following formula is used to calculate the

generator frequency:

Generator Frequency = [(LO + Offset1) * Numerator / Denominator)] +

Offset2

Select the parameters of the generator so that the frequency range of the

generator is not exceeded with the above formula.

Note:

(1) This function is available only in conjunction with option FS-K30 and

FS-B10.

Remote-control command(s):

SOURce:EXTernal:FREQuency:OFFSet<1|2>

SOURce:EXTernal:FREQuency[:FACTor]:NUMerator

SOURce:EXTernal:FREQuency[:FACTor]:DENominator

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

offset1

Variable Type       ViReal64

Defines the frequency offset of the selected generator with reference

to the receive frequency.

Valid Range: none

Default Value: 0.0 Hz

numerator

Variable Type       ViReal64

Defines the numerator of the factor with which the analyzer frequency

is multiplied in order to obtain the transmit frequency of the

selected generator.

Valid Range: none

Default Value: 1.0

denominator

Variable Type       ViReal64

Defines the denominator of the factor with which the analyzer

frequency is multiplied in order to obtain the transmit frequency of

the selected generator.

Valid Range: none

Default Value: 1.0

offset2

Variable Type       ViReal64

Defines the frequency offset of the selected generator with reference

to the receive frequency.

Valid Range: none

Default Value: 0.0 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Offset 1) out of range.

BFFC0003  Parameter 3 (Numerator) out of range.

BFFC0004  Parameter 4 (Denominator) out of range.

BFFC0005  Parameter 5 (Offset 2) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseGeneratorLevel

ViStatus rsfsu_confNoiseGeneratorLevel (ViSession instrumentHandle,

ViReal64 level);

Purpose

This function sets the output power of the selected generator.

Note:

(1) This function is available only in conjunction with option FS-K30 and

FS-B10.

Remote-control command(s):

SOURce:EXTernal:POWer[:LEVel]

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

level

Variable Type       ViReal64

Sets the output power of the selected generator.

Valid Range: none

Default Value: 5.0 dBm

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Level) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseGeneratorSettings

ViStatus rsfsu_confNoiseGeneratorSettings (ViSession instrumentHandle,

ViString generatorSelect,

ViInt32 GPIBAddress,

ViBoolean initBeforeMeas);

Purpose

This function controls the set-up of external signal generator.

Note:

(1) This function is available only in conjunction with option FS-K30 and

FS-B10.

Remote-control command(s):

SYSTem:COMMunicate:RDEVice:GENerator:TYPE <name>

SYSTem:COMMunicate:GPIB:RDEVice:GENerator:ADDRess

SYSTem:CONFigure:GENerator:INITialise:AUTO ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

generatorSelect

Variable Type       ViString

Selects the type of external signal generator.

Valid Range: none

Default Value: ""

Notes:

The following list shows the available generator types:

SME02, SME03, SME06, SMG, SMGL, SMGU, SMH, SMHU, SMIQ02B, SMIQ02E,

SMIQ03B, SMIQ03E, SMIQ04B, SMIQ06B, SML01, SMR20, SMR20B11, SMR27,

SMR27B11, SMR30, SMR30B11, SMR40, SMR40B11, SMP02, SMP03, SMP04,

SMP22, SMT02, SMT03, SMT06, SMX, SMY01, SMY02, HP8340A, HP ESG-A

Series 1000A, HP ESG-A Series 2000A, HP ESG-A Series 3000A, HP ESG-A

Series 4000A, HP ESG-D SERIES E4432B

For more details refer also to the documentation.

GPIBAddress

Variable Type       ViInt32

Changes the IEC/IEEE-bus address of the external signal generator.

Valid Range: 0 to 30

Default Value: 28

initBeforeMeas

Variable Type       ViBoolean

Specify whether an initialisation sequence of GPIB commands is sent

to an external signal generator prior to performing each measurement.

Sending an initialisation sequence before each measurement ensures

that the signal generator will be in the correct state to receive

control commands during a measurement sequence, however initialising

an external signal generator before each measurement adds a time

overhead to each measurement.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Generator Select) out of range.

BFFC0003  Parameter 3 (GPIB Address) out of range.

BFFC0004  Parameter 4 (Init Before Meas) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseImageRejection

ViStatus rsfsu_confNoiseImageRejection (ViSession instrumentHandle,

ViReal64 imageRejection);

Purpose

This function allows an image rejection value associated with DUT to be

entered which will be effective across the complete range of measurement

frequencies.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CORRection:IREJection

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

imageRejection

Variable Type       ViReal64

Image rejection value.

Valid Range:

0 to 999.99 dB

Defaul Value: 999.99 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Image Rejection) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseLOFrequency

ViStatus rsfsu_confNoiseLOFrequency (ViSession instrumentHandle,

ViInt32 localOscillator,

ViReal64 LOFrequency);

Purpose

This function can be used to specify whether the local oscillator will be

used as a fixed frequency source (fixed local oscillator frequency for a

new frequency measurement list that requires computing) or a variable

frequency source.

This function will not have any immediate effect if a direct frequency

list is being used which is set by the remote control command

SENS:CONF:MODE:DUT AMPL.

See the remote command SENS:CONF:MODE:DUT for an explanation on how the

measurement lists are created.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CONFigure:MODESYSTem:LOSCillator FIXed | VARiable

[SENSe:]CONFigure:MODESYSTem:LOSCillator:FREQuency

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

localOscillator

Variable Type       ViInt32

Specify whether the local oscillator will be used as a fixed

frequency source or a variable frequency source.

Valid Range:

0 - Fixed

1 - Variable

Default Value: 0

Notes:

(1) Fixed:

Local Oscillator is used as a fixed frequency source with value

entered to LO Frequency parameter.

(2) Variable:

Local Oscillator is used as a variable frequency source.

LOFrequency

Variable Type       ViReal64

Fixed local oscillator frequency value.

Valid Range:

FSQ 3  - 0.0 Hz to 3.6 GHz

FSQ 8  - 0.0 Hz to 8.0 GHz

FSQ 26 - 0.0 Hz to 26.5 GHz

Default Value: 0.0 Hz

Note(s):

(1) This control is not range checked.

(2) This setting will not have any immediate effect if a variable

local osciallator frequency has been specified.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Local Oscillator) out of range.

BFFC0003  Parameter 3 (LO Frequency) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseLossInputSettings

ViStatus rsfsu_confNoiseLossInputSettings (ViSession instrumentHandle,

ViInt32 selection,

ViReal64 lossInputConstant);

Purpose

This function is used to identify whether a single input loss value

applies for all measured frequencies or whether an input loss list is to

be used. (Loss Input values specified at specific input frequencies).

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CORRection:LOSS:INPut:MODE SPOT | TABLe

[SENSe:]CORRection:LOSS:INPut:SPOT

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

selection

Variable Type       ViInt32

The mode selection is used to specify whether a single input loss

value applies for all measured frequencies or whether an input loss

list is to be used.

Valid Range:

0 - Constant

1 - Table

Default Value: 0

lossInputConstant

Variable Type       ViReal64

This Loss Input Constant allows the loss input constant to be set to

a value applicable for all input frequencies to be measured. This

value provides a simple mechanism to enter a loss input value that is

applicable throughout a range of measurement frequencies. This value

will however not be effective if the remote control command

SENS:CORR:LOSS:INP:MODE TABL was previously entered.

Valid Range:

-999.99 to 999.99 dB

Default Value: 0.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Selection) out of range.

BFFC0003  Parameter 3 (Loss Input Constant) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseLossInputTable

ViStatus rsfsu_confNoiseLossInputTable (ViSession instrumentHandle,

ViInt32 no_ofValues,

ViReal64 _VI_FAR frequency[],

ViReal64 _VI_FAR lossInput[]);

Purpose

This function allows a new user defined input loss lookup table list to

be specified. The list arguments specified will completely overwrite all

current input loss list entries regardless of how many entries are

present and how many entries are being supplied for the new list. This

list will be interpreted by the K30 option to determine the correct input

loss to use for the frequency that requires measuring.

This function will however not be effective if the remote control command

SENS:CORR:LOSS:INP:MODE SPOT was previously entered.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CORRection:LOSS:INPut:TABLe

<frequency_value>,<loss_figure_value>,...

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

no_ofValues

Variable Type       ViInt32

Defines number of values in the table up to a maximum of argument 100

pairs.

Valid Range:

1 to 100

Default Value: 0

frequency

Variable Type       ViReal64[]

The frequency pair of arguments. The frequency is specified in Hz

with a maximum of two decimal places.

Valid Range: none

Default Value: none

lossInput

Variable Type       ViReal64[]

The loss figure pair of arguments. The loss figure must be a value in

dB between -999.99 and 999.99 accurate to two decimal places.

Valid Range: none

Default Value: none

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (No of Values) out of range.

BFFC0003  Parameter 3 (Frequency) out of range.

BFFC0004  Parameter 4 (Loss Input) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseLossOutputSettings

ViStatus rsfsu_confNoiseLossOutputSettings (ViSession instrumentHandle,

ViInt32 selection,

ViReal64 lossOutputConstant);

Purpose

This function is used to identify whether a single output loss value

applies for all measured frequencies or whether an output loss list is to

be used. (Loss Output values specified at specific input frequencies).

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CORRection:LOSS:OUTPut:MODE SPOT | TABLe

[SENSe:]CORRection:LOSS:OUTPut:SPOT

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

selection

Variable Type       ViInt32

The mode selection is used to specify identify whether a single

output loss value applies for all measured frequencies or whether an

output loss list is to be used.

Valid Range:

0 - Constant

1 - Table

Default Value: 0

lossOutputConstant

Variable Type       ViReal64

This loss output constant to be set to a value applicable for all

input frequencies to be measured. This function provides a simple

mechanism to enter a loss output value that is applicable throughout

a range of measurement frequencies. This value will however not be

effective if the remote control command SENS:CORR:LOSS:OUTP:MODE TABL

was previously entered.

Valid Range:

-999.99 to 999.99 dB

Default Value: 0.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Selection) out of range.

BFFC0003  Parameter 3 (Loss Output Constant) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseLossOutputTable

ViStatus rsfsu_confNoiseLossOutputTable (ViSession instrumentHandle,

ViInt32 no_ofValues,

ViReal64 _VI_FAR frequency[],

ViReal64 _VI_FAR lossOutput[]);

Purpose

This function allows a new user defined output loss lookup table list to

be specified. The list arguments specified willcompletely overwrite all

current output loss list entries regardless of how many entries are

present and how many entries are being supplied for the new list. This

list will be interpreted by the K30 option to determine the correct

output loss to use for the frequency that requires measuring. This

function will however not be effective if the remote control command

SENS:CORR:LOSS:INP:MODE SPOT was previously entered.

This function will however not be effective if the remote control command

SENS:CORR:LOSS:OUTP:MODE SPOT was previously entered.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]CORRection:LOSS:OUTPut:TABLe

<frequency_value>,<loss_figure_value>,...

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

no_ofValues

Variable Type       ViInt32

Defines number of values in the table up to a maximum of argument 100

pairs.

Valid Range:

1 to 100

Default Value: 0

frequency

Variable Type       ViReal64[]

The frequency pair of arguments. The frequency is specified in Hz

with a maximum of two decimal places.

Valid Range: none

Default Value: none

lossOutput

Variable Type       ViReal64[]

The loss figure pair of arguments. The loss figure must be a value in

dB between -999.99 and 999.99 accurate to two decimal places.

Valid Range: none

Default Value: none

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (No of Values) out of range.

BFFC0003  Parameter 3 (Frequency) out of range.

BFFC0004  Parameter 4 (Loss Output) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoisePreamplifier

ViStatus rsfsu_confNoisePreamplifier (ViSession instrumentHandle,

ViBoolean preamplifier);

Purpose

This function switches the pre-amplifier on and off. If only the B2 or

B25 option is present then the pre-amplifier of the installed option is

controlled. If both the B2 and B25 options are present then the

pre-amplifier controlled by this function depends on the setting for the

INPut:PRESelection[:STATe] command. If the INPut:PRESelection[:STATe] is

set to ON then the pre-amplifier for the B2 option is controlled by this

function, otherwise the pre-amplifier of the B25 option is controlled.

Note:

(1) This function is available only in conjunction with option FS-K30 and

FS-B2 (or FS-B25).

Remote-control command(s):

INPut:GAIN:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

preamplifier

Variable Type       ViBoolean

Switches the pre-amplifier on and off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Pre-amplifier) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoisePreselector

ViStatus rsfsu_confNoisePreselector (ViSession instrumentHandle,

ViBoolean preselector);

Purpose

This function switches the preselection on or off.

Note:

(1) This function is available only in conjunction with option FS-K30 and

FS-B2.

Remote-control command(s):

INPut:PRESelection[:STATe] ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

preselector

Variable Type       ViBoolean

Switches the preselection on or off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Pre-selector) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseRBW

ViStatus rsfsu_confNoiseRBW (ViSession instrumentHandle,

ViReal64 resolutionBandwidth);

Purpose

This function allows the resolution bandwidth of the analyser to be

specified.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]BANDwidth|BWIDth:RESolution

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

resolutionBandwidth

Variable Type       ViReal64

Resolution bandwidth of the analyser.

Valid Range:

10.0 to 1.0e7 Hz

Default Value: 1.0e6 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Resolution Bandwidth) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseRefLevel

ViStatus rsfsu_confNoiseRefLevel (ViSession instrumentHandle,

ViBoolean refLevelAuto,

ViReal64 refLevelValue);

Purpose

This function enables or disables automatic reference level detection. If

reference level detection is enabled, the reference level for the Y-axis

display value for all trace displays is set.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

DISPlay[:WINDow<1>]:TRACe<1|2>:Y[:SCALe]:RLEVel:AUTO ON | OFF

DISPlay[:WINDow<1>]:TRACe<1|2>:Y[:SCALe]:RLEVel

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

refLevelAuto

Variable Type       ViBoolean

Enables or disables automatic reference level detection.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

refLevelValue

Variable Type       ViReal64

If reference level detection is enabled, the reference level for the

Y-axis display value for all trace displays is set.

Valid Range: none

Default Value: -30.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Ref Level Auto) out of range.

BFFC0003  Parameter 3 (Ref Level Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseRFAttenuation

ViStatus rsfsu_confNoiseRFAttenuation (ViSession instrumentHandle,

ViReal64 RFAttenuation);

Purpose

This function specifies the RF attenuation that the analyser imposes.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

INPut:ATTenuation

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

RFAttenuation

Variable Type       ViReal64

Specifies the RF attenuation value.

Valid Range:

0 to 75

Default Value: 0.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (RF Attenuation) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseStartFrequency

ViStatus rsfsu_confNoiseStartFrequency (ViSession instrumentHandle,

ViReal64 startFrequency);

Purpose

This function can be used to specify the starting frequency for a new

frequency measurement list that requires computing. This value will form

the basis of one of the input criteria for computing a new frequency

list.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]FREQuency:STARt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

startFrequency

Variable Type       ViReal64

Starting frequency for a new frequency measurement list that requires

computing.

Valid Range:

FSQ 3  - 0.0 Hz to 3.6 GHz

FSQ 8  - 0.0 Hz to 8.0 GHz

FSQ 26 - 0.0 Hz to 26.5 GHz

Default Value: 550.0e6 Hz

Note:

(1) This control is not range checked.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Start Frequency) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseStepFrequency

ViStatus rsfsu_confNoiseStepFrequency (ViSession instrumentHandle,

ViReal64 stepFrequency);

Purpose

This function can be used to specify the step frequency for a new

frequency measurement list that requires computing. This value will form

the basis of one of the input criteria for computing a new frequency

list.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]FREQuency:STEP

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

stepFrequency

Variable Type       ViReal64

Step frequency for a new frequency measurement list that requires

computing.

Valid Range:

FSQ 3  - 0.0 Hz to 3.6 GHz

FSQ 8  - 0.0 Hz to 8.0 GHz

FSQ 26 - 0.0 Hz to 26.5 GHz

Default Value: 2.0e6 Hz

Note:

(1) This control is not range checked.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Step Frequency) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseStopFrequency

ViStatus rsfsu_confNoiseStopFrequency (ViSession instrumentHandle,

ViReal64 stopFrequency);

Purpose

This function can be used to specify the ending frequency for a new

frequency measurement list that requires computing. This value will form

the basis of one of the input criteria for computing a new frequency

list.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]FREQuency:STOP

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

stopFrequency

Variable Type       ViReal64

Stop frequency for a new frequency measurement list that requires

computing.

Valid Range:

FSQ 3  - 0.0 Hz to 3.6 GHz

FSQ 8  - 0.0 Hz to 8.0 GHz

FSQ 26 - 0.0 Hz to 26.5 GHz

Default Value: 550.0e6 Hz

Note:

(1) This control is not range checked.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Stop Frequency) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseSweepTime

ViStatus rsfsu_confNoiseSweepTime (ViSession instrumentHandle,

ViReal64 sweepTime);

Purpose

This function allows the sweep time to be specified.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

[SENSe:]SWEep:TIME

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

sweepTime

Variable Type       ViReal64

Sweep time value.

Valid Range:

1e-006 to 16000.0

Default Value: 0.1 s

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Sweep Time) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseTraceDisplay

ViStatus rsfsu_confNoiseTraceDisplay (ViSession instrumentHandle,

ViInt32 traceDisplay);

Purpose

This function toggles the display of traces between being displayed in

separate graphs or displayed in a single combined graph.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

DISPlay:FORMat SPLit | SINGle

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

traceDisplay

Variable Type       ViInt32

Toggles the display of traces between being displayed in separate

graphs or displayed in a single combined graph.

Valid Range:

0 - Single

1 - Split

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Trace Display) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNoiseTraceSettings

ViStatus rsfsu_confNoiseTraceSettings (ViSession instrumentHandle,

ViInt32 yaxis,

ViBoolean automaticScaling,

ViReal64 minYAxis,

ViReal64 maxYAxis,

ViBoolean symbols);

Purpose

This function sets parameters associated with the graphical display of

Noise (Noise Figure and Noise Temperature) results.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

DISPlay[:WINDow<1>]:TRACe<1|2>[:STATe] ON | OFF

DISPlay[:WINDow<1>]:TABLe ON | OFF

DISPlay:DATA:TRACe<1> NFIGure | TEFFective

DISPlay[:WINDow<1>]:TRACe<1|2>:Y[:SCALe]:AUTO ON | OFF

DISPlay[:WINDow<1>]:TRACe<1|2>:SYMBols ON | OFF

DISPlay[:WINDow<1>]:TRACe<1|2>:Y[:SCALe]:BOTTom

DISPlay[:WINDow<1>]:TRACe<1|2>:Y[:SCALe]:TOP

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

yaxis

Variable Type       ViInt32

Allows the user to specify the type of data to be displayed in trace

1. The user can select either to display either Noise Figure results

(NFIGure), Noise temperature (TEFFective - to keep compatibility with

the NFA) or switch it to Off - Noise results are not displayed

graphically then.

Valid Range:

0 - Off

1 - Noise Figure

2 - Noise Temperature

Default Value: 1

automaticScaling

Variable Type       ViBoolean

Switches on or off automatic scaling of the Y-axis for the specified

trace display. Automatic scaling set the Y-axis to automatically

scale to best fit the measurement results. Note that the specified

trace must be active for this function to have an immediate effect.

This can be achieved using the command DISP:WIND:TRAC<1|2>:STAT ON.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

minYAxis

Variable Type       ViReal64

Sets the minimum (bottom) Y-axis display value for the specified

trace display. Note that this parameter has no affect if automatic

scaling of the Y-axis is enabled or the specified trace is not

currently active.

Valid Range:

-75 to 75 dB

Default Value: 0.0 dB

maxYAxis

Variable Type       ViReal64

Sets the maximum (top) Y-axis display value for the specified trace

display. Note that this parameter has no affect if automatic scaling

of the Y-axis is enabled or the specified trace is not currently

active.

Valid Range:

-75 to 75 dB

Default Value: 0.0 dB

symbols

Variable Type       ViBoolean

Switches on or off the display of the symbols to mark the measurement

points for the specified trace (Noise).

This parameter will only have a noticeable effect if the required

trace is active as set by the command DISP:WIND:TRAC<1|2>:STAT ON.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Y-axis) out of range.

BFFC0003  Parameter 3 (Automatic Scaling) out of range.

BFFC0004  Parameter 4 (Min Y-Axis) out of range.

BFFC0005  Parameter 5 (Max Y-Axis) out of range.

BFFC0006  Parameter 6 (Symbols) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNTHCDevice

ViStatus rsfsu_confNTHCDevice (ViSession instrumentHandle,

ViInt32 deviceNumber, ViInt32 dataFormat,

ViInt32 destination);

Purpose

This function selects the interface and configuration for the hardcopy.

The destination and data format are also configured.

Note:

(1) This function is reserved for compatibility with FSEx driver. Prefer

to use function rsfsu_confNTHCDeviceExt.

Remote-control command(s):

HCOPy:DEVice:LANGuage<1|2> WMF | GDI | EWMF | BMP

HCOPy:DESTination<1|2> 'MMEM'

HCOPy:DESTination<1|2> 'SYST:COMM:PRIN'

HCOPy:DESTination<1|2> 'SYST:COMM:CLIP'

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

deviceNumber

Variable Type       ViInt32

Selects the device to be configured.

Valid Range:

1 - 1 (Default Value)

2 - 2

dataFormat

Variable Type       ViInt32

This control determines the data format of the printout.

Valid Range:

0 - WMF (Default Value)

1 - GDI

2 - EWMF

3 - BMP

Notes:

WMF           (WINDOWS Metafile Format) and

EWMF          Data formats for output files which can at a later

time be integrated in corresponding programs for

documentation purposes.

BMP           (Bitmap) Data format for output files.

GDI           (Graphics Device Interface) Default format for the

output to a printer configured under Windows. For

the output file the printer driver configured

under Windows is used and thus a printer-specific

format is generated.

destination

Variable Type       ViInt32

This control selects the interface for the hardcopy device.

Valid Range:

0 - 'MMEM' (Creates a file for hardcopy output)

1 - 'SYST:COMM:PRIN'

2 - 'SYST:COMM:CLIP'

Default Value: 0

Notes:

(1) 'MMEM':

Creates a file for the hardcopy output.

(2) 'SYST:COMM:PRIN':

Directs the hardcopy to the printer. directs the hardcopy to the

printer. The printer is selected with command

SYSTEM:COMMunicate:PRINter:SELect. GDI should be selected for

HCOPy:DEVice:LANGuage.

(3) 'SYST:COMM:CLIP':

directs the hardcopy to the clipboard. EWMF should be selected for

data format.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Device Number) out of range.

BFFC0003  Parameter 3 (Data Format) out of range.

BFFC0004  Parameter 4 (Destination) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confNTHCDeviceExt

ViStatus rsfsu_confNTHCDeviceExt (ViSession instrumentHandle,

ViInt32 deviceNumber,

ViInt32 dataFormat,

ViInt32 destination,

ViBoolean pageOrientation);

Purpose

This function selects the interface and configuration for the hardcopy.

The destination and data format are also configured.

Remote-control command(s):

HCOPy:DEVice:LANGuage<1|2> WMF | GDI | EWMF | BMP

HCOPy:PAGE:ORIentation<1|2> PORTrait | LANDscape

HCOPy:DESTination<1|2> 'MMEM' | 'SYST:COMM:PRIN' | 'SYST:COMM:CLIP'

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

deviceNumber

Variable Type       ViInt32

Selects the device to be configured.

Valid Range:

1 - 1 (Default Value)

2 - 2

dataFormat

Variable Type       ViInt32

This control determines the data format of the printout.

Valid Range:

0 - WMF (Default Value)

1 - GDI

2 - EWMF

3 - BMP

Notes:

WMF           (WINDOWS Metafile Format) and

EWMF          Data formats for output files which can at a later

time be integrated in corresponding programs for

documentation purposes.

BMP           (Bitmap) Data format for output files.

GDI           (Graphics Device Interface) Default format for the

output to a printer configured under Windows. For

the output file the printer driver configured

under Windows is used and thus a printer-specific

format is generated.

destination

Variable Type       ViInt32

This control selects the interface for the hardcopy device.

Valid Range:

0 - 'MMEM'

1 - 'SYST:COMM:PRIN'

2 - 'SYST:COMM:CLIP'

Default Value: 0

Notes:

(1) 'MMEM':

Directs the hardcopy to a file. To define file name use function File

Operations (rsfsu_actMemOp()) with required settings. Page

Orientation settings will be ignored.

(2) 'SYST:COMM:PRIN':

Directs the hardcopy to the printer. directs the hardcopy to the

printer. The printer is selected with command

SYSTEM:COMMunicate:PRINter:SELect. GDI should be selected for

HCOPy:DEVice:LANGuage.

(3) 'SYST:COMM:CLIP':

directs the hardcopy to the clipboard. EWMF should be selected for

data format. The Page Orientation settings will be ignored.

pageOrientation

Variable Type       ViBoolean

The command selects the format of the output (portrait and landscape)

(hardcopy unit 1 or 2).

Valid Range:

VI_TRUE  (1) - Portrait (Default Value)

VI_FALSE (0) - Landscape

Note:

(1)  This setting will be ignored for

Destination <> 'SYST:COMM:PRIN'.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Device Number) out of range.

BFFC0003  Parameter 3 (Data Format) out of range.

BFFC0004  Parameter 4 (Destination) out of range.

BFFC0005  Parameter 5 (Page Orientation) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confPMetFrequency

ViStatus rsfsu_confPMetFrequency (ViSession instrumentHandle,

ViInt32 window, ViReal64 frequency);

Purpose

This function sets the frequency of the power sensor. The limit values

fmin and fmax are set by the power sensor that is connected.

Note:

(1) This function is available only in conjunction with option FS-K9.

Remote-control command(s):

[SENSe<1|2>:]PMETer:FREQuency

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

frequency

Variable Type       ViReal64

Sets the frequency of the power sensor. The limit values fmin and

fmax are set by the power sensor that is connected.

Valid Range: fmin to fmax

Default Value: 50.0e6 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Frequency) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confPMetFrequencyCoupling

ViStatus rsfsu_confPMetFrequencyCoupling (ViSession instrumentHandle,

ViInt32 window,

ViInt32 frequencyCoupling);

Purpose

This function sets the coupling for the frequency of the power sensor.

Note:

(1) This function is available only in conjunction with option FS-K9.

Remote-control command(s):

[SENSe<1|2>:]PMETer:FREQuency:LINK OFF | CENTer | MARKer1

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

frequencyCoupling

Variable Type       ViInt32

Sets the coupling for the frequency of the power sensor.

Valid Range:

0 - Off

1 - Center

2 - Marker 1

Notes:

(1) Off:

Reactivates the coupling of the frequency.

(2) Center:

Couples the frequency to the center frequency of the analyzer.

(3) Marker 1:

Couples the frequency to the position of marker 1.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Frequency Coupling) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confPMetMeasTime

ViStatus rsfsu_confPMetMeasTime (ViSession instrumentHandle,

ViInt32 window, ViInt32 measTime);

Purpose

This function determines the measurement time of the power sensor.

Note:

(1) This function is available only in conjunction with option FS-K9.

Remote-control command(s):

[SENSe<1|2>:]PMETer:MTIMe SHORt | NORMal | LONG

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

measTime

Variable Type       ViInt32

Determines the measurement time of the power sensor.

Valid Range:

0 - Short

1 - Normal

2 - Long

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Meas Time) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confPMetResultDisplay

ViStatus rsfsu_confPMetResultDisplay (ViSession instrumentHandle,

ViInt32 window,

ViInt32 resultDisplay,

ViReal64 referenceLevelOffset);

Purpose

This function controls the selection and presentation of text and

graphics information as well as measurement data on the screen.

Note:

(1) This function is available only in conjunction with option FS-K9.

Remote-control command(s):

DISPlay[:WINDow<1|2>]:TRACe:Y[:SCALe]:RLEVel:OFFSet

DISPlay[:WINDow<1|2>]:PMETer:STATe ON | OFF

CALCulate<1|2>:PMETer:RELative:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

resultDisplay

Variable Type       ViInt32

Selects type of result display.

Valid Range:

0 - Bargraph Display Off

1 - Bargraph Display On

2 - Relative Display

3 - Absolute Display

4 - Reference Level Offset

Notes:

(1) Bargraph Display:

Switches bargraph display on or off for measurements with a power

sensor. If bargraph display is switched off, the measurement value

appears in the Marker Info field.

(2) Relative / Absolute Display:

Switches between relative and absolute display of the result.

(3) Reference Level Offset:

Defines the reference level offset in the selected measurement

window. Depending on the coupling of the measurement windows, this

function is valid for both screens (INSTrument:COUPle ALL) or only

for the selected measurement window (INSTrument:COUPle NONE).

referenceLevelOffset

Variable Type       ViReal64

Defines the reference level offset in the selected measurement

window.

Valid Range: -200.0 dB to 200.0 dB

Default Value: 0.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Result Display) out of range.

BFFC0004  Parameter 4 (Reference Level Offset) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confPMetUnits

ViStatus rsfsu_confPMetUnits (ViSession instrumentHandle,

ViInt32 window, ViInt32 meterUnits);

Purpose

This function selects the unit for absolute or relative measurements with

a power sensor.

Note:

(1) This function is available only in conjunction with option FS-K9.

Remote-control command(s):

UNIT<1|2>:PMETer:POWer DBM | WATT | W

UNIT<1|2>:PMETer:POWer:RATio DB | PCT

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

meterUnits

Variable Type       ViInt32

Selects the unit for measurements with a power sensor.

Valid Range:

Absolute:

0 - dBm

1 - Watt

Relative:

2 - dB

3 - %

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Meter Units) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confPVTLimitLine

ViStatus rsfsu_confPVTLimitLine (ViSession instrumentHandle,

ViInt32 measMode,

ViReal64 expectedPower);

Purpose

This function switches on or off the automatic selection of limit lines

for Power vs Time measurements.

Note:

(1) This function is available only in conjunction with option FS-K84.

Remote-control command(s):

CALCulate:LIMit:PVTime:REFerence AUTO | MANual | ONCE

CALCulate:LIMit:PVTime:RVALue

CALCulate:LIMit:PVTime:RESTore

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measMode

Variable Type       ViInt32

This control switches on or off the automatic selection of limit

lines for power vs time measurements.

Valid Range:

0 - Auto (Default Value)

1 - Manual

2 - Once

3 - Restore

Notes:

(1) AUTO:

Set limit line depends on the measured channel power.

(2) MANUAL:

One of fixed limit lines is activated. The limit line is selected

using Expected Power control.

(3) ONCE:

With the parameter ONCE the current mean power value of the averaged

time response is used to set as the fixedreference value for the

limit lines. The mode is switched from AUTO to MANUAL.

(4) RESTORE:

Restores the standard limit lines for the power versus time

measurement. All changes made to the standard limit lines are lost

and the state of these limit lines as they left the factory is

restored.

expectedPower

Variable Type       ViReal64

The reference level in dBm for calculating the limit lines.

Valid Range: not checked

Default Value: 0.0 dBm

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Meas Mode) out of range.

BFFC0003  Parameter 3 (Expected Power) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confReferenceOsc

ViStatus rsfsu_confReferenceOsc (ViSession instrumentHandle,

ViBoolean source, ViInt32 frequency);

Purpose

This function selects the reference oscillator.

Remote-control command(s):

ROSCillator:SOURce INTernal | EXTernal

ROSCillator:EXTernal:FREQuency

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

source

Variable Type       ViBoolean

Selects the source of the reference oscillator.

Valid Range:

VI_FALSE (0) - Internal (Default Value)

VI_TRUE  (1) - External

frequency

Variable Type       ViInt32

This control sets the external reference oscillator frequency.

Valid Range: 2000000 Hz to 20000000 Hz

Default Value: 0

Note:

(1) Zero (0) is here for backward compatibility. It sets the 10 MHz

(instrument default reference frequency).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Source) out of range.

BFFC0003  Parameter 3 (Frequency) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRefLevel

ViStatus rsfsu_confRefLevel (ViSession instrumentHandle, ViInt32 window,

ViReal64 referenceLevelValue);

Purpose

This function configures the reference level for the signal analyzer

mode.

Remote-control command(s):

DISPlay:WINDow<1|2>:TRACe:Y:RLEVel

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

referenceLevelValue

Variable Type       ViReal64

Specifies the reference level, or if selected  defines the upper

limit of the display range.

Valid Range: -1.0e12 to 1.0e12

Default Value: 0.0

Notes:

(1) To set reference level unit use function Reference Level Opt.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Reference Level Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRefLevelOpt

ViStatus rsfsu_confRefLevelOpt (ViSession instrumentHandle,

ViInt32 window,

ViInt32 referenceLevelFunctionMode,

ViReal64 refLevelPosition_,

ViReal64 referenceLevelOffset_dB,

ViInt32 unit);

Purpose

This function configures the reference level, reference offset, and input

units for the instrument when it is in spectrum analyzer mode.

Remote-control command(s):

DISPlay:WINDow<1|2>:TRACe:Y:RLEVel:OFFSet

DISPlay:WINDow<1|2>:TRACe:Y:SCALe:RPOSition

CALCulate<1|2>:UNIT:POWER

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

referenceLevelFunctionMode

Variable Type       ViInt32

Selects the parameter(s) to set to the instrument.

Valid Range:

0 - Reference Level Offset (Default Value)

1 - RESERVED

2 - RESERVED

3 - RESERVED

4 - RESERVED

5 - Reference Level Unit

6 - Reference Level Position

refLevelPosition_

Variable Type       ViReal64

Specifies the reference level position.

Valid Range: -200.0 % to 200.0 %

Default Value: 100.0 %

referenceLevelOffset_dB

Variable Type       ViReal64

Sets the reference level offset.

Valid Range: -200.0 dB to 200.0 dB

Default Value: 0.0 dB

unit

Variable Type       ViInt32

This control selects the reference level unit.

Valid Range:

0 - dBm (Default Value)

1 - dBuV

2 - dBuA

3 - dBpW

4 - V

5 - A

6 - W

7 - dBmV

8 - No Units

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Reference Level Function ) out of range.

BFFC0004  Parameter 4 (Ref Level Position (%)) out of range.

BFFC0005  Parameter 5 (Reference Level Offset (d) out of range.

BFFC0006  Parameter 6 (Unit) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confResbw

ViStatus rsfsu_confResbw (ViSession instrumentHandle, ViInt32 window,

ViReal64 bwValueHz);

Purpose

This function controls the setting of the analyzer's resolution filter

bandwidths. Automatically sets the coupling to manual.

Remote-control command(s):

SENSe<1|2>:BANDwidth:RESolution

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

bwValueHz

Variable Type       ViReal64

Sets manually the resolution bandwidth.

Valid Range:

Analog Filter: 10.0 Hz to 50.0e6 Hz

FFT-Filter:     1.0 Hz to 50.0e6 Hz

Default Value: 1.0e6 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Bw Value [Hz]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRFInAtt

ViStatus rsfsu_confRFInAtt (ViSession instrumentHandle, ViInt32 window,

ViReal64 manualAttenuatorValuedB);

Purpose

This function configures the attenuator at the instrument's RF input. It

switches automatically to manual mode.

Note:

In the default state, the attenuation set on the step attenuator is

coupled to the reference level of the instrument. If the attenuation is

programmed directly, the coupling to the reference level is switched off.

Remote-control command(s):

INPut<1|2>:ATTenuation

INPut<1|2>:ATTenuation:AUTO OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

manualAttenuatorValuedB

Variable Type       ViReal64

This control sets the manual input attenuation level.

Valid Range: 0.0 to 75.0 dB

Default Value: 10.0 dB

Note:

The step width is 5 dB, the range is 0 dB to 75 dB (with or without

option B25 electronic attenuator)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Manual Attenuator Value [) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRFInAttAuto

ViStatus rsfsu_confRFInAttAuto (ViSession instrumentHandle,

ViInt32 window,

ViInt32 RFAttenuation_not_used);

Purpose

This function configures the attenuator at the instrument's RF input to

the auto mode.

Note:

The minimum input attenuation set with the coupling switched on is 10 dB

(with electronic attenuator option: 5 dB).

Remote-control command(s):

INPut<1|2>:ATTenuation:AUTO ON

INPut<1|2>:ATTenuation:AUTO:MODE NORMal | LNOise | LDIStortion

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

RFAttenuation_not_used

Variable Type       ViInt32

Selects the attenuation mode.

Valid Range:

0 - Auto Normal (Default Value)

1 - Auto Low Noise

2 - Auto Low Distortion

Note:

(1) This control is reserved for compatibility with FSEx driver  and

does not apply to the instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (RF Attenuation (not used)) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRFInEAtt

ViStatus rsfsu_confRFInEAtt (ViSession instrumentHandle, ViInt32 window,

ViReal64 manualAttenuatorValuedB);

Purpose

This function programs the attenuation of the electronic input

attenuator. The attenuation can be varied in 5 dB steps from 0 to 30 dB.

Other entries are rounded to the next lower integer value. If the

attenuation is programmed directly, the coupling to the reference level

is switched off. If the defined reference level cannot be set with the

given RF attenuation, this level is adapted to the maximum possible

value.

Notes:

(1) The electronic attenuator is swiched off in the default state.

(2) The function is only available with the electronic attenuator option

FSU-B25.

Remote-control command(s):

INPut<1|2>:EATT

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

manualAttenuatorValuedB

Variable Type       ViReal64

This control sets the manual input electronic attenuation level.

Valid Range: 0.0 to 30.0 dB

Default Value: 0.0 dB

Note:

The attenuation can be varied in 5 dB steps.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Manual Attenuator Value [) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRFInEAttAuto

ViStatus rsfsu_confRFInEAttAuto (ViSession instrumentHandle,

ViInt32 window,

ViBoolean autoAttenuator);

Purpose

This function automatically couples the electronic input attenuation to

the reference level and the attenuation of the mechanical attenuator

(state ON) or switches the input attenuation to manual entry (state OFF).

Note:

The function is only available with the electronic attenuator option

FSU-B25.

Remote-control command(s):

INPut<1|2>:EATT:AUTO ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

autoAttenuator

Variable Type       ViBoolean

This control automatically couples the electronic input attenuation

to the reference level and the attenuation of the mechanical

attenuator (state ON) or switches the input attenuation to manual

entry (state OFF).

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Auto Attenuator) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRFInEAttState

ViStatus rsfsu_confRFInEAttState (ViSession instrumentHandle,

ViInt32 window,

ViBoolean attenuatorState);

Purpose

This function switches the electronic input attenuation into the signal

path (state ON) or removes it from the signal path (state OFF).

Note:

The function is only available with the option electronic attenuator

FSU-B25.

Remote-control command(s):

INPut<1|2>:EATT:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

attenuatorState

Variable Type       ViBoolean

This control switches the electronic input attenuation into the

signal path (state ON) or removes it from the signal path (state

OFF).

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Attenuator State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRFInGain

ViStatus rsfsu_confRFInGain (ViSession instrumentHandle, ViInt32 window,

ViBoolean preamplifierState);

Purpose

This function switches on the preamplifier for the instrument. The

switchable gain is fixed to 20 dB.

Note:

The function is only available with the option electronic attenuator

FSU-B25.

Remote-control command(s):

INPut<1|2>:GAIN:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

preamplifierState

Variable Type       ViBoolean

This control switches the electronic input attenuation into the

signal path (state ON) or removes it from the signal path (state

OFF).

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Preamplifier State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRFInImp

ViStatus rsfsu_confRFInImp (ViSession instrumentHandle, ViInt32 window,

ViInt32 inputImpedance);

Purpose

This function configures the input impedance of the instrument's RF

input.

Remote-control command(s):

INPut<1|2>:IMPedance 50 | 75

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

inputImpedance

Variable Type       ViInt32

Specifies the nominal input impedance.

Valid Range:

0 - 50 Ohm (Default Value)

1 - 75 Ohm

2 - RESERVED

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Input Impedance) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRFInputCoupling

ViStatus rsfsu_confRFInputCoupling (ViSession instrumentHandle,

ViInt32 window, ViBoolean coupling);

Purpose

This function switches the input coupling of the RF input between AC and

DC.

Remote-control command(s):

INPut<1|2>:COUPling AC | DC

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

coupling

Variable Type       ViBoolean

This control switches the input coupling of the RF input between AC

and DC.

Valid Range:

VI_FALSE (0) - DC (Default Value)

VI_TRUE  (1) - AC

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Coupling) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confRFInputProtectRst

ViStatus rsfsu_confRFInputProtectRst (ViSession instrumentHandle,

ViInt32 window);

Purpose

This instrument is equipped with an overload protection mechanism. This

mechanism becomes active as soon as the power at the input mixer exceeds

a value of 27 dBm. It ensures that the connection between RF input and

input mixer is cut off. The function resets the attenuator into the state

that it had before the overload condition was detected. It re-connects

the RF input with the input mixer.

Note:

This function comes into effect only if the reason for the overload

condition has been eliminated. Otherwise the connection between RF input

and input mixer is left open.

Remote-control command(s):

INPut<1|2>:ATTenuation:PROTection:RESet

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMCarrierCount

ViStatus rsfsu_confSAMCarrierCount (ViSession instrumentHandle,

ViInt32 window, ViInt32 count);

Purpose

This function selects the number of carrier signals.

Note:

This function is available only for multicarrier channel and

adjacent-channel power measurements (CALC:MARK:FUNC:POW:SEL MCAC) in the

frequency domain (span > 0).

Remote-control command(s):

SENSe<1|2>:POWer:ACHannel:TXCHannel:COUNt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

count

Variable Type       ViInt32

Selects the number of carrier signals.

Valid Range: 1 to 12

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Count) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMChannelSpacing

ViStatus rsfsu_confSAMChannelSpacing (ViSession instrumentHandle,

ViInt32 window, ViReal64 spacing);

Purpose

This functions defines the channel spacing for the carrier signals.

Note:

(1) This function is only available in the frequency domain (span > 0).

Remote-control command(s):

SENSe<1|2>:POWer:ACHanel:SPACing:CHANnel

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

spacing

Variable Type       ViReal64

Defines the channel spacing for the carrier signals.

Valid Range: 100.0 Hz to 2.0e9 Hz

Default Value: 20.0e3 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Spacing) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMDeltaMarkFixRef

ViStatus rsfsu_confSAMDeltaMarkFixRef (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViBoolean referenceFixed);

Purpose

This function activates/deactivates reference fixed marker.

Remote-control command(s):

CALCulate<1|2>:DELTamarker<1..4>:FUNCtion:FIXed:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

referenceFixed

Variable Type       ViBoolean

Toggles between reference fixed on and off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Reference Fixed) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMDeltaMarkFixRefMax

ViStatus rsfsu_confSAMDeltaMarkFixRefMax (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function sets the reference point level for all delta markers in the

selected measurement window for a measurement with fixed reference point.

(CALC:DELT:FUNC:FIX:STAT ON) to the peak of the selected trace.

For phase-noise measurements (CALCulate:DELTamarker:FUNCtion:PNOise:STATe

ON), the command defines a new reference point level for delta marker 2

in the selected measurement window.

Remote-control command(s):

CALCulate<1|2>:DELTamarker:FUNCtion:FIXed:RPOint:MAXimum:PEAK

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMDeltaMarkFixRefPoint

ViStatus rsfsu_confSAMDeltaMarkFixRefPoint (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViInt32 referencePointSel,

ViReal64 refPointValue);

Purpose

This function configures reference point for delta marker measurement

reference fixed.

Note:

(1) This function is active only when reference fixed marker is switched

On.

Remote-control command(s):

CALCulate<1|2>:DELTamarker<1..4>:FUNCtion:FIXed:RPOint:X

CALCulate<1|2>:DELTamarker<1..4>:FUNCtion:FIXed:RPOint:Y

CALCulate<1|2>:DELTamarker<1..4>:FUNCtion:FIXed:RPOint:Y:OFFSet

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

referencePointSel

Variable Type       ViInt32

If corresponding number is set, the appropriate parameter of the

reference point parameter is changed.

Valid Range:

0 - Ref Point Level (default Value)

1 - Ref Point Level Offset

2 - Ref Point Freq Hz Or Time

refPointValue

Variable Type       ViReal64

Sets the reference point level, reference point level offset,

reference point frequency.

Valid Range: -200.0 to 200.0 dBm (or in selected display unit)

Default Value: 0.0 dBm

--

Sets the reference point level offset.

Valid Range: -100.0 to 100.0 dB

Default Value: 0.0 dB

---

Sets the reference point frequency.

Valid Range for Frequency Domain: 0.0 to 7.0e9 Hz

Valid Range for Time Domain: -100.0 sec to 16000.0 sec

Default Value: 3.0E9 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Reference Point Sel) out of range.

BFFC0005  Parameter 5 (Ref Point Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMDetector

ViStatus rsfsu_confSAMDetector (ViSession instrumentHandle,

ViInt32 traceNumber,

ViInt32 traceDetector);

Purpose

This function controls the recording of measurement values via the type

of detector selected for each trace.

Note:

(1) This function is reserved for compatibility with FSEx driver. Prefer

to use function rsfsu_confSAMDetectorExt.

Remote-control command(s):

SENSe:DETector<1|2|3>:FUNCtion:AUTO ON | OFF

SENSe:DETector<1|2|3>:FUNCtion APEak

SENSe:DETector<1|2|3>:FUNCtion NEGative

SENSe:DETector<1|2|3>:FUNCtion POSitive

SENSe:DETector<1|2|3>:FUNCtion SAMPle

SENSe:DETector<1|2|3>:FUNCtion RMS

SENSe:DETector<1|2|3>:FUNCtion AVERage

SENSe:DETector<1|2|3>:FUNCtion QPEak

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

traceNumber

Variable Type       ViInt32

Trace number to select detector.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

traceDetector

Variable Type       ViInt32

This command switches the detector for recording of the measured

value.

Valid Range:

0 - Auto Select On (Default Value)

1 - Auto Select Off

2 - Auto Peak

3 - Negative Peak

4 - Positive Peak

5 - Sample

6 - RMS

7 - Average

8 - Q Peak

Notes:

(1) Auto Select On/Off:

This command either couples the detector to the current trace setting

or turns coupling off.

(2) Auto Peak:

The value "APEak" (AutoPeak) displays both the positive peak value

and the negative peak value when noise is present. The positive peak

value is displayed when one signal is present.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Trace Number) out of range.

BFFC0003  Parameter 3 (Trace Detector) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMDetectorExt

ViStatus rsfsu_confSAMDetectorExt (ViSession instrumentHandle,

ViInt32 window, ViInt32 traceNumber,

ViInt32 traceDetector);

Purpose

This function controls the recording of measurement values via the type

of detector selected for each trace.

Remote-control command(s):

SENSe<1|2>:DETector<1..3>:FUNCtion:AUTO ON |OFF

SENSe<1|2>:DETector<1..3>:FUNCtion APEak |NEGative | POSitive | SAMPle |

RMS | AVERage | QPEak

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

traceNumber

Variable Type       ViInt32

Trace number to select detector.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

traceDetector

Variable Type       ViInt32

This command switches the detector for recording of the measured

value.

Valid Range:

0 - Auto Select On (Default Value)

1 - Auto Select Off

2 - Auto Peak

3 - Negative Peak

4 - Positive Peak

5 - Sample

6 - RMS

7 - Average

8 - Q Peak

Notes:

(1) Auto Select On/Off:

This command either couples the detector to the current trace setting

or turns coupling off.

(2) Auto Peak:

The value "APEak" (AutoPeak) displays both the positive peak value

and the negative peak value when noise is present. The positive peak

value is displayed when one signal is present.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Trace Number) out of range.

BFFC0004  Parameter 4 (Trace Detector) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMGateSet

ViStatus rsfsu_confSAMGateSet (ViSession instrumentHandle,

ViInt32 window, ViBoolean gateSweep,

ViReal64 gateLevel_not_used,

ViBoolean gateMode,

ViBoolean gatePolarity,

ViReal64 gateDelays,

ViReal64 gateLengths,

ViBoolean gateSource);

Purpose

This function configures the gate parameters.

Remote-control command(s):

SENSe<1|2>:SWEep:EGATe:TYPE EDGE | LEVel

SENSe<1|2>:SWEep:EGATe:POLarity POSitive | NEGative

SENSe<1|2>:SWEep:EGATe:HOLD

SENSe<1|2>:SWEep:EGATe:SOURce EXTernal | IFPower | RFPower

SENSe<1|2>:SWEep:EGATe ON | OFF

SENSe<1|2>:SWEep:EGATe:LENGth

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

gateSweep

Variable Type       ViBoolean

This control controls the sweep with the external gate signal.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

gateLevel_not_used

Variable Type       ViReal64

This control determines the threshold for the external gate signal.

Valid Range: -5.0 V to +5.0 V

Default Value: 0.0 V

Note:

(1) This control is reserved for compatibility with FSEx driver  and

does not apply to the instrument.

gateMode

Variable Type       ViBoolean

This command sets the type of triggering (level or edge) by the

external gate signal.

Valid Range:

VI_FALSE (0) - Edge

VI_TRUE  (1) - Level (Default Value)

gatePolarity

Variable Type       ViBoolean

This control determines the polarity of the external gate signal.

Valid Range:

VI_FALSE (0) - Negative

VI_TRUE  (1) - Positive (Default Value)

gateDelays

Variable Type       ViReal64

This control defines the delay time between the external gate signal

and the continuation of the sweep.

Valid Range: 0.0 to 100.0 s

Default Value: 0.0 s

gateLengths

Variable Type       ViReal64

In case of edge triggering, this control determines the time interval

in which the analyzer sweeps.

Valid Range: 125.0e-9 to 100.0 s

Default Value: 125.0e-9 s

Note:

(1) Gate Length is only used in edge mode.

gateSource

Variable Type       ViBoolean

This control toggles between external gate signal and RF power

signal.

Valid Values:

0 - IF Power (Default Value)

1 - External

2 - RF Power

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Gate Sweep) out of range.

BFFC0004  Parameter 4 (Gate Level (not used)) out of range.

BFFC0005  Parameter 5 (Gate Mode) out of range.

BFFC0006  Parameter 6 (Gate Polarity) out of range.

BFFC0007  Parameter 7 (Gate Delay [s]) out of range.

BFFC0008  Parameter 8 (Gate Length [s]) out of range.

BFFC0009  Parameter 9 (Gate Source) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkChPowChanChannels

ViStatus rsfsu_confSAMMarkChPowChanChannels (ViSession instrumentHandle,

ViInt32 window,

ViInt32 numberOfAdjacentChannels);

Purpose

This function sets the number of adjacent channels (upper and lower

channel in pairs).The figure 0 stands for pure channel power measurement.

Note:

The function is only available in the frequency domain (span > 0).

Remote-control command(s):

SENSe<1|2>:POWer:ACHannel:ACPairs

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

numberOfAdjacentChannels

Variable Type       ViInt32

Sets the number of adjacent channels to be considered for adjacent

channel measurement.

Valid Range:

0 - 0 (Default Value)

1 - 1

2 - 2

3 - 3

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Number Of Adjacent Channe) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkChPowChanOpt

ViStatus rsfsu_confSAMMarkChPowChanOpt (ViSession instrumentHandle,

ViInt32 window,

ViInt32 numberOfAdjacentChannels,

ViReal64 CHBandwidthHz,

ViReal64 ADJBandwidthHz,

ViReal64 ADJChannelSpacingHz,

ViReal64 ALT1BandwidthHz,

ViReal64 ALT1ChannelSpacingHz,

ViReal64 ALT2BandwidthHz,

ViReal64 ALT2ChannelSpacingHz);

Purpose

This function configures the channel bandwidth and channel spacing for

channel power measurement.

Remote-control command(s):

SENSe<1|2>:POWer:ACHannel:ACPairs

SENSe<1|2>:POWer:ACHannel:BWIDth

SENSe<1|2>:POWer:ACHannel:SPACing:ACHannel

SENSe<1|2>:POWer:ACHannel:BWIDth:ACHannel

SENSe<1|2>:POWer:ACHannel:SPACing:ALTernate<1|2>

SENSe<1|2>:POWer:ACHannel:BWIDth:ALTernate<1|2>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

numberOfAdjacentChannels

Variable Type       ViInt32

Sets the number of adjacent channels to be considered for adjacent

channel measurement.

Valid Range:

0 - 0 (Default Value)

1 - 1

2 - 2

3 - 3

CHBandwidthHz

Variable Type       ViReal64

Specifies channel bandwidth for the transmission channel.

Valid Range: 100.0 Hz to 1.0e9 Hz

Default Value: 24.0e3 Hz

ADJBandwidthHz

Variable Type       ViReal64

Specifies adjacent channel bandwidth.

Valid Range: 100.0 Hz to 1.0e9 Hz

Default Value: 24.0e3 Hz

ADJChannelSpacingHz

Variable Type       ViReal64

This control defines the spacing of the adjacent channel relative to

the carrier signal.

Valid Range: 100.0 Hz to 2.0e9 Hz

Default Value: 24.0e3 Hz

ALT1BandwidthHz

Variable Type       ViReal64

Specifies first alternate channel bandwidth.

Valid Range: 100.0 Hz to 1.0e9 Hz

Default Value: 24.0e3 Hz

ALT1ChannelSpacingHz

Variable Type       ViReal64

This cotrol defines the spacing of the first alternate adjacent

channel  (ALTernate 1) relative to the carrier signal.

Valid Range: 100.0 Hz to 2.0e9 Hz

Default Value: 48.0e3 Hz

ALT2BandwidthHz

Variable Type       ViReal64

Specifies second alternate channel bandwidth.

Valid Range: 100.0 Hz to 1.0e9 Hz

Default Value: 24.0e3 Hz

ALT2ChannelSpacingHz

Variable Type       ViReal64

This control defines the spacing of the second alternate adjacent

channel (ALTernate 2) relative to the carrier signal.

Valid Range: 100.0 Hz to 2.0e9 Hz

Default Value: 72.0e3 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Number Of Adjacent Channe) out of range.

BFFC0004  Parameter 4 (CH Bandwidth [Hz]) out of range.

BFFC0005  Parameter 5 (ADJ Bandwidth [Hz]) out of range.

BFFC0006  Parameter 6 (ADJ Channel Spacing [Hz]) out of range.

BFFC0007  Parameter 7 (ALT1 Bandwidth [Hz]) out of range.

BFFC0008  Parameter 8 (ALT1 Channel Spacing [Hz]) out of range.

BFFC0009  Parameter 9 (ALT2 Bandwidth [Hz]) out of range.

BFFC000A  Parameter 10 (ALT2 Channel Spacing [Hz]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkChPowChanSet

ViStatus rsfsu_confSAMMarkChPowChanSet (ViSession instrumentHandle,

ViInt32 window,

ViBoolean CPACPMode);

Purpose

This function controls relative/absolute mode of channel power

measurement.

Remote-control command(s):

SENSe<1|2>:POWer:ACHannel:MODE ABSolute | RELative

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

CPACPMode

Variable Type       ViBoolean

Toggles between absolute and relative measurement.

Valid Range:

VI_FALSE (0) - Abs (Default Value)

VI_TRUE  (1) - Rel

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (CP/ACP Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkChPowChanStandard

ViStatus rsfsu_confSAMMarkChPowChanStandard (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViInt32 ACPStandard);

Purpose

This function configures the standard for channel power measurement.

Note:

FAST ACP is not available if a WLAN standard is selected.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:POWer:PRESet  NONE | NADC | TETRA |

PDC | PHS | CDPD  | F8CDma | R8CDma | F19CDma | R19CDma | FWCDma | RWCDma

| FW3Gppcdma | RW3Gppcdma | S2CDma | D2CDma | M2CDma | FIS95C0 | RIS95C0

| FIS95C1 | RIS95C1 | TCDMa | AWLan | BWLan

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

ACPStandard

Variable Type       ViInt32

Selects the ACP (adjacent channel power) standard.

Valid Range:

0 - None (Default Value)

1 - NADC

2 - TETRA

3 - PDC

4 - PHS

5 - CDPD

6 - CDMA 800 forward

7 - CDMA 800 reverse

8 - CDMA 1900 forward

9 - CDMA 1900 reverse

10 - WCDMA forward

11 - WCDMA reverse

12 - W-CDMA 3GPP forward

13 - W-CDMA 3GPP reverse

14 - CDMA2000 Multi Carrier with carrier 1

15 - CDMA2000 Direct Sequence

16 - CDMA2000 Multi Carrier with carrier 3

17 - CDMA IS95C Class 0 forward

18 - CDMA IS95C Class 0 reverse

19 - CDMA IS95C Class 1 forward

20 - CDMA IS95C Class 1 reverse

21 - TD-SCDMA

22 - WLAN 802.11A

23 - WLAN 802.11B

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (ACP Standard) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkChPowHiSpeed

ViStatus rsfsu_confSAMMarkChPowHiSpeed (ViSession instrumentHandle,

ViInt32 window,

ViBoolean highSpeedMeasurement);

Purpose

This function switches on or off the high-speed channel/adjacent channel

power measurement. The measurement itself is performed in the time domain

on the center frequencies of the different

channels. The function automatically switches to the time domain and

back. Depending on the selected mobile radio standard, weighting filters

with SQRT(cos) characteristic or very steep-sided channel filters are

used for band limitation. If no standard is selected, weighting is

performed by means of the normal resolution filters.

Notes:

(1) The function is only available in the frequency domain (span > 0).

(2) It is not available if no standard is active.

Remote-control command(s):

SENSe<1|2>:POWer:HSPeed ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

highSpeedMeasurement

Variable Type       ViBoolean

This control switches on or off the high-speed channel/adjacent

channel power measurement.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (High Speed Measurement) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkChPowLim

ViStatus rsfsu_confSAMMarkChPowLim (ViSession instrumentHandle,

ViInt32 window,

ViInt32 numberOfAdjacentChannels,

ViReal64 ADJLowLimit_not_used,

ViReal64 ALT1LowLimit_not_used,

ViReal64 ALT2LowLimit_not_used,

ViReal64 ADJHighLimit,

ViReal64 ALT1HighLimit,

ViReal64 ALT2HighLimit,

ViBoolean ADJLimitState,

ViBoolean ALT1LimitState,

ViBoolean ALT2LimitState);

Purpose

This function configures the limit check for channel power measurement.

Note:

(1) This function is reserved for compatibility with FSEx driver. Prefer

to use function rsfsu_confSAMMarkChPowLimExt.

Remote-control command(s):

SENSe<1|2>:POWer:ACHannel:ACP

CALCulate<1|2>:LIMit:ACP:ALTernate<1|2>

CALCulate<1|2>:LIMit:ACP:ACHannel

CALCulate<1|2>:LIMit:ACP:ACHannel:STATe ON | OFF

CALCulate<1|2>:LIMit:ACP:ALTernate<1|2>:STATe ON | OFF

CALCulate<1|2>:LIMit:ACP:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

numberOfAdjacentChannels

Variable Type       ViInt32

Sets the number of adjacent channels to be considered for adjacent

channel measurement.

Valid Range:

0 - 0 (Default Value)

1 - 1

2 - 2

3 - 3

ADJLowLimit_not_used

Variable Type       ViReal64

Sets the lower limit for adjacent channel power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ALT1LowLimit_not_used

Variable Type       ViReal64

Sets the lower limit for the first alternate adjacent channel power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ALT2LowLimit_not_used

Variable Type       ViReal64

Sets the lower limit for the second alternate adjacent channel power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ADJHighLimit

Variable Type       ViReal64

Sets the upper limit for adjacent channel power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ALT1HighLimit

Variable Type       ViReal64

Sets the upper limit for the first alternate adjacent channel power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ALT2HighLimit

Variable Type       ViReal64

Sets the upper limit for the second alternate adjacent channel power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ADJLimitState

Variable Type       ViBoolean

Enables or disables limit check for the adjacent channel power.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

ALT1LimitState

Variable Type       ViBoolean

Enables or disables limit check for the first alternate adjacent

channel power.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

ALT2LimitState

Variable Type       ViBoolean

Enables or disables limit check for the second alternate adjacent

channel power.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Number Of Adjacent Channe) out of range.

BFFC0004  Parameter 4 (ADJ Low Limit (not used)) out of range.

BFFC0005  Parameter 5 (ALT1 Low Limit (not used)) out of range.

BFFC0006  Parameter 6 (ALT2 Low Limit (not used)) out of range.

BFFC0007  Parameter 7 (ADJ High Limit) out of range.

BFFC0008  Parameter 8 (ALT1 High Limit) out of range.

BFFC0009  Parameter 9 (ALT2 High Limit) out of range.

BFFC000A  Parameter 10 (ADJ Limit State) out of range.

BFFC000B  Parameter 11 (ALT1 Limit State) out of range.

BFFC000C  Parameter 12 (ALT2 Limit State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkChPowLimExt

ViStatus rsfsu_confSAMMarkChPowLimExt (ViSession instrumentHandle,

ViInt32 window, ViInt32 measMode,

ViInt32 numberOfAdjacentChannels,

ViReal64 ADJRelativeLimitdB,

ViReal64 ALT1RelativeLimitdB,

ViReal64 ALT2RelativeLimitdB,

ViReal64 ADJAbsoluteLimitdBm,

ViReal64 ALT1AbsoluteLimitdBm,

ViReal64 ALT2AbsoluteLimitdBm,

ViBoolean ADJLimitState,

ViBoolean ALT1LimitState,

ViBoolean ALT2LimitState,

ViBoolean ACPLimitCheck);

Purpose

This function configures the limit check for channel power measurement.

Remote-control command(s):

SENSe<1|2>:POWer:ACHannel:ACP

CALCulate<1|2>:LIMit:ACPower:ACHannel:ABSolute

CALCulate<1|2>:LIMit:ACPower:ACHannel:RELative

CALCulate<1|2>:LIMit:ACPower:ALTernate<1|2>:RELative

CALCulate<1|2>:LIMit:ACPower:ALTernate<1|2>:ABSolute

CALCulate<1|2>:LIMit:ACPower:ACHannel:RELative:STATe ON | OFF

CALCulate<1|2>:LIMit:ACPower:ACHannel:ABSolute:STATe ON | OFF

CALCulate<1|2>:LIMit:ACPower:ALTernate<1|2>:RELative:STATe ON | OFF

CALCulate<1|2>:LIMit:ACPower:ALTernate<1|2>:ABSolute:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

measMode

Variable Type       ViInt32

Selects absolute and/or relative measurement.

Valid Range:

0 - Set Rel (Default Value)

1 - Set Abs

2 - Set Abs and Rel

Note:

(1) Set Rel:

Sets up relative limit checking values and state.

(2) Set Abs:

Sets up absolute limit checking values and state.

(3) Set Abs and Rel:

Sets up both absolute and relative limit checking values and state.

numberOfAdjacentChannels

Variable Type       ViInt32

Sets the number of adjacent channels to be considered for adjacent

channel measurement.

Valid Range:

0 - 0 (Default Value)

1 - 1

2 - 2

3 - 3

ADJRelativeLimitdB

Variable Type       ViReal64

Sets the relative limit for adjacent channel power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ALT1RelativeLimitdB

Variable Type       ViReal64

Sets the relative limit for the first alternate adjacent channel

power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ALT2RelativeLimitdB

Variable Type       ViReal64

Sets the relative limit for the second alternate adjacent  channel

power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ADJAbsoluteLimitdBm

Variable Type       ViReal64

Sets the absolute limit for adjacent channel power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ALT1AbsoluteLimitdBm

Variable Type       ViReal64

Sets the absolute limit for the first alternate adjacent channel

power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ALT2AbsoluteLimitdBm

Variable Type       ViReal64

Sets the absolute limit for the second alternate adjacent channel

power.

Valid Range: -500.0 dB to 500.0 dB

Default Value: 0.0 dB

ADJLimitState

Variable Type       ViBoolean

Enables or disables limit check for the adjacent channel power.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

ALT1LimitState

Variable Type       ViBoolean

Enables or disables limit check for the first alternate adjacent

channel power.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

ALT2LimitState

Variable Type       ViBoolean

Enables or disables limit check for the second alternate adjacent

channel power.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

ACPLimitCheck

Variable Type       ViBoolean

Enables or disables ACP limit checking.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Meas Mode) out of range.

BFFC0004  Parameter 4 (Number Of Adjacent Channe) out of range.

BFFC0005  Parameter 5 (ADJ Relative Limit [dB]) out of range.

BFFC0006  Parameter 6 (ALT1 Relative Limit [dB]) out of range.

BFFC0007  Parameter 7 (ALT2 Relative Limit [dB]) out of range.

BFFC0008  Parameter 8 (ADJ Absolute Limit [dBm]) out of range.

BFFC0009  Parameter 9 (ALT1 Absolute Limit [dBm]) out of range.

BFFC000A  Parameter 10 (ALT2 Absolute Limit [dBm]) out of range.

BFFC000B  Parameter 11 (ADJ Limit State) out of range.

BFFC000C  Parameter 12 (ALT1 Limit State) out of range.

BFFC000D  Parameter 13 (ALT2 Limit State) out of range.

BFFC000E  Parameter 14 (ACP Limit Check) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkChPowMeas

ViStatus rsfsu_confSAMMarkChPowMeas (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViInt32 channelPowerType,

ViBoolean channelPowerFunctionMode);

Purpose

This function sets the channel power measurement modes.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:POWer OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:POWer:SELect CPOWer | CN | CN0 |

ACPower | OBWidth | MCACpower

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

channelPowerType

Variable Type       ViInt32

Specifies the channel power function to measure.

Valid Range:

0 - Channel Power (Default Value)

1 - C / N

2 - C / No

3 - Adjacent Channel Power

4 - Occupied Bandwidth Power

5 - Multi Carrier Adjacent Channel Power

channelPowerFunctionMode

Variable Type       ViBoolean

Sets the operation performed by this function.

Valid Range:

VI_FALSE (0) - Measurement Off (Default Value)

VI_TRUE  (1) - Measurement On

Notes:

(1) Measurement Off - turns off power measurement.

(2) Measurement On - turns on power measurement specified by the

control Channel Power Type.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Channel Power Type) out of range.

BFFC0005  Parameter 5 (Channel Power Function Mo) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkChPowOccPowBW

ViStatus rsfsu_confSAMMarkChPowOccPowBW (ViSession instrumentHandle,

ViInt32 window,

ViReal64 powerBandwidthPerc);

Purpose

This function configures the bandwidth for occupied bandwidth

measurement.

Remote-control command(s):

SENSe<1|2>:POWer:BANDwidth

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

powerBandwidthPerc

Variable Type       ViReal64

Defines the percentage of the power with respect to the total power

in the displayed frequency range. This defines the occupied bandwidth

(percent of the total power).

Valid Range: 10.0 % to 99.9 %

Default Value: 99.0 %

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Power Bandwidth [Perc]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkDemod

ViStatus rsfsu_confSAMMarkDemod (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViReal64 markerStopTimes,

ViBoolean demodulationState,

ViBoolean modulationType,

ViReal64 speakerVolumePerc);

Purpose

This function configures the demodulation markers. The sweep stops at the

frequency determined by the set marker for the selected time and the RF

signal is demodulated.

Note:

This function is available only if option FSU-B3 is installed.

Remote-control command(s):

SYSTem:SPEaker:VOLume

CALCulate<1|2>:MARKer<1..4>:FUNCtion:DEModulation:HOLD

CALCulate<1|2>:MARKer<1..4>:FUNCtion:DEModulation:STATe ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:DEModulation:SELect AM |FM

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

markerStopTimes

Variable Type       ViReal64

Specifies the time for which the sweep will be stopped at the marker

frequency.

Valid Range: 10.0 ms to 1000.0 s

Default Value: 0.5 s

demodulationState

Variable Type       ViBoolean

Turns on or off the demodulation.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

modulationType

Variable Type       ViBoolean

Specifies the modulation type.

Valid Range:

VI_FALSE (0) - Frequency Modulation (Default Value)

VI_TRUE  (1) - Amplitude Modulation

speakerVolumePerc

Variable Type       ViReal64

Specifies the speaker volume.

Valid Range: 0.0 % to 100.0 %

Default Value: 50.0 %

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Marker Stop Time [s]) out of range.

BFFC0005  Parameter 5 (Demodulation State) out of range.

BFFC0006  Parameter 6 (Modulation Type) out of range.

BFFC0007  Parameter 7 (Speaker Volume [Perc]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkFreqCnt

ViStatus rsfsu_confSAMMarkFreqCnt (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber_not_used,

ViInt32 counterResolution,

ViBoolean signalCounterState);

Purpose

This function configures the frequency counter. The frequency counter

measures the frequency of the RF signal at the intermediate frequency.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:COUNt ON | OFF

CALCulate<1|2>:MARKer<1..4>:COUNt:RESolution 0.1 | 1 | 10 | 100 | 1000

|10000

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber_not_used

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

Note:

(1) This control is reserved for compatibility with FSEx driver  and

does not apply to the instrument.

counterResolution

Variable Type       ViInt32

Specifies the counter resolution. The time which the frequency

counter requires for a measurement is proportional to the selected

resolution.

Valid Range:

0 - 0.1 Hz

1 - 1.0 Hz

2 - 10.0 Hz

3 - 100.0 Hz

4 - 1.0 kHz (Default Value)

5 - 10.0 kHz

signalCounterState

Variable Type       ViBoolean

Turns on or off the frequency counter.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number (not used)) out of range.

BFFC0004  Parameter 4 (Counter Resolution) out of range.

BFFC0005  Parameter 5 (Signal Counter State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkModDepth

ViStatus rsfsu_confSAMMarkModDepth (ViSession instrumentHandle,

ViInt32 window,

ViBoolean modDepthMeasState);

Purpose

This Function switches on the measurement of the AM modulation depth. An

AM-modulated carrier is required on the screen for correct functioning.

If necessary, marker 1 is previously activated and set to the largest

signal available.

Notes:

(1) The level value of marker 1 is taken as the carrier level. On

activating the function, marker 2 and marker 3 are automatically set to

the adjacent maxima of the trace as delta markers symetrically to the

carrier and marker 2 is activated for the entry.

(2) If the position of delta marker 2 is changed, delta marker 3 is moved

symetrically to the reference marker (marker 1). If the position of delta

marker 3 is changed, fine adjustment can then be performed independently

of delta marker 2.

(3) FSU calculates the power at the marker positions from the measured

levels. The AM modulation depth is calculated from the ratio of power

values at the reference marker and the delta markers. If the two AM

sidebands differ in power, the average value of the two power values is

used for calculating the AM modulation depth.

Remote-control command(s):

CALCulate<1|2>:MARKer:FUNCtion:MDEPth ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

modDepthMeasState

Variable Type       ViBoolean

Turns On or Off the AM modulation depth measurement.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Mod Depth Meas State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkNoise

ViStatus rsfsu_confSAMMarkNoise (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViBoolean noiseMeasurementState);

Purpose

This function configures the noise power density measurement.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:NOISe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

noiseMeasurementState

Variable Type       ViBoolean

Turns on or off the noise measurement.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Noise Measurement State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkPhaseNoise

ViStatus rsfsu_confSAMMarkPhaseNoise (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViBoolean phaseNoiseMeasurementState);

Purpose

This function configures the phase noise measurement.

Remote-control command(s):

CALCulate<1|2>:DELTamarker<1..4>:FUNCtion:PNOise ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

phaseNoiseMeasurementState

Variable Type       ViBoolean

Turns on or off the phase noise measurement.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Notes:

(1) If set to On, Reference Fixed (Config Delta Marker Fixed

reference) is automatically set to On.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Phase Noise Measurement S) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkSearchLimits

ViStatus rsfsu_confSAMMarkSearchLimits (ViSession instrumentHandle,

ViInt32 window,

ViInt32 searchLimit,

ViReal64 searchLimitValue);

Purpose

This function allows separate control of threshold or each search limit

line.

Remote-control command(s):

CALCulate<1|2>:MARKer:X:SLIMits?

CALCulate<1|2>:MARKer:X:SLIMits ON | OFF

CALCulate<1|2>:THReshold:STATe ON | OFF

CALCulate<1|2>:MARKer:X:SLIMits:LEFT

CALCulate<1|2>:MARKer:X:SLIMits:RIGHt

CALCulate<1|2>:THReshold

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

0 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

searchLimit

Variable Type       ViInt32

Selects search limit parameter.

Valid Range:

0 - Off

1 - Left Limit

2 - Right Limit

3 - Threshold

Default Value: 0

Notes:

(1) Off:

This command switches threshold line (OFF) and unlimited (OFF) search

range in the selected measurement window. If the time domain power

measurement is active, this command limits the evaluation range on

the trace.

(2) Left Limit:

This command sets the left limit of the search range for markers and

delta markers in the selected measurement window. Depending on the x

axis domain the indicated value defines a frequency (span > 0) or

time (span = 0). If the time domain power measurement is active, this

command limits the evaluation range on the trace.

(3) Right Limit:

This command sets the right limit of the search range for markers and

delta markers in the selected measurement window. Depending on the x

axis domain the indicated value defines a frequency (span > 0) or

time (span = 0). If the time domain power measurement is active, this

command limits the evaluation range to the trace.

(4) Threshold:

This command switches on the threshold line in the selected

measurement window. This command also defines the threshold value for

the maximum/minimum search of markers with marker search functions

MAX PEAK, NEXT PEAK, etc in the selected measurement window. The

associated display line is automatically switched on.

searchLimitValue

Variable Type       ViReal64

This control sets value for parameter selected via Search Limit

control.

Valid Range when Search Limit is:

Off:         not applied

Left Limit:  0.0 to MAX (frequency | sweep time)

Right Limit: 0.0 to MAX (frequency | sweep time)

Threshold:   MIN to MAX (depending on current unit)

Default Value: 0.0

Note:

Range checking for this parameter is not applied.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Search Limit) out of range.

BFFC0004  Parameter 4 (Search Limit Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkSearchNdB

ViStatus rsfsu_confSAMMarkSearchNdB (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViBoolean n_dBDown,

ViReal64 n_dBDownValuedB);

Purpose

This function configures the measurement function ndB down.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:NDBD:STATe ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:NDBD

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

0 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

n_dBDown

Variable Type       ViBoolean

Activates and deactivates temporary markers which are located n dB

below the active reference marker. The value n is entered using N dB

Down Value.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

n_dBDownValuedB

Variable Type       ViReal64

Enters the N dB value.

Valid Range: 0.0 dB to 100.0 dB

Default Value: 6.0 dB

Notes:

(1) This control is active only if N dB Down is set to On.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (N dB Down) out of range.

BFFC0005  Parameter 5 (N dB Down Value [dB]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkSearchParam

ViStatus rsfsu_confSAMMarkSearchParam (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViReal64 peakExcursiondB,

ViBoolean excludeLO,

ViBoolean searchLimit);

Purpose

This function configures the marker search operations.

Note:

(1) This function is reserved for compatibility with FSEx driver. Prefer

to use function rsfsu_confSAMMarkSearchParamExt.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:LOEXclude ON | OFF

CALCulate<1|2>:MARKer<1..4>:PEXCursion

CALCulate<1|2>:MARKer<1..4>:X:SLIMits ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

0 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

peakExcursiondB

Variable Type       ViReal64

Specifies the minimum amount a signal level must decrease or increase

before it is recognized by the search function as a minimum or

maximum.

Valid Range: 0.0 dB to 80.0 dB

Default Value: 6.0 dB

excludeLO

Variable Type       ViBoolean

Switches between a restricted (On) and unrestricted search range.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

searchLimit

Variable Type       ViBoolean

Switches between a limited (On) and unlimited search range.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Peak Excursion [dB]) out of range.

BFFC0005  Parameter 5 (Exclude LO) out of range.

BFFC0006  Parameter 6 (Search Limit) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMMarkSearchParamExt

ViStatus rsfsu_confSAMMarkSearchParamExt (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViReal64 peakExcursiondB,

ViBoolean excludeLO,

ViBoolean searchLimits,

ViReal64 leftSearchLimit,

ViReal64 rightSearchLimit,

ViReal64 threshold);

Purpose

This function configures the marker search operations.

Remote-control command(s):

INSTrument<1|2>:SELect?

CALCulate<1|2>:MARKer<1..4>:LOEXclude ON | OFF

CALCulate<1|2>:MARKer<1..4>:PEXCursion

CALCulate<1|2>:MARKer<1..4>:X:SLIMits ON | OFF

CALCulate<1|2>:MARKer<1..4>:X:SLIMits:LEFT

CALCulate<1|2>:MARKer<1..4>:X:SLIMits:RIGHT

CALCulate<1|2>:THReshold:STATe ON | OFF

CALCulate<1|2>:THReshold

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

0 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

peakExcursiondB

Variable Type       ViReal64

Specifies the minimum amount a signal level must decrease or increase

before it is recognized by the search function as a minimum or

maximum.

Valid Range: 0.0 dB to 80.0 dB

Default Value: 6.0 dB

excludeLO

Variable Type       ViBoolean

Switches between a restricted (On) and unrestricted search range.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

searchLimits

Variable Type       ViBoolean

Switches between a limited (On) or unlimited (Off) searching range.

This control also switches the Threshold line.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

leftSearchLimit

Variable Type       ViReal64

This control sets the left limit of the search range for markers and

delta markers in the selected measurement window to the indicated

frequency (span > 0) or time (span = 0). The function is independent

of the selection of a marker.

Valid Range: not checked

Default Value: 0.0

Notes:

(1) If the time domain power measurement is active, this command

limits the evaluation range of the trace.

(2) The settings in this control will only be accepted if Search

Limits is set to On.

rightSearchLimit

Variable Type       ViReal64

This command sets the right limit of the search range for markers and

delta markers in the selected measurement window to the indicated

frequency (span > 0) or time (span = 0).

Valid Range: not checked

Default Value: 0.0

Notes:

(1) If the time domain power measurement is active, this command

limits the evaluation range to the trace.

(2) The settings in this control will only be accepted if Search

Limits is set to On.

threshold

Variable Type       ViReal64

This control sets the threshold value for the maximum/minimum search

of markers with marker search functions MAX PEAK, NEXT PEAK, etc in

the selected measurement window. The associated display line is

switched to ON in case Search Limit is also switched to ON.

Valid Range: not checked

Default Value: 0.0

Note:

(1) The settings in this control will only be accepted if Search

Limits is set to On.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Peak Excursion [dB]) out of range.

BFFC0005  Parameter 5 (Exclude LO) out of range.

BFFC0006  Parameter 6 (Search Limits) out of range.

BFFC0007  Parameter 7 (Left Search Limit) out of range.

BFFC0008  Parameter 8 (Right Search Limit) out of range.

BFFC0009  Parameter 9 (Threshold) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMNoiseCorr

ViStatus rsfsu_confSAMNoiseCorr (ViSession instrumentHandle,

ViInt32 window,

ViBoolean noiseCorrectionState);

Purpose

This function switches on or off the correction of the instrument

inherent noise for channel power measurement. On activating this

function, a reference measurement of the instrument inherent noise is

performed. The measured noise power is then subtracted from the power in

the examined channel. The measurement of the instrument inherent noise is

then repeated after any change of the center frequency, resolution

bandwidth, sweep time and level setting.

Note:

This function is only available in the frequency domain (span > 0).

Remote-control command(s):

SENSe<1|2>:POWer:NCORrection ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

noiseCorrectionState

Variable Type       ViBoolean

Turns on or off the noise correction.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Noise Correction State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMPowerTrace

ViStatus rsfsu_confSAMPowerTrace (ViSession instrumentHandle,

ViInt32 window, ViInt32 traceNumber);

Purpose

This function assigns the channel/adjacent channel power measurement of

the indicated trace in the selected measurement window. The corresponding

trace should be active, ie its states should differ from blank.

Remote-control command(s):

SENSe<1|2>:POWer:TRACe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

traceNumber

Variable Type       ViInt32

Selects trace number.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Trace Number) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMReferenceAuto

ViStatus rsfsu_confSAMReferenceAuto (ViSession instrumentHandle,

ViInt32 window, ViInt32 reference);

Purpose

This function selects the method for auto selection of the reference for

relative adjacent channel power measurement.

Note:

This function is only available in the frequency domain (span > 0).

Remote-control command(s):

SENSe<1|2>:POWer:ACHanel:REFerence:TXCHannel:AUTO MINimum | MAXimum |

LHIGhest

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

reference

Variable Type       ViInt32

This control sets the type of reference selection.

Valid Values:

0 - Minimum (Default Value)

1 - Maximum

2 - Highest

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Reference) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMReferenceMan

ViStatus rsfsu_confSAMReferenceMan (ViSession instrumentHandle,

ViInt32 window, ViInt32 channel);

Purpose

This function manually selects the reference channel for relative

adjacent channel power measurement.

Note:

This function is only available in the frequency domain (span > 0).

Remote-control command(s):

SENSe<1|2>:POWer:ACHanel:REFerence:TXCHannel:MANual

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

channel

Variable Type       ViInt32

This control select the reference channel for relative adjacent

channel power measurement.

Valid Range: 1 to 12

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Channel) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMSigStat

ViStatus rsfsu_confSAMSigStat (ViSession instrumentHandle,

ViInt32 window_RESERVED,

ViInt32 statisticsFunction,

ViInt32 no_ofSamples);

Purpose

This function controls the statistical measurements in the instrument.

Remote-control command(s):

CALCulate:STATistics:APD ON | OFF

CALCulate:STATistics:CCDF ON | OFF

CALCulate:STATistics:NSAMples

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window_RESERVED

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

(2) This parameter is reserved for future use.

statisticsFunction

Variable Type       ViInt32

This control switches statistics functions on/off and selects the

corresponding number of measurement samples.

Valid Range:

0 - All Off (Default Value)

1 - APD On

2 - CCDF On

3 - APD Off

4 - CCDF Off

Notes:

(1) All Off:

Switches Off APD and CCDF measurements if they are active.

(2) APD On / Off:

Switches On / Off the measurement of amplitude distribution (APD). On

activating this function, the CCDF measurement is switched off.

(3) CCDF On / Off:

Switches On / Off the measurement of the complementary cumulative

distribution function (CCDF). On activating this function, the APD

mesurement is switched off.

(4) All measurements will be executed in time domain (span = 0 Hz).

no_ofSamples

Variable Type       ViInt32

This control sets the number of test points to be recorded for the

statistical measurement functions.

Valid Range: 100 to 1e9

Default Value: 100000

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window (RESERVED)) out of range.

BFFC0003  Parameter 3 (Statistics Function) out of range.

BFFC0004  Parameter 4 (No of Samples) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMSigStatScal

ViStatus rsfsu_confSAMSigStatScal (ViSession instrumentHandle,

ViInt32 window_RESERVED,

ViInt32 scalingMode,

ViReal64 scalingValue);

Purpose

This function changes the scaling parameters for statistical

measurements.

Remote-control command(s):

CALCulate:STATistic:PRESet

CALCulate:STATistic:SCALe:AUTO ONCE

CALCulate:STATistic:SCALe:X:RLEVel

CALCulate:STATistic:SCALe:X:RANGe

CALCulate:STATistic:SCALe:Y:UPPer

CALCulate:STATistic:SCALe:Y:LOWer

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window_RESERVED

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

(2) This parameter is reserved for future use.

scalingMode

Variable Type       ViInt32

This control selects scaling mode.

Valid Range:

0 - Default Settings (Default value)

1 - Adjust Settings

2 - X-axis Ref Level

3 - X-axis Range

4 - Y-axis Max Value

5 - Y-axis Min Value

Notes:

(1) Default Settings:

This command resets the scaling of the X- and Y-axes in a statistical

measurement. The following values are set:

X-axis ref level: -20 dBm

X-axis range APD: 100 dB

X-axis range CCDF: 20 dB

Y-axis upper limit: 1.0

Y-axis lower limit: 1E-6

(2) Adjust Settings:

This command optimizes the level setting of FSU depending on the

measured peak power to obtain maximum instrument sensitivity. To

obtain maximum resolution, the level range is set as a function of

the measured spacing between peak power and the minimum power with

the APD measurement and of the spacing between the peak power and the

mean power with CCDF measurement. In addition, the probability scale

for the number of test points is adapted.

(3) X-axis Ref Level:

This command defines the reference level for the X-axis of the

measurement diagram. The setting is identical to the reference level

setting using the command DISPlay:WINDow:TRACe:Y:RLEVel.

With the reference level offset <> 0 the indicated value range of the

reference level is modified by the offset.

(4) X-axis Range:

This command defines the level range for the X-axis of the

measurement diagram.

(5) Y-axis Max Value:

This command defines the upper limit for the Y-axis of the diagram in

statitistical measurements. Since probabilities are specified on the

Y-axis, the entered numerical values are dimensionless.

(6) Y-axis Min Value:

This command defines the lower limit for th Y-axis of the diagram in

statistical measurements. Since probabilities are specified on the

Y-axis, the entered numerical values are dimensionless.

scalingValue

Variable Type       ViReal64

This control sets the scaling values for the selected scaling mode.

Valid Ranges:

X-axis Ref Level: -130.0 dBm to 30.0 dBm

X-axis Range: 10.0 dB to 200.0 dB

Y-axis Max Value: 1.0E-8 to 1.0

Y-axis Min Value: 1.0E-9 to 0.1

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window (RESERVED)) out of range.

BFFC0003  Parameter 3 (Scaling Mode) out of range.

BFFC0004  Parameter 4 (Scaling Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMTraceMath

ViStatus rsfsu_confSAMTraceMath (ViSession instrumentHandle,

ViInt32 window, ViInt32 traceNumber,

ViBoolean mathState,

ViString mathExpression);

Purpose

This function allows to process data from the SENSe-subsystem in numeric

expressions.

Note:

(1) This function is reserved for compatibility with FSEx driver. Prefer

to use function rsfsu_confSAMTraceMathExt.

Remote-control command(s):

CALCulate<1|2>:MATH

CALCulate<1|2>:MATH:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

traceNumber

Variable Type       ViInt32

Selects the trace to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

mathState

Variable Type       ViBoolean

This command switches the mathematical relation of traces on or off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

mathExpression

Variable Type       ViString

This control defines the mathematical expression for relating traces

and reference line.

Valid Range: not checked

Default Value: ""

Notes:

Parameter: <expr>::= "OP1 - OP2"

OP1 ::= TRACE1

OP2 ::= TRACE2 | TRACE3

Example:

"TRACE1-TRACE3"

The strings must be exact such as above listed.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Trace Number) out of range.

BFFC0004  Parameter 4 (Math State) out of range.

BFFC0005  Parameter 5 (Math Expression) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMTraceMathExt

ViStatus rsfsu_confSAMTraceMathExt (ViSession instrumentHandle,

ViInt32 window, ViBoolean mathState,

ViString mathExpression,

ViBoolean mathMode,

ViReal64 mathPosition);

Purpose

This function allows to process data from the SENSe-subsystem in numeric

expressions.

Remote-control command(s):

CALCulate<1|2>:MATH:STATe ON |OFF

CALCulate<1|2>:MATH

CALCulate<1|2>:MATH:MODE LINear | LOGarithmic

CALCulate<1|2>:POSition

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

mathState

Variable Type       ViBoolean

This command switches the mathematical relation of traces on or off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

mathExpression

Variable Type       ViString

This control defines the mathematical expression for relating traces

and trace1.

Valid Range: not checked

Default Value: ""

Notes:

Parameter: <expr>::= "OP1 - OP2"

OP1 ::= TRACE1

OP2 ::= TRACE2 | TRACE3

Example:

"TRACE1-TRACE3"

The strings must be exact such as above listed.

mathMode

Variable Type       ViBoolean

This control selects linear or logarithmic (= video) calculation of

the mathematical functions related to the traces. The calculation of

the average is one of such functions.

Valid Range:

VI_FALSE (0) - Log (Default Value)

VI_TRUE  (1) - Lin

mathPosition

Variable Type       ViReal64

This control defines the position of the result of the trace

mathematics in the selected measurement window. The indication is in

% of the screen height, 100% corresponding to the upper

diagram edge.

Valid Range: -100.0 to 200.0 %

Default Value: 50.0 %

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Math State) out of range.

BFFC0004  Parameter 4 (Math Expression) out of range.

BFFC0005  Parameter 5 (Math Mode) out of range.

BFFC0006  Parameter 6 (Math Position [%]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMTrigCount

ViStatus rsfsu_confSAMTrigCount (ViSession instrumentHandle,

ViInt32 window, ViInt32 count);

Purpose

This function sets the trigger count.

Notes:

(1) This function is reserved for compatibility with previous versions of

the driver. Prefer to use function Channel Power Carrier Count

(rsfsu_confSAMCarrierCount).

(2) This function is only available in the frequency domain (span > 0).

Remote-control command(s):

SENSe<1|2>:POWer:ACHannel:TXCHannel:COUNt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

count

Variable Type       ViInt32

This conrol specifies the trigger count.

Valid Range:

1 to 12

Default Value:

1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Count) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMTrigSpacing

ViStatus rsfsu_confSAMTrigSpacing (ViSession instrumentHandle,

ViInt32 window, ViReal64 spacing);

Purpose

This function sets the channel spacing of adjacent channel to trigger

signal.

Notes:

(1) This function is reserved for compatibility with previous versions of

the driver. Prefer to use function Channel Power Channel Spacing

(rsfsu_confSAMChannelSpacing).

(2) This function is only available in the frequency domain (span > 0).

Remote-control command(s):

SENSe<1|2>:POWer:ACHanel:SPACing:CHANnel

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

spacing

Variable Type       ViReal64

This conrol specifies the channel spacing of adjacent channel to

trigger signal.

Valid Range:

100 Hz to 2.0E+9 Hz

Default Value:

20.0E+3 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Spacing) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSAMVibw

ViStatus rsfsu_confSAMVibw (ViSession instrumentHandle, ViInt32 window,

ViReal64 bwValueHz);

Purpose

This function controls the setting of the analyzer's video filter

bandwidths. Automatically sets the coupling to manual.

Remote-control command(s):

SENSe<1|2>:BANDwidth:VIDeo

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

bwValueHz

Variable Type       ViReal64

Sets manually the video bandwidth.

Valid Range: 1.0 Hz to 10.0e6 Hz

Default Value: 1.0e6 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Bw Value [Hz]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confScreenCpl

ViStatus rsfsu_confScreenCpl (ViSession instrumentHandle,

ViInt32 coupling);

Purpose

This function defines the coupling between the two measurement windows

screen A and B.

Remote-control command(s):

INSTrument:COUPle NONE | RLEVel | CF_B | CF_A

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

coupling

Variable Type       ViInt32

This control defines the coupling between the two measurement windows

screen A and B.

Valid Range:

0 - None (Default Value)

1 - Ref Level

2 - Center Freq of B

3 - Center Freq of A

(1) None:

No coupling. The two measurement windows are operated like two

independent "virtual" devices. Switches off the coupling of

measurement windows. Two independent "virtual" devices are then

obtained.

(2) Ref Level:

The reference levels of the two measurement windows are

coupled.

(3) Center Freq of B:

The center frequency of screen B is coupled to the frequency of

marker 1 in screen A.

(4) Center Frequency of A:

The center frequency of screen A is coupled to the frequency of

marker 1 in screen B.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Coupling) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEAtt

ViStatus rsfsu_confSEAtt (ViSession instrumentHandle, ViInt32 range,

ViReal64 manualAttenuatorValuedB);

Purpose

This function defines the RF attenuation of a range in the spurious

measurement.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:INPut:ATTenuation:AUTO OFF

[SENSe<1|2>:]LIST:RANGe<1..20>:INPut:ATTenuation <numeric_value>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

manualAttenuatorValuedB

Variable Type       ViReal64

This control sets the RF attenuation of a range

Valid Range: 0.0 dB to 75.0 dB

Default Value: 15.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Manual Attenuator Value [) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEAttAuto

ViStatus rsfsu_confSEAttAuto (ViSession instrumentHandle,

ViInt32 range);

Purpose

This function switches the autoranging of a range in the spurious

measurement on.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:INPut:ATTenuation:AUTO ON

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEBreakSweep

ViStatus rsfsu_confSEBreakSweep (ViSession instrumentHandle,

ViInt32 range, ViBoolean state);

Purpose

This function determines whether a sweep is stopped in the spurious

measurement when a range is changed.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:BREAK

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

state

Variable Type       ViBoolean

This control sets the break state. It determines whether a sweep is

stopped in the spurious measurement when a range is changed.

Valid Values:

VI_FALSE: Off (Default Value)

VI_TRUE:  On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEDeleteRange

ViStatus rsfsu_confSEDeleteRange (ViSession instrumentHandle,

ViInt32 range);

Purpose

This function deletes a range.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:DELete

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control specifies the range that will be deleted.

Valid Values: 1 to 20

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEDetector

ViStatus rsfsu_confSEDetector (ViSession instrumentHandle,

ViInt32 range, ViInt32 detector);

Purpose

This function sets the detector in the spurious measurement.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:DETector APEak | NEGative | POSitive |

SAMPle | RMS | AVERage

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

detector

Variable Type       ViInt32

This control sets the detector in the spurious measurement.

Valid Values:

0 - Auto Peak

1 - Negative

2 - Positive (Default Value)

3 - Sample

4 - RMS

5 - Average

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Detector) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEFilter

ViStatus rsfsu_confSEFilter (ViSession instrumentHandle, ViInt32 range,

ViInt32 filter);

Purpose

This function sets the filter in the spurious measurement.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:FILTer:TYPE NORMal | CHANnel | RRC | P5

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

filter

Variable Type       ViInt32

This control sets the filter in the spurious measurement.

Valid Values:

0 - Normal (Default Value)

1 - Channel

2 - RRC

3 - P5

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Filter) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEMLimitLine

ViStatus rsfsu_confSEMLimitLine (ViSession instrumentHandle,

ViInt32 measMode,

ViReal64 expectedPower);

Purpose

This function switches on or off the automatic selection of limit lines

for spectrum emission mask measurements.

Note:

(1) This function is available only in conjunction with options FS-K72,

FS-K73, FS-K76, FS-K77, FS-K82, FS-K83, FS-K84 or FS-K85.

Remote-control command(s):

CALCulate:LIMit:ESPectrum:MODE AUTO | MANual | USER

CALCulate:LIMit:ESPectrum:VALue

CALCulate:LIMit:ESPectrum:RESTore

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measMode

Variable Type       ViInt32

This control switches on or off the automatic selection of limit

lines for spectrum emission mask measurements.

Valid Range:

0 - Auto (Default Value)

1 - Manual

2 - Restore

3 - User

Notes:

(1) Auto:

Set limit line depends on the measured channel power.

(2) Manual:

One of fixed limit lines is activated. The limit line is selected

using Expected Power control. This selection is not supported with

FS-K77, FS-K83 and FS-K85 option.

(3) Restore:

Restores the standard limit lines for the spectrum emission mask

measurement. All changes made to the standard limit lines are lost

and the state of these limit lines as they left the factory is

restored.

(4) User:

User-defined limit lines are enabled.

expectedPower

Variable Type       ViReal64

The limit line is selected by entering the expected power as a value.

Depending on the value entered, one of four possible limit lines is

selected:

Valid Range (FS-K72, FS-K73): coerced by instrument

Expected power in dBm  Selected imit line  Retrieved value

value >= 43            "P >= 43"           43.0

39 <= value < 43       "39 <= P < 43"      39.0

31 <= value < 39       "31 <= P < 39"      31.0

value < 31             "P < 31"             0.0

Valid Range (FS-K76): coerced by instrument

Expected power in dBm  Selected imit line  Retrieved value

value >= 34            "P >= 34"           34.0

26 <= value < 34       "26 <= P < 34"      26.0

value < 26             "P < 26"             0.0

Valid Range (FS-K82, FS-K84): coerced by instrument

Expected power in dBm  Selected imit line  Retrieved value

value >= 33            "P >= 33"           33.0

28 <= value < 33       "28 <= P < 33"      28.0

value < 28             "P < 28"             0.0

Default Value: 0.0 dBm

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Meas Mode) out of range.

BFFC0003  Parameter 3 (Expected Power) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEPreamplifier

ViStatus rsfsu_confSEPreamplifier (ViSession instrumentHandle,

ViInt32 range,

ViBoolean preamplifier);

Purpose

This function switches the preamplifier of a range in the spurious

measurement on or off.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:INPut:GAIN:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

preamplifier

Variable Type       ViBoolean

Switches the preamplifier of a range on and off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Pre-amplifier) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSERefLevel

ViStatus rsfsu_confSERefLevel (ViSession instrumentHandle,

ViInt32 range, ViReal64 refLevelValue);

Purpose

This function defines the reference level of a range in the spurious

measurement.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:RLEVel <numeric_value>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

refLevelValue

Variable Type       ViReal64

Specifies the reference level of a range in the spurious measurement.

Valid Range: -130.0 dBm to 30.0 dBm

Default Value: -30.0 dBm

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Ref Level Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEResolutionBW

ViStatus rsfsu_confSEResolutionBW (ViSession instrumentHandle,

ViInt32 range,

ViReal64 resolutionBandwidth);

Purpose

This function selects the resolution bandwidth (RBW) of a range in the

spurious measurement.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:BANDwidth[:RESolution]

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

resolutionBandwidth

Variable Type       ViReal64

Sets the the resolution bandwidth,

Valid Range: 10.0 Hz to 50.0e6 Hz

Default Value: 10.0e3 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Resolution Bandwidth) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSESearchPeaks

ViStatus rsfsu_confSESearchPeaks (ViSession instrumentHandle,

ViInt32 subrange, ViReal64 margin);

Purpose

This function determines the list of the subrange maximums from the

existing sweep results.

Note(s):

(1) The list can be read with command TRACe? SPURious

(rsfsu_dataReadTraceSE).

Remote-control command(s):

CALCulate:PEAK:SUBR

CALCulate:PEAK:MARG

CALCulate:PEAK

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

subrange

Variable Type       ViInt32

This control sets he number of peaks per range that are stored in the

list. Once the selected number of peaks has been reached, the peak

search is stopped in the current range and continued in the next

range.

Valid Values: 1 to 50

Default Value: 25

margin

Variable Type       ViReal64

This control sets the margin, i.e. the acceptance threshold for

determining the peak list. The limit line is shifted by this value

when the maximums are determined.

Valid Values: -200.0 dB to 200.0 dB

Default Value: 6.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Subrange) out of range.

BFFC0003  Parameter 3 (Margin) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEStartStopFrq

ViStatus rsfsu_confSEStartStopFrq (ViSession instrumentHandle,

ViInt32 range,

ViInt32 frequencyParameter,

ViReal64 startFrequencyHz,

ViReal64 stopFrequencyHz);

Purpose

This function configures start and stop frequencies of a range in the

spurious measurement.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>[:FREQuency]:STARt

[SENSe<1|2>:]LIST:RANGe<1..20>[:FREQuency]:STOP

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

frequencyParameter

Variable Type       ViInt32

Specifies which parameter will be set.

Valid Range:

0 - Start And Stop (Default Value)

1 - Start Only

2 - Stop Only

startFrequencyHz

Variable Type       ViReal64

Sets the start frequency of a range in the spurious measurement.

Valid Range: not checked

Default Value: 20.0 Hz

stopFrequencyHz

Variable Type       ViReal64

Sets the stop frequency of a range in the spurious measurement.

Valid Range: not checked

Default Value: 1.0e+9 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Frequency Parameter) out of range.

BFFC0004  Parameter 4 (Start Frequency [Hz]) out of range.

BFFC0005  Parameter 5 (Stop Frequency [Hz]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSESweepMode

ViStatus rsfsu_confSESweepMode (ViSession instrumentHandle,

ViInt32 mode);

Purpose

This function controls the sweep continuation.

Remote-control command(s):

[SENSe<1|2>:]SWEep:MODE AUTO | LIST

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

mode

Variable Type       ViInt32

This control selects the sweep continuation.

Valid Range:

0 - Auto (Default Value)

1 - List

Notes:

(1) Auto:

Free-running sweep

(2) List:

Activates the spurious emissions measurement.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSESweepPoints

ViStatus rsfsu_confSESweepPoints (ViSession instrumentHandle,

ViInt32 range, ViInt32 sweepPoints);

Purpose

This function defines the number of sweep points of a range in the

spurious measurement.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:POINts

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

sweepPoints

Variable Type       ViInt32

Sets the number of sweep points of a range in the spurious

measurement.

Valid Range: 155, 313, 625, 1251, 1999, 2501, 5001, 10001, 20001

Default Value: 625

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Sweep Points) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSESweepTime

ViStatus rsfsu_confSESweepTime (ViSession instrumentHandle,

ViInt32 range, ViReal64 sweepTime);

Purpose

This function defines the sweep duration of a range in the spurious

measurement.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:SWEep:TIME <numeric_value>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

sweepTime

Variable Type       ViReal64

Sets the sweep duration of a range in the spurious measurement.

Valid Range: 0.001 s to 16000.0 s

Default Value: 0.1 s

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Sweep Time) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSESweepTimeAuto

ViStatus rsfsu_confSESweepTimeAuto (ViSession instrumentHandle,

ViInt32 range, ViBoolean state);

Purpose

This function the automatic coupling of the sweep duration on the

frequency span and the bandwidth settings in a spurious measurement

range.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:SWEep:TIME:AUTO ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

state

Variable Type       ViBoolean

Sets the automatic coupling of the sweep duration on the frequency

span and the bandwidth settings in a spurious measurement range.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSetActiveWindow

ViStatus rsfsu_confSetActiveWindow (ViSession instrumentHandle,

ViInt32 window);

Purpose

This function selects the active measurement window.

Functionality depends on the instruments operating modes where:

(1) Analyzer Mode, WLAN Mode (FSQ-K90/K91 option):

WINDOW 1 corresponds to SCREEN A, WINDOW 2 to SCREEN B. In FULL SCREEN

mode, the measurements are only performed in the active measurement

window. Measurements are therefore triggered in the active window and

result queries (marker, trace data and other results) answered also in

the active window. Triggering measurements and querying results in the

inactive window yields an error message (execution error). In split

screen mode, the selection of the active window for result queries is

irrelevant.

Note:

In FULL SCREEN mode, settings can also be performed in the inactive

measurement window. They become effective as soon as the corresponding

window becomes active.

(2) All measurement modes (SAN, VSA, GSM/EDGE, WCDP, ...):

The associated measurement window for all trace data functions is

selected with the numeric suffix of command ":TRACe<1|2>[:DATA] ..."

(Trace data depends on the window parameter). To allow select active

measurement window within the existing trace data functions, call this

function. WINDOW 1 (SCREEN A) is a default value when the communication

with the instrument is established (default instrument setup).

Note:

Possibility to select associated measurement window is mentioned in all

functions related to this feature.

Remote-control command(s):

INSTrument:NSELect?

DISPlay:<WINDow[1|2]>:SELect

DISPlay:<WINDow[1|2]>:SSELect

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be used as active measurement

(associated) window.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSetPrinter

ViStatus rsfsu_confSetPrinter (ViSession instrumentHandle,

ViInt32 deviceNumber,

ViString printerName);

Purpose

This function selects one of the printers installed under Windows NT.

Prior to use this function get list of available printers using function

Get Printer List.

Remote-control command(s):

SYSTem:COMMunicate:PRINter:SELect

SYSTem:COMMunicate:PRINter:SELect<1|2>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

deviceNumber

Variable Type       ViInt32

Selects the device to be configured.

Valid Range:

1 - 1 (Default Value)

2 - 2

Note:

(1) Device Number 2 is currently not supported.

printerName

Variable Type       ViString

This control sets printer name.

Valid Range: up to 128 chars

Default Value: ""

Note:

Valid names are taken using function Get Printer List.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Device Number) out of range.

BFFC0003  Parameter 3 (Printer Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSETransducer

ViStatus rsfsu_confSETransducer (ViSession instrumentHandle,

ViInt32 range, ViString factor);

Purpose

This function sets the transducer factor for a range in the spurious

measurement.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:TRANsducer

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

factor

Variable Type       ViString

Sets the transducer factor for a range in the spurious measurement.

Valid Range: not checked

Default Value: ""

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Factor) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSEVideoBW

ViStatus rsfsu_confSEVideoBW (ViSession instrumentHandle, ViInt32 range,

ViReal64 bandwidth);

Purpose

This function selects the video bandwidth (RBW) of a range in the

spurious measurement.

Remote-control command(s):

[SENSe<1|2>:]LIST:RANGe<1..20>:BANDwidth:VIDeo

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

range

Variable Type       ViInt32

This control sets the spurious measurement range.

Valid Values: 1 to 20

Default Value: 1

bandwidth

Variable Type       ViReal64

Sets the the video bandwidth,

Valid Range: 1.0 Hz to 10.0e6 Hz

Default Value: 30.0e3 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Range) out of range.

BFFC0003  Parameter 3 (Bandwidth) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSignalTrack

ViStatus rsfsu_confSignalTrack (ViSession instrumentHandle,

ViInt32 window, ViInt32 trace,

ViInt32 signalTrackFunction,

ViReal64 bandwidthValue,

ViReal64 thresholdValue);

Purpose

This function switches the signal-track for the selected measurement

window on or off and sets the track bandwidth and threshold to the

indicated values. The function is independent of the selected marker; it

will use the indicated trace for searching the signal.

Notes:

(1) With signal track activated, the maximum signal is determined after

each frequency sweep and the center frequency of this signal is set. With

drifting signals the center frequency follows the signal.

(2) Using the function in single sweep mode requires the function Send

Trigger and Wait for OPC to be used in order to synchronize subsequent

commands correctly to the end of the track action.

(3) This function is only available in the frequency domain (span > 0).

Remote-control command(s):

CALCulate<1|2>:MARKer:FUNCtion:STRack:TRACe

CALCulate<1|2>:MARKer:FUNCtion:STRack:STATe OFF | ON

CALCulate<1|2>:MARKer:FUNCtion:STRack:BANDwidth

CALCulate<1|2>:MARKer:FUNCtion:STRack:THReshold

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

trace

Variable Type       ViInt32

This control defines the trace on which the largest signal is

searched for in the selected measurement window. It is only available

in the frequency domain (span > 0).

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

signalTrackFunction

Variable Type       ViInt32

This control selects signal track operation or parameter to set.

Valid Range:

0 - Signal Track Off (Default Value)

1 - Signal Track On

2 - Bandwidth Value

3 - Threshold Value

4 - Set All

Notes:

(1) Signal Track On/Off:

With signal track activated, the maximum signal is determined after

each frequency sweep and the center frequency of this signal is set.

With drifting signals the center frequency follows the signal.

(2) Bandwidth Value:

This value defines the bandwidth at the center frequency within which

the largest signal is searched for in the selected measurement

window. Prior to set this value switch signal track to ON.

(3) Threshold Value:

This value defines the threshold above which the largest signal is

searched for in the selected measurement window. Prior to set this

value switch signal track to ON.

(4) Set All:

Sets bandwidth and threshold values together with signal track to On.

bandwidthValue

Variable Type       ViReal64

This control inputs bandwidth value.

Valid Range: 10.0 Hz to MAX (SPAN) Hz

Default Value: 10.0 Hz

Note:

Range is not checked.

thresholdValue

Variable Type       ViReal64

This control inputs  threshold value.

Valid Range: -330.0 dBm to +30.0 dBm

Default Value: 0.0

Note:

Range will not be checked. The threshold value depends on the level

units which the instrument is currently using.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Trace) out of range.

BFFC0004  Parameter 4 (Signal Track Function) out of range.

BFFC0005  Parameter 5 (Bandwidth Value) out of range.

BFFC0006  Parameter 6 (Threshold Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSlotsMulti

ViStatus rsfsu_confSlotsMulti (ViSession instrumentHandle,

ViInt32 slot);

Purpose

This function defines the used slots of the mobile.

The multislot setting defines how many adjacent slots are active and

which of the active slots should be used for synchronization.

Remote-control command(s):

CONFigure:CHANnel:SLOT:MULTi ACT1SYNC1 | ACT2SYNC1 | ACT2SYNC2 |

ACT3SYNC1 | ACT3SYNC2 | ACT3SYNC3 | ACT4SYNC1 | ACT4SYNC2 | ACT4SYNC3 |

ACT4SYNC4

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

slot

Variable Type       ViInt32

This control defines the active slots and synchronization slot.

Valid Range:

0 - ACT1SYNC1 (Default Value)

1 - ACT2SYNC1

2 - ACT2SYNC2

3 - ACT3SYNC1

4 - ACT3SYNC2

5 - ACT3SYNC3

6 - ACT4SYNC1

7 - ACT4SYNC2

8 - ACT4SYNC3

9 - ACT4SYNC4

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Slot) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSpanCenterFrq

ViStatus rsfsu_confSpanCenterFrq (ViSession instrumentHandle,

ViInt32 window,

ViInt32 frequencyParameter,

ViReal64 frequencySpanHz,

ViReal64 centerFrequencyHz);

Purpose

This function configures frequency span and center frequency for the

spectrum analyzer.

Remote-control command(s):

SENSe<1|2>:FREQuency:CENTer

SENSe<1|2>:FREQuency:SPAN

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

frequencyParameter

Variable Type       ViInt32

Specifies which parameter will be set.

Valid Range:

0 - Span And Center (Default Value)

1 - Span Only

2 - Center Only

Note:

In Vector Analyzer mode "Center Only" is possible to select.

frequencySpanHz

Variable Type       ViReal64

Sets the frequency span of the analyzer.

Valid Range:

FSU 3  - 0.0 Hz to 3.6 GHz

FSU 8  - 0.0 Hz to 8.0 GHz

FSU 26 - 0.0 Hz to 26.5 GHz

FSU 46 - 0.0 Hz to 46.0 GHz

FSU 50 - 0.0 Hz to 50.0 GHz

Default Value: 1.5E9 Hz

centerFrequencyHz

Variable Type       ViReal64

Sets the center frequency of the analyzer.

Valid Range (Only for Frequency Offset = 0.0):

FSU 3  - 0.0 Hz to 3.6 GHz

FSU 8  - 0.0 Hz to 8.0 GHz

FSU 26 - 0.0 Hz to 26.5 GHz

FSU 46 - 0.0 Hz to 46.0 GHz

FSU 50 - 0.0 Hz to 50.0 GHz

Default Value: 1.5E9 Hz

Note:

(1) This control is not range checked.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Frequency Parameter) out of range.

BFFC0004  Parameter 4 (Frequency Span [Hz]) out of range.

BFFC0005  Parameter 5 (Center Frequency [Hz]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confStartStopFrq

ViStatus rsfsu_confStartStopFrq (ViSession instrumentHandle,

ViInt32 window,

ViInt32 frequencyParameter,

ViReal64 startFrequencyHz,

ViReal64 stopFrequencyHz);

Purpose

This function configures start and stop frequencies for the spectrum

analyzer.

Remote-control command(s):

SENSe<1|2>:FREQuency:STARt

SENSe<1|2>:FREQuency:STOP

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

frequencyParameter

Variable Type       ViInt32

Specifies which parameter will be set.

Valid Range:

0 - Start And Stop (Default Value)

1 - Start Only

2 - Stop Only

startFrequencyHz

Variable Type       ViReal64

Sets the start frequency of the analyzer.

Valid Range (Only for Frequency Offset = 0.0):

FSU 3  - 0.0 Hz to 3.6 GHz

FSU 8  - 0.0 Hz to 8.0 GHz

FSU 26 - 0.0 Hz to 26.5 GHz

FSU 46 - 0.0 Hz to 46.0 GHz

FSU 50 - 0.0 Hz to 50.0 GHz

Default Value: 0.0 Hz

Note:

(1) This control is not range checked.

stopFrequencyHz

Variable Type       ViReal64

Sets the stop frequency of the analyzer.

Valid Range (Only for Frequency Offset = 0.0):

FSU 3  - 0.0 Hz to 3.6 GHz

FSU 8  - 0.0 Hz to 8.0 GHz

FSU 26 - 0.0 Hz to 26.5 GHz

FSU 46 - 0.0 Hz to 46.0 GHz

FSU 50 - 0.0 Hz to 50.0 GHz

Default Value: 3.0E9 Hz

Note:

(1) This control is not range checked.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Frequency Parameter) out of range.

BFFC0004  Parameter 4 (Start Frequency [Hz]) out of range.

BFFC0005  Parameter 5 (Stop Frequency [Hz]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSumMark

ViStatus rsfsu_confSumMark (ViSession instrumentHandle, ViInt32 window,

ViInt32 markerNumber,

ViInt32 summMkrFunction,

ViBoolean stateSwitch);

Purpose

This function configures the time domain power measurement.

Note:

This function selects absolute time domain power measurement in the

indicated measurement window.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:AOFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary OFF | ON

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:MODE ABSolute

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:RMS ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:MEAN ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:MAXimum ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:PPEak ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:MPEak ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:MIDDle ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:PHOLd ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:AVERage ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:AOFFset ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:SDEViation ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

summMkrFunction

Variable Type       ViInt32

Selects desired summary marker function.

Valid Range:

0 - Time Domain Power State (Default Value)

1 - RMS State

2 - Mean State

3 - RESERVED

4 - Positive Peak

5 - RESERVED

6 - RESERVED

7 - Peak Hold State

8 - Average State

9 - Summ All Off

10 - Standard Deviation

Notes:

(1) Time Domain Power State:

This command switches on or off the time domain power measurement in

total. It is only available in the time domain (span = 0).

(2) RMS State:

This command switches on or off the measurement of the effective

(rms) value of the total trace. It is only available in the time

domain (span = 0).

(3) Mean State:

This command switches on or off the measurement of the mean value of

the total trace. It is only available in the time domain (span = 0).

(4) Positive Peak:

This command switches on or off the measurement of the positive peak

value of the total trace. It is only available in the time domain

(span = 0).

(5) Standard Deviation:

This command switches on or off the measurement of the standard

deviation of the total trace in the selected measurement window. It

is only available in the time domain (span = 0).

(6) Peak Hold State:

This command switches on or off the peak-hold function. It is only

available in the time domain (span = 0).

(7) Average State:

This command switches the calculation of the average value on or off.

It is only available in the time domain (span = 0).

(8) Summ All Off:

This command switches off all measuring functions. It is only

available in the time domain (span = 0).

stateSwitch

Variable Type       ViBoolean

Switches between On and Off states of summary marker functions.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Summ Mkr Function) out of range.

BFFC0005  Parameter 5 (State Switch) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSumMarkRel

ViStatus rsfsu_confSumMarkRel (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 summMkrFunction,

ViBoolean stateSwitch);

Purpose

This function configures the time domain power measurements. This

function is only available for measurements in the time domain (SPAN =

0).

Notes:

(1) This function selects relative time domain power measurement in the

indicated measurement window.

(2) The reference power for relative measurement is defined with

function Set Reference. If the reference power is not defined explicitly,

the value 0 is used.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:AOFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary OFF | ON

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:MODE RELative

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:RMS ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:MEAN ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:MAXimum ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:PPEak ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:MPEak ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:MIDDle ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:PHOLd ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:AVERage ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:AOFFset ON | OFF

CALCulate<1|2>:MARKer<1..4>:FUNCtion:SUMMary:SDEViation ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

4 - 4

summMkrFunction

Variable Type       ViInt32

Selects desired summary marker function.

Valid Range:

0 - Time Domain Power State (Default Value)

1 - RMS State

2 - Mean State

3 - RESERVED

4 - Positive Peak

5 - RESERVED

6 - RESERVED

7 - Peak Hold State

8 - Average State

9 - Summ All Off

10 - Standard Deviation

Notes:

(1) Time Domain Power State:

This command switches on or off the time domain power measurement in

total. This function is only available in the time domain (span = 0).

(2) RMS State:

This command switches on or off the measurement of the effective

(rms) value of the total trace. It is only available in the time

domain (span = 0).

(3) Mean State:

This command switches on or off the measurement of the mean value of

the total trace. It is only available in the time domain (span = 0).

(4) Positive Peak:

This command switches on or off the measurement of the positive peak

value of the total trace. It is only available in the time domain

(span = 0).

(5) Standard Deviation:

This command switches on or off the measurement of the standard

deviation of the total trace in the selected measurement window. It

is only available in the time domain (span = 0).

(6) Peak Hold State:

This command switches on or off the peak-hold function. It is only

available in the time domain (span = 0).

(7) Average State:

This command switches the calculation of the average value on or off.

It is only available in the time domain (span = 0).

(8) Summ All Off:

This command switches off all measuring functions. It is only

available in the time domain (span = 0).

stateSwitch

Variable Type       ViBoolean

Switches between On and Off states of summary marker functions.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Summ Mkr Function) out of range.

BFFC0005  Parameter 5 (State Switch) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSweep

ViStatus rsfsu_confSweep (ViSession instrumentHandle, ViInt32 window,

ViInt32 sweep, ViInt32 sweepCountValue);

Purpose

This function configures the sweep subsystem.

Remote-control command(s):

INITiate<1|2>:CONTinuous ON |OFF

INITiate<1|2>:DISPlay ON |OFF

SENSe<1|2>:SWEep:COUNt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

sweep

Variable Type       ViInt32

This control selects the sweep parameter.

Valid Range:

0 - Continuous Sweep (Default Value)

1 - Single Sweep

2 - Sweep Count Value

3 - Single Sweep Disp On

4 - Single Sweep Disp Off

Notes:

(1) Continuous Sweep, Single Sweep:

This command determines if the trigger system is continuously

initiated ("Free Run").

(2) Sweep Count Value:

Allows to set up sweep count value.

(3) Single Sweep Disp On/Off:

This command switches the display on or off during a single sweep.

sweepCountValue

Variable Type       ViInt32

This control defines the number of sweeps started with single sweep.

Valid Range: 0 to 32767

Default Value: 0

Note:

(1) This parameter defines the number of sweeps or the number of

averaging procedures. In the average mode, the value 0 defines a

running averaging of measurement data over 10 sweeps.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Sweep) out of range.

BFFC0004  Parameter 4 (Sweep Count Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSweepPoints

ViStatus rsfsu_confSweepPoints (ViSession instrumentHandle,

ViInt32 window, ViInt32 sweepPoints);

Purpose

This function defines the number of measurement points collected during a

sweep.

Remote-control command(s):

SENSe<1|2>:SWEep:POINts

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

sweepPoints

Variable Type       ViInt32

Sets the number of measurement points collected during a sweep.

Valid Range: 155, 313, 625, 1251, 2501, 5001, 10001

Default Value: 625

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Sweep Points) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSweepTime

ViStatus rsfsu_confSweepTime (ViSession instrumentHandle,

ViInt32 window, ViReal64 sweepTime);

Purpose

This function controls the setting of the analyzer's sweeptime.

Automatically sets the coupling to manual.

Note:

If SWEep:TIME is directly programmed, automatic coupling is switched off.

Remote-control command(s):

SENSe<1|2>:SWEep:TIME

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

sweepTime

Variable Type       ViReal64

Sets manually the sweeptime.

Valid Range:

5 ms to  1000 sec  for frequency domain

1 us to 16000 sec  for time domain

Default Value: 0.005 sec

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Sweep Time) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSweepTimeAuto

ViStatus rsfsu_confSweepTimeAuto (ViSession instrumentHandle,

ViInt32 window,

ViBoolean sweepTimeAuto);

Purpose

This function controls the automatic coupling of the sweep time to the

frequency span and bandwidth settings.

Note:

If SWEep:TIME is directly programmed, automatic coupling is switched off.

Remote-control command(s):

[SENSe<1|2>:]SWEep:TIME:AUTO ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

sweepTimeAuto

Variable Type       ViBoolean

Controls the automatic coupling of the sweep time to the frequency

span and bandwidth settings.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On

Default Value: VI_TRUE (1)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Sweep Time Auto) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confSystemTime

ViStatus rsfsu_confSystemTime (ViSession instrumentHandle,

ViBoolean timeDate, ViInt32 hour,

ViInt32 minute, ViInt32 second,

ViInt32 day, ViInt32 month,

ViInt32 year);

Purpose

This function configures the instrument's internal clock and  calendar.

Remote-control command(s):

SYSTem:TIME

SYSTem:DATE

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

timeDate

Variable Type       ViBoolean

This control selects TIME or DATE to be updated.

Valid Range:

VI_FALSE (0) - Date

VI_TRUE  (1) - Time (Default Value)

hour

Variable Type       ViInt32

Sets the clock hour.

Valid Range: 0 to 23

Default Value: 0

minute

Variable Type       ViInt32

Sets the clock minutes.

Valid Range: 0 to 59

Default: 0

second

Variable Type       ViInt32

Sets the clock seconds.

Valid Range: 0 to 59

Default Value: 0

day

Variable Type       ViInt32

Sets the calendar day.

Valid Range: 1 to 31

Default Value: 1

month

Variable Type       ViInt32

Sets calendar month.

Valid Range:

1 - January (Default Value)

2 - February

3 - March

4 - April

5 - May

6 - June

7 - July

8 - August

9 - September

10 - October

11 - November

12 - December

year

Variable Type       ViInt32

Sets the year for the date function.

Valid Range: 1980 to 2099

Default Value: 1998

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Time/Date) out of range.

BFFC0003  Parameter 3 (Hour) out of range.

BFFC0004  Parameter 4 (Minute) out of range.

BFFC0005  Parameter 5 (Second) out of range.

BFFC0006  Parameter 6 (Day) out of range.

BFFC0007  Parameter 7 (Month) out of range.

BFFC0008  Parameter 8 (Year) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTOI

ViStatus rsfsu_confTOI (ViSession instrumentHandle, ViInt32 window,

ViBoolean stateSwitch);

Purpose

This function initiates the measurement of the third-order intercept. A

two-tone signal with equal carrier levels is expected at the input of the

FSU. Marker 1 and marker 2 (both normal markers) are set to the maximum

of the two signals. Delta marker 3 and delta marker 4 are positioned to

the intermodulation products. FSU calculates the third-order intercept

from the level spacing between the normal markers and the delata markers.

Remote-control command(s):

CALCulate<1|2>:MARKer:FUNCtion:TOI ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

stateSwitch

Variable Type       ViBoolean

Switches between On and Off states of TOI.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (State Switch) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrace

ViStatus rsfsu_confTrace (ViSession instrumentHandle, ViInt32 window,

ViInt32 traceNumber, ViInt32 traceMode);

Purpose

This function defines the type of display and the evaluation of the

traces.

Remote-control command(s):

DISPlay:WINDow<1|2>:TRACe<1..3>:STATe OFF

DISPlay:WINDow<1|2>:TRACe<1..3>:MODE:HCONtinuous OFF | ON

DISPlay:WINDow<1|2>:TRACe<1..3>:MODE:ANALog OFF | ON

DISPlay:WINDow<1|2>:TRACe<1..3>:MODE WRITe | VIEW | AVERage | MAXHold |

MINHold

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

traceNumber

Variable Type       ViInt32

Selects the trace to configure.

Valid Range:

1 - 1 (Default Value)

2 - 2

3 - 3

traceMode

Variable Type       ViInt32

This control defines the type of display and the evaluation of the

traces.

Valid Range:

0 - Clear/Write (Default Value)

1 - View

2 - Average

3 - Blank

4 - Max Hold

5 - Min Hold

6 - RESERVED

7 - RESERVED

8 - RESERVED

9 - RESERVED

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Trace Number) out of range.

BFFC0004  Parameter 4 (Trace Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTraceIQ

ViStatus rsfsu_confTraceIQ (ViSession instrumentHandle,

ViReal64 samplingRateHz,

ViInt32 triggerMode, ViBoolean triggerSlope,

ViReal64 bandwidthHz,

ViInt32 pretriggerSamples,

ViInt32 no_ofSamples);

Purpose

This function defines the settings of the analyzer hardware for the

measurement of I/Q data. This allows setting the bandwidth of the analog

filters in front of the A/D converter as well as setting the sample rate,

trigger conditions and the record length.

Notes:

(1) Function Trace I/Q State (confTraceIQState) has to be switched ON

before using this function.

Remote-control command(s):

TRACe:IQ:SET

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

samplingRateHz

Variable Type       ViReal64

This control sets the sampling rate for the I/Q data acquisition.

Valid Range: 10.0e3 Hz to 100.0e6 Hz

Value range with bandwidth extension (R&S FSQ-B72): 10.0e3 Hz to

326.4e6 Hz

Default Value: 32.0e6 Hz

Notes:

(1) Typical values are: 15.625 kHz, 31.25 kHz, 62.5 kHz, 125 kHz, 250

kHz, 500 kHz, 1 MHz, 2 MHz, 4 MHz, 8 MHz, 16 MHz, 32 MHz.

(2) At sampling rates > 81.6 MHz, a bandwidth of 50 MHz must always

be specified.

triggerMode

Variable Type       ViInt32

This parameter selects the trigger source for starting data

acquisition.

Valid Range:

0 - Immediate (Default Value)

1 - External

triggerSlope

Variable Type       ViBoolean

This parameter selects the slope of the trigger signal for external

trigger.

Valid Range:

VI_FALSE (0) - Negative

VI_TRUE  (1) - Positive (Default Value)

Note:

Positive value is currently the only value supported.

bandwidthHz

Variable Type       ViReal64

This control sets bandwidth of the analog filters in front of the A/D

converter.

Valid Range: 300.0e3 to 10.0e6 Hz in steps of 1, 3, 10

Default Value: 10.0e6 Hz

pretriggerSamples

Variable Type       ViInt32

This control sets the number of measurement values to be recorded

before the trigger point.

Valid Range: -16744447 to 65023

Default Value: 0

Note:

For <trigger mode> = Immediate the value must be 0.

no_ofSamples

Variable Type       ViInt32

This control sets the number of measurement values to record.

Valid Range: 1 to 523776 (= 512*1024 - 512)

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Sampling Rate [Hz]) out of range.

BFFC0003  Parameter 3 (Trigger Mode) out of range.

BFFC0004  Parameter 4 (Trigger Slope) out of range.

BFFC0005  Parameter 5 (Bandwidth [Hz]) out of range.

BFFC0006  Parameter 6 (Pretrigger Samples) out of range.

BFFC0007  Parameter 7 (No of Samples) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTraceIQSrate

ViStatus rsfsu_confTraceIQSrate (ViSession instrumentHandle,

ViReal64 samplingRateHz);

Purpose

This function sets the sampling rate for the I/Q data acquisition. Thus

the sample rate can be modified without affecting the other settings.

Notes:

(1) Function Trace I/Q State (confTraceIQState) has to be switched ON

before using this function.

Remote-control command(s):

TRACe:IQ:SRATe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

samplingRateHz

Variable Type       ViReal64

This control sets the sampling rate for the I/Q data acquisition.

Valid Range: 10.0e3 Hz to 100.0e6 Hz

Value range with bandwidth extension (R&S FSQ-B72): 10.0e3 Hz to

326.4e6 Hz

Default Value: 32.0e6 Hz

Notes:

(1) Typical values are: 15.625 kHz, 31.25 kHz, 62.5 kHz, 125 kHz, 250

kHz, 500 kHz, 1 MHz, 2 MHz, 4 MHz, 8 MHz, 16 MHz, 32 MHz.

(2) At sampling rates > 81.6 MHz, a bandwidth of 50 MHz must always

be specified.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Sampling Rate [Hz]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTraceIQState

ViStatus rsfsu_confTraceIQState (ViSession instrumentHandle,

ViBoolean traceState);

Purpose

This function switches IQ data acquisition on or off.

Note:

The I/Q data acquisition is not compatible with other measurement

functions. Therefore all other measurement functions will be switched off

as soon as the I/Q measurement function is switched on. Additionally a

trace display is not possible in this operating mode. Therefore all

traces are set to "BLANK". Finally split screen operation will

automatically be stopped.

Remote-control command(s):

TRACe:IQ OFF | ON

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

traceState

Variable Type       ViBoolean

This control switches IQ data acquisition on or off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Trace State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrackExtFreqSweep

ViStatus rsfsu_confTrackExtFreqSweep (ViSession instrumentHandle,

ViInt32 window, ViInt32 generator,

ViBoolean frequencySweep,

ViReal64 frequencyOffset,

ViReal64 denominator,

ViReal64 numerator);

Purpose

This function is used for setting the generator multiplier and the offset

used to derive the generator frequency from the analyzer frequency.

Notes:

(1) If no external generator is selected, if the IECBUS address is not

correct or the generator is not ready for operation, an execution error

will be generated.

(2) This function is only valid in combination with option Ext. Generator

Control FSU-B10.

Remote-control command(s):

SOURce<1|2>:EXTernal<1|2>:FREQuency:NUMerator

SOURce<1|2>:EXTernal<1|2>:FREQuency:DENominator

SOURce<1|2>:EXTernal<1|2>:FREQuency:OFFSet

SOURce<1|2>:EXTernal<1|2>:FREQuency:SWEep ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

generator

Variable Type       ViInt32

This control selects the external generator 1 or 2.

Valid Range: 1 to 2

Default Value: 1

frequencySweep

Variable Type       ViBoolean

This control activates or deactivates the frequency sweep for

generator 1 or 2 in the selected window.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

frequencyOffset

Variable Type       ViReal64

This control defines the frequency offset of the selected generator 1

or 2 with reference to the receive frequency in the selected window.

Valid Range: none

Default Value: 0.0

Note:

Select the multiplication factor so that the frequency range of the

generator is not exceeded if the following formula:

F(generator) = F(analyzer) * (numerator/denominator) + F(offset)

is applied to the start and stop frequency of the analyzer.

This parameter is not range-checked.

denominator

Variable Type       ViReal64

This control defines the denominator of the factor with which the

analyzer frequency is multiplied in order to obtain the transmit

frequency of the selected generator 1 or 2 in the selected window.

Valid Range: none

Default Value: 1.0

Note:

Select the multiplication factor so that the frequency range of the

generator is not exceeded if the following formula:

F(generator) = F(analyzer) * (numerator/denominator) + F(offset)

is applied to the start and stop frequency of the analyzer.

This parameter is not range-checked.

numerator

Variable Type       ViReal64

This control defines the numerator of the factor with which the

analyzer frequency is multiplied to obtain the transmit frequency of

the selected generator 1 or 2 in the selected window.

Valid Range: none

Default Value: 1.0

Note:

Select the multiplication factor so that the frequency range of the

generator is not exceeded if the following formula:

F(generator) = F(analyzer) * (numerator/denominator) + F(offset)

is applied to the start and stop frequency of the analyzer.

This parameter is not range-checked.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Generator) out of range.

BFFC0004  Parameter 4 (Frequency Sweep) out of range.

BFFC0005  Parameter 5 (Frequency Offset) out of range.

BFFC0006  Parameter 6 (Denominator) out of range.

BFFC0007  Parameter 7 (Numerator) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrackExtPower

ViStatus rsfsu_confTrackExtPower (ViSession instrumentHandle,

ViInt32 window, ViInt32 generator,

ViReal64 powerLeveldBm);

Purpose

This function sets the output power of the selected generator 1 or 2 in

the selected window.

Note:

This function is only valid in combination with option Ext. Generator

Control FSU-B10.

Remote-control command(s):

SOURce<1|2>:EXTernal<1|2>:POWer

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

generator

Variable Type       ViInt32

This control selects the external generator 1 or 2.

Valid Range: 1 to 2

Default Value: 1

powerLeveldBm

Variable Type       ViReal64

This control sets the output power of the selected generator 1 or 2

in the selected window.

Valid Range: none

Default Value: -20.0 dBm

Note:

(1) This parameter is not range-checked.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Generator) out of range.

BFFC0004  Parameter 4 (Power Level [dBm]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrackExtSel

ViStatus rsfsu_confTrackExtSel (ViSession instrumentHandle,

ViInt32 generator,

ViInt32 generatorAddress,

ViInt32 generatorLink,

ViString generatorType);

Purpose

This function sets parameters for the external generator. When switching

on the external generator, the FSU switches off the internal tracking

generator and starts programming the generator settings via the IECBUS.

At the same time, the maximum stop frequency is limited to the maximum

generator frequency. This upper limit is automatically reduced by the set

frequency offset of the generator and a set multiplication factor. With

the external generator switched on, the FFT filters (FILTER TYPE FFT in

the menu BW) are not available.

Notes:

(1) If there is an error on the IECBUS during programming of the external

generator, the generator is automatically switched off and error is

generated.

(2) This function is only available with option Ext. Generator Control

FSU-B10.

Remote-control command(s):

SYSTem:COMMunicate:GPIB:RDEVice:GENerator<1|2>:ADDRess

SYSTem:COMMunicate:RDEVice:GENerator<1|2>:TYPE

SYSTem:COMMunicate:RDEVice:GENerator<1|2>:LINK GPIB | TTL

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

generator

Variable Type       ViInt32

This control selects the external generator 1 or 2.

Valid Range: 1 to 2

Default Value: 1

generatorAddress

Variable Type       ViInt32

This control changes the IEC/IEEE-bus address of the device selected

as external generator 1 or 2.

Valid Range: 0 to 30

Default Value: 28

Note:

If two generators are connected at the same time to IECBUS 2 their

addresses must be different.

generatorLink

Variable Type       ViInt32

This control selects the interface type of the external generator 1

or 2.

Valid Range:

0 - GPIB (Default Value)

1 - TTL

Notes:

(1) IECBUS alone (= GPIB, for all the generators of other

manufacturers and some Rohde & Schwarz units).

(2) IECBUS and TTL interface for synchronization (= TTL, for most of

the Rohde & Schwarz generators).

(3) The difference between the two operating modes is the execution

speed. While, during IECBUS operation, each settable frequency is

transmitted separately to the generator, a whole frequency list can

be programmed in one go if the TTL interface is also used. Frequency

switching can then be performed per TTL handshake which results in

considerable speed advantages.

(4) Only one of the two generators can be operated via the TTL

interface at a time. The other generator must be configured for

IECBUS (GPIB).

generatorType

Variable Type       ViString

This control selects the type of external generator 1 or 2. The

available generator types including the associated interface are

listed in the documentation.

Valid Values:

SME02, SME03, SME06, SMG, SMGL, SMGU, SMH, SMHU, SMIQ02B, SMIQ02E,

SMIQ03B, SMIQ03E, SMIQ04B, SMIQ06B, SML01, SMR20, SMR20B11, SMR27,

SMR27B11, SMR30, SMR30B11, SMR40, SMR40B11, SMP02, SMP03, SMP04,

SMP22, SMT02, SMT03, SMT06, SMU02, SMU02B31, SMU03, SMU03B31, SMU04,

SMU04B31, SMU06, SMU06B31.

Default Value: "SME03"

Notes:

(1) Generators with TTL interface can also be operated via IECBUS (=

GPIB) alone.

(2) With "NONE" selected, the corresponding generator 1 or 2 is

deactivated.

(3) There is a minimum version requirement of the R&S SMU firmware.

Please install V1.10 or higher on the R&S SMU!

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Generator) out of range.

BFFC0003  Parameter 3 (Generator Address) out of range.

BFFC0004  Parameter 4 (Generator Link) out of range.

BFFC0005  Parameter 5 (Generator Type) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrackExtSrc

ViStatus rsfsu_confTrackExtSrc (ViSession instrumentHandle,

ViInt32 window,

ViBoolean externalSource);

Purpose

This function activates or deactivates the external generator in the

selected window.

Notes:

(1) This function is only available in connection with option Ext.

Generator Control FSU-B10.

(2) If no external generator is selected, if the IECBUS address is not

correct or the generator is not ready for operation, an execution error

will be generated.

Remote-control command(s):

SOURce<1|2>:EXTernal ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

externalSource

Variable Type       ViBoolean

This command switches on or off external source.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (External Source) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrackExtSrcRef

ViStatus rsfsu_confTrackExtSrcRef (ViSession instrumentHandle,

ViInt32 window, ViInt32 reference);

Purpose

This function configures the reference signal for the external

generator.

Note:

(1) This command is only valid in conjunction with option Tracking

Generator, FSQ-B9 and Ext. Generator Control FSQ-B10.

Remote-control command(s):

SOUR:EXT1:ROSC INT | EXT

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

reference

Variable Type       ViInt32

This control sets the reference signal for the external

generator.

Valid Range:

0 - Internal (Default Value)

1 - External

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Reference) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrackModulation

ViStatus rsfsu_confTrackModulation (ViSession instrumentHandle,

ViInt32 window, ViInt32 modulation,

ViBoolean modState);

Purpose

This function switches on or off the external modulations of the tracking

generator.

Note:

This function is only valid in conjunction with option Tracking

Generator, FSU-B9.

Remote-control command(s):

SOURce<1|2>:AM:STATe ON | OFF

SOURce<1|2>:FM:STATe ON | OFF

SOURce<1|2>:DM:STATe ON | OFF

SOURce<1|2>:POWer:ALC:SOURce ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

modulation

Variable Type       ViInt32

This control switches on or off the external modulations of the

tracking generator.

Valid Range:

0 - Ext AM (Default Value)

1 - Ext FM

2 - Ext I/Q

3 - RESERVED

Notes:

(1) Ext AM:

This command switches on or off the external amplitude modulation of

the tracking generator. External ALC and external I/Q-modulation is

switched off, if active.

(2) Ext FM:

This command switches on or off the external frequency modulation of

the tracking generator. External AM, external I/Q-modulation and

frequency offset are switched off, if active.

(3) Ext I/Q:

This command switches on or off the external delta modulation of the

tracking generator. External AM, external ALC, external FM and

external frequency offset are switched off, if active. This is only

available with option FSE-B9 or FSE-B11.

modState

Variable Type       ViBoolean

This command switches on or off selected modulation.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Modulation) out of range.

BFFC0004  Parameter 4 (Mod State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrackRefLevel

ViStatus rsfsu_confTrackRefLevel (ViSession instrumentHandle,

ViInt32 window, ViReal64 refValue);

Purpose

This function configures the reference value position.

Note:

This command is only valid in conjunction with option Tracking Generator,

FSU-B9.

Remote-control command(s):

DISPlay:WINDow<1|2>:TRACe:Y:RVALue

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

refValue

Variable Type       ViReal64

This control defines the reference value for the y-axis of the

measurement diagram. This defines the upper limit of the display

range, the corresponding parameters of the manual control

vary according to the measurement mode.

Valid Range: depends on the unit.

This value is not range-checked.

Default Value: 0.0 dB

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Ref Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrackRefPosition

ViStatus rsfsu_confTrackRefPosition (ViSession instrumentHandle,

ViInt32 window,

ViReal64 refValuePosition);

Purpose

This function configures the reference value position.

Note:

This command is only valid in conjunction with option Tracking Generator,

FSU-B9.

Remote-control command(s):

DISPlay:WINDow<1|2>:TRACe:Y:RPOSition

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

refValuePosition

Variable Type       ViReal64

This function defines the position of the reference value.

Valid Range: -200.0 % to 200.0 %

Default Value: 100.0 %

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Ref Value Position [%]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrackSrc

ViStatus rsfsu_confTrackSrc (ViSession instrumentHandle, ViInt32 window,

ViInt32 setSourceValue,

ViReal64 sourceValue);

Purpose

This function configures tracking generator source values.

Note:

This function is only valid in conjunction with option Tracking

Generator, FSU-B9.

Remote-control command(s):

SOURce<1|2>:POWer:OFFSet

SOURce<1|2>:POWer

SOURce<1|2>:FREQuency:OFFSet

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

setSourceValue

Variable Type       ViInt32

This control sets desired tracking generator source value.

Valid Range:

0 - Power Offset

1 - Power Amplitude (Default Value)

2 - Frequency Offset

Notes:

(1) Power Offset:

This command defines a level offset for the tracking generator.

(2) Power Amplitude:

This command defines the level of the tracking generator.

(3) Frequency Offset:

This command defines a frequency offset of the tracking generators to

the current analyzer frequency.

sourceValue

Variable Type       ViReal64

This is source value control.

Valid Ranges:

- Power Offset: -200.0 dB to +200.0 dB

- Power Amplitude: -30.0 dBm to 0.0 dBm

- Frequency Offset: -150.0e6 Hz to 150.0e6 Hz

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Set Source Value) out of range.

BFFC0004  Parameter 4 (Source Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTransducer

ViStatus rsfsu_confTransducer (ViSession instrumentHandle,

ViInt32 window,

ViInt32 transducerOperation,

ViString transducerName,

ViBoolean transducerState);

Purpose

This function (de)activates a transducer factor or deletes the selected

factor.

Remote-control command(s):

SENSe<1|2>:CORRection:TRANsducer:SELect

SENSe<1|2>:CORRection:TRANsducer:STATe ON | OFF

SENSe<1|2>:CORRection:TRANsducer:DELete

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

transducerOperation

Variable Type       ViInt32

This control selects the transducer factor or delete it.

Valid Range:

0 - Select Transducer Factor (Default Value)

1 - Delete Transducer Factor

transducerName

Variable Type       ViString

This command selects the transducer factor designated by <name>. If

<name> does not exist yet, a new transducer factor or set is created.

Valid Range: up to 8 characters

Default Value: ""

transducerState

Variable Type       ViBoolean

This command switches the selected transducer factor on or off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Transducer Operation) out of range.

BFFC0004  Parameter 4 (Transducer Name) out of range.

BFFC0005  Parameter 5 (Transducer State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTransducerDef

ViStatus rsfsu_confTransducerDef (ViSession instrumentHandle,

ViInt32 window,

ViString transducerName,

ViInt32 transducerScaling,

ViInt32 transducerFactorUnit,

ViInt32 no_ofTestPoints,

ViReal64 _VI_FAR frequencyData[],

ViReal64 _VI_FAR levelData[]);

Purpose

This function creates a new transducer factor. The transducer is stored

on harddisk any may be activated by function rsfsu_setupTransducer.

Remote-control command(s):

CORRection:TRANsducer:SELect

CORRection:TRANsducer:SCALing LINear | LOGarithmic

CORRection:TRANsducer:DATA

CORRection:TRANsducer:UNIT DB | DBM | DBMV | DBUV | DBUV/M | DBUA |

DBUA/M | DBPW | DBPT

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

(2) This control is not used.

transducerName

Variable Type       ViString

This command selects the transducer factor designated by <name>. If

<name> does not exist yet, a new transducer factor is created.

Valid Range: up to 8 characters

Default Value: ""

transducerScaling

Variable Type       ViInt32

This control defines whether the frequency scaling of the transducer

factor is linear or logarithmic.

Valid Values:

0 - Logarithmic

1 - Linear (Default Value)

transducerFactorUnit

Variable Type       ViInt32

This control defines the unit of the transducer factor selected.

Valid Ranges:

0 - DB (Default Value)

1 - DBM

2 - DBMV

3 - DB�V

4 - DB�V/M

5 - DB�A

6 - DB�A/M

7 - DBPW

8 - DBPT

no_ofTestPoints

Variable Type       ViInt32

Specifies the number of frequency,level pairs. This value corresponds

with number of arrays members (Frequency and Level controls).

Valid Range:

0 to max memory

Default Value:

0

frequencyData

Variable Type       ViReal64[]

This array defines the frequency part (in Hz) of the test points for

the selected transducer factor.  The frequencies must be in ascending

order.

Valid Range: not range checked

Default Value: none

levelData

Variable Type       ViReal64[]

This array defines the level part of the test points for the selected

transducer factor.

The level values are sent as dimensionless numbers; the unit is

specified by control Unit.

Valid Range: not range checked

Default Value: none

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Transducer Name) out of range.

BFFC0004  Parameter 4 (Transducer Scaling) out of range.

BFFC0005  Parameter 5 (Transducer Factor Unit) out of range.

BFFC0006  Parameter 6 (No of Test Points) out of range.

BFFC0007  Parameter 7 (Frequency Data) out of range.

BFFC0008  Parameter 8 (Level Data) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTransducerFactor

ViStatus rsfsu_confTransducerFactor (ViSession instrumentHandle,

ViInt32 window,

ViString transducerName);

Purpose

This function generates a transducer factor <name> using normalized trace

data.

Note:

(1) The function is only available when normalization is switched on.

(2) The associated measurement window is selected with the function Set

Active Window (rsfsu_confSetActiveWindow).

Remote-control command(s):

[SENSe<1|2>:]CORRection:TRANsducer:GENerate

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

transducerName

Variable Type       ViString

This command selects the transducer factor name.

Valid Range: up to 8 characters

Default Value: ""

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Transducer Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTransducerRefLevAdj

ViStatus rsfsu_confTransducerRefLevAdj (ViSession instrumentHandle,

ViInt32 window,

ViBoolean refLevAdj);

Purpose

This function allows automatic reference level offset adaption to restore

the original dynamic range by also shifting the reference level by the

maximum value of the transducer factor.

Note:

Command SENS:CORR:TRAN:SEL has to be sent prior to this function.

Remote-control command(s):

[SENSe<1|2>:]CORRection:TRANsducer:ADJust:RLEVel[:STATe] ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

refLevAdj

Variable Type       ViBoolean

Allows automatic reference level offset adaption to restore the

original dynamic range by also shifting the reference level by the

maximum value of the transducer factor.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Ref Lev Adj) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrg

ViStatus rsfsu_confTrg (ViSession instrumentHandle, ViInt32 window,

ViInt32 trigger, ViReal64 triggerValue,

ViBoolean triggerSlope);

Purpose

This function configures the trigger source and trigger slope.

Remote-control command(s):

TRIGger<1|2>:SEQuence:LEVel:VIDeo

TRIGger<1|2>:SEQuence:LEVel:FM

TRIGger<1|2>:SEQuence:LEVel:IFPower

TRIGger<1|2>:SEQuence:LEVel:AM

TRIGger<1|2>:SEQuence:LEVel:RFPower

TRIGger<1|2>:SEQuence:SLOPe POSitive | NEGative

TRIGger<1|2>:SEQuence:SOURce IMMediate | LINE | VIDeo | EXTernal |

RFPower | AF | IFPower

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

trigger

Variable Type       ViInt32

This control selects the trigger source for the start of a sweep.

Valid Range:

0 - Free Run (Default Value)

1 - RESERVED

2 - Video

3 - Video Trg Value

4 - External

5 - RESERVED

6 - RF Power

7 - AF Signal

8 - AF Signal Value

9 - IF Power

10 - IF Power Value

11 - TV

12 - AM

13 - AM Value

14 - RF Power Value

Notes:

(1) Free Run:

Automatic triggering the next measurement at the end of the

previous one.

(2) Video:

The next measurement is triggered by the detection of a signal at the

video filter output.

(3) Video Trg Value:

Sets trigger level value.

(4) External:

The next measurement is triggered by the signal at the external

trigger input.

(5) RF Power:

The next measurement is triggered by the detect�on of a signal at the

instrument RF (80 MHz bandwidth). RFPower cannot be selected in FM

DEMOD mode. This selection is valid only with option FS-B6 (TV and

RF-Trigger) installed.

(6) RF Power Value:

Sets the RF Power trigger value.

(6) AF Signal:

The next measurement is triggered on detection of an FM. This

selection is valid only with option FS-K7 (FM Demodulator) installed.

(7) AF Signal Value:

Sets trigger level value, when FM modulated signals are used as

trigger source. This selection is valid only with option FS-K7 (FM

Demodulator) installed.

(8) IF Power:

The next measurement is triggered by the detect�on of a signal at the

instrument IF (10 MHz bandwidth).

(9) IF Power Value:

Sets the IF Power Level value.

(10) TV:

The next measurement is triggered by the detect�on of a TV signal

according to the settings of the TRIGger:SEQuence:VIDeo-

subsystem. TV cannot be selected in FM DEMOD mode.

(11) AM:

The next measurement is triggered on detection of an            audio

signal after AM demodulation. This selection is valid only with

option FS-K7 (FM Demodulator) installed.

(12) AM Value:

Sets the AM trigger value. This selection is valid only with option

FS-K7 (FM Demodulator) installed.

triggerValue

Variable Type       ViReal64

This command sets the level of the video trigger source.

Valid Range:

0.0 % to 100.0 %   (Video Trigger Level)

-70 dB to 30.0 dB  (IF Power Level)

Default Value: 0.0

triggerSlope

Variable Type       ViBoolean

This control selects the slope of the trigger signal only when

trigger source is selected.

Valid Range:

VI_FALSE (0) - Negative

VI_TRUE  (1) - Positive (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Trigger) out of range.

BFFC0004  Parameter 4 (Trigger Value) out of range.

BFFC0005  Parameter 5 (Trigger Slope) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrgDelay

ViStatus rsfsu_confTrgDelay (ViSession instrumentHandle, ViInt32 window,

ViReal64 triggerValues);

Purpose

This function defines the length of the trigger delay.

Remote-control command(s):

TRIGger<1|2>:SEQuence:HOLD

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

triggerValues

Variable Type       ViReal64

This control defines the length of the trigger delay.

Valid Range: -100.0 to 100.0 s

Default Value: 0.0 s

Notes:

(1) A negative delay time (pretrigger) can be set in the time domain

(SPAN < 0 Hz) only. The maximum permissible range and the maximum

effective resolution of the pretrigger are limited by the set sweep

time (max range= - 499/500 x sweep time; max. resolution = sweep

time/500). Pretriggering is not possible when the rms or the average

detector is activated.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Trigger Value [s]) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confTrgDelayComp

ViStatus rsfsu_confTrgDelayComp (ViSession instrumentHandle,

ViInt32 window,

ViBoolean delayCompensation);

Purpose

This function enables the filter group delay compensation for the

external trigger. If a bursted signal is analyzed in zero span and the

delay compensation is on, a change of the RBW will not change the rising

slope of the signal.

Note:

The GSM/EDGE application FS-K5 and the 1xEV-DO BTS application FSK84 are

implicitly using delay compensation active since their first version.

Remote-control command(s):

TRIGger<1|2>:HOLD:ADJ:AUTO ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

delayCompensation

Variable Type       ViBoolean

Enables the filter group delay compensation for the external trigger.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Delay Compensation) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMADemod

ViStatus rsfsu_confVSAMADemod (ViSession instrumentHandle,

ViInt32 window, ViBoolean state);

Purpose

This function controls analog demodulation state.

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

SENSe1:ADEMod:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) Only Window A can be used now.

state

Variable Type       ViBoolean

Activates the FM demodulator of the instrument. The instrument will

be set to time-domain measurement (span = 0) at the current center

frequency. The detector will be set to  SAMPle, the demodulator

itself will be set up according to the parameters of command Analog

Demodulation Set.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMADemodMeasTime

ViStatus rsfsu_confVSAMADemodMeasTime (ViSession instrumentHandle,

ViInt32 window,

ViReal64 measuringTime);

Purpose

This function sets analog demodulation measurement time.

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

SENSe1:ADEMod:MTIMe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) Only Window A can be used now.

measuringTime

Variable Type       ViReal64

This control sets the measuring time for the analog demodulation.

Valid Range:

Default Value:

62.625E-6 s

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Measuring Time) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMADemodParam

ViStatus rsfsu_confVSAMADemodParam (ViSession instrumentHandle,

ViInt32 window,

ViInt32 demodulationParameter,

ViInt32 AFCoupling,

ViInt32 sideBand,

ViBoolean realTime);

Purpose

This function controls AF coupling of the FM demodulation function.

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

SENSe1:ADEMod:AF:COUPling DC | AC

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) Only Window A can be used now.

demodulationParameter

Variable Type       ViInt32

This control selects the measured data that are to be controlled.

Valid Range:

0 - Realtime

1 - Side Band

2 - AF Coupling (Default Value)

Note:

(1) Only the particular parameter specified by the Demodulation

Parameter is set.

(2) Only the AF Coupling is implemented now. Other parameters are for

future use.

AFCoupling

Variable Type       ViInt32

This control selects selects the coupling of the AF path.

Valid Range:

0 - DC

1 - AC (Default Value)

sideBand

Variable Type       ViInt32

This control selects the side band for the demodulation.

Valid Range:

0 - Normal (Default Value)

1 - Inverse

Note:

(1) This parameter is not used now.

realTime

Variable Type       ViBoolean

This control selects whether the demodulation performed in real time

or in blocks.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Note:

(1) This parameter is not used now.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Demodulation Parameter) out of range.

BFFC0004  Parameter 4 (AF Coupling) out of range.

BFFC0005  Parameter 5 (Side Band) out of range.

BFFC0006  Parameter 6 (Real Time) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMADemodPhaseWrap

ViStatus rsfsu_confVSAMADemodPhaseWrap (ViSession instrumentHandle,

ViInt32 window,

ViBoolean phaseWrap);

Purpose

This function activates the phase wrapping from -PI to +PI (-180 to

+180).

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

CALCulate<1|2>:FORMat PHASe | UPHase

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

phaseWrap

Variable Type       ViBoolean

Activates / Deactivates the phase wrapping to -PI to +PI (-180 to

+180).

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Phase Wrap) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMADemodPMUnits

ViStatus rsfsu_confVSAMADemodPMUnits (ViSession instrumentHandle,

ViInt32 window, ViInt32 PMUnits);

Purpose

This function selects the unit for PM traces.

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

UNIT<1|2>:ANGL RAD | DEG

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

PMUnits

Variable Type       ViInt32

Selects the unit for PM traces.

Valid Range:

0 - RAD

1 - DEG

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (PM Units) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMADemodRFParam

ViStatus rsfsu_confVSAMADemodRFParam (ViSession instrumentHandle,

ViInt32 window,

ViInt32 frequencyParameter,

ViReal64 bandwidthResolution,

ViInt32 span);

Purpose

This function configures bandwidth resolution and span for display of the

RF spectrum.

Remote-control command(s):

SENSe1:ADEMod:SPECtrum:BANDwidth:RESolution

SENSe1:ADEMod:SPECtrum:SPAN

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

frequencyParameter

Variable Type       ViInt32

Specifies which parameter will be set.

Valid Range:

0 - BW Resolution And Center (Default Value)

1 - BW Resolution Only

2 - Span Only

bandwidthResolution

Variable Type       ViReal64

Sets the resolution bandwidth for display of the RF spectrum from the

data of analog demodulation.

Valid Range:

1 to 10E+6 Hz

Default Value: 61.2E+3 Hz

span

Variable Type       ViInt32

This control sets the frequency range for display of the RF spectrum

resulting from the data of FM demodulation.

Valid Range:

0 - 10 MHz

1 - 8 MHz

2 - 5 MHz   (Default Value)

3 - 3 MHz

4 - 1.6 MHz

5 - 800 kHz

6 - 400 kHz

7 - 200 kHz

8 - 100 kHz

9 - 50 kHz

10 - 25 kHz

11 - 12.5 kHz

12 - 6.4 kHz

13 - 3.2 kHz

14 - 1.6 kHz

15 - 800 Hz

16 - 400 Hz

17 - 200 Hz

18 - 100 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Frequency Parameter) out of range.

BFFC0004  Parameter 4 (Bandwidth Resolution) out of range.

BFFC0005  Parameter 5 (Span) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMADemodSet

ViStatus rsfsu_confVSAMADemodSet (ViSession instrumentHandle,

ViInt32 window, ViInt32 sampleRate,

ViInt32 recordLength,

ViInt32 triggerSource,

ViInt32 triggerSlope,

ViInt32 pretriggerSamples,

ViInt32 numberOfMeas);

Purpose

This function configures the FM demodulator.

Note:

(1) This command is only available with option FS-K7 (FM Demodulator).

Remote-control command(s):

SENSe1:ADEMod:SET

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) Only Window A can be used now.

sampleRate

Variable Type       ViInt32

Selects the sampling rate for the data acquisition.

Valid Ranges:

0 - 15.625 kHz (Default Value)

1 -  31.25 kHz

2 -  62.5 kHz

3 -  125 kHz

4 -  250 kHz

5 -  500 kHz

6 -  1 MHz

7 -  2 MHz

8 -  4 MHz

9 -  8 MHz

10 -  16 MHz

11 -  32 MHz

recordLength

Variable Type       ViInt32

Number of samples to be stored in I/Q memory.

Valid Ranges:

1 to 130560

Default Value:

1

triggerSource

Variable Type       ViInt32

Selects the trigger source used for the measurement.

Valid Ranges:

0 - Immediate (Default Value)

1 - External

triggerSlope

Variable Type       ViInt32

Selecs the trigger slope.

Valid Ranges:

0 - Negative

1 - Positive (Default Value)

pretriggerSamples

Variable Type       ViInt32

Number of samples to be used as an offset to the trigger signal.

Valid Ranges:

-65024  to 130560

Default Value:

0

Note:

(1) This value will be ingnored for trigger = 'Immediate'.

numberOfMeas

Variable Type       ViInt32

Number of measurements. The value indicated here is particularly

necessary for the AVERage/MAXHold/MINHold function.

Valid Ranges:

0 to 32767

Default Value:

0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Sample Rate) out of range.

BFFC0004  Parameter 4 (Record Length) out of range.

BFFC0005  Parameter 5 (Trigger Source) out of range.

BFFC0006  Parameter 6 (Trigger Slope) out of range.

BFFC0007  Parameter 7 (Pretrigger Samples) out of range.

BFFC0008  Parameter 8 (Number Of Meas) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMADemodType

ViStatus rsfsu_confVSAMADemodType (ViSession instrumentHandle,

ViInt32 window,

ViInt32 demodulationType,

ViInt32 resultType1,

ViInt32 resultType2,

ViInt32 resultType3);

Purpose

This function sets the result types to be measured in parallel with the

selected demodulation.

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

[SENSe1:]ADEMod:AM | FM | PM | SPECtrum[:TYPE] <WRITe | AVERage | MAXHold

| MINHold | VIEW | OFF>,<WRITe | AVERage | MAXHold | MINHold | VIEW |

OFF>,<WRITe | AVERage | MAXHold | MINHold | VIEW | OFF>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

demodulationType

Variable Type       ViInt32

This parameter selects the demodulation.

Valid Range:

0 - AM

1 - FM

2 - RF Spectrum

3 - PM

Notes:

(1) AM: Selects the result types to be measured in parallel by AM

demodulation.

(2) FM: Selects the result types to be created in parallel by FM

demodulation.

(3) RF Spectrum: Selects the result types to be measured

simultaneously with display of the RF

(2) PM: Selects the result types to be created in parallel by PM

demodulation.

resultType1

Variable Type       ViInt32

Selects the result type 1, to be measured in parallel by selected

demodulation.

Valid Range:

0 - Current Meas Value

1 - Average

2 - Maximum

3 - Minimum

4 - OFF (Default Value)

5 - View

Notes:

(1) Current Meas Value:

The current measured values are recorded.

(2) Average:

The measurement results are averaged over the given number of

measurements.

(3) Maximum:

The maximum result over the given number of measurements is obtained.

(4) Minimum:

The minimum result over the given number of measurements is obtained.

(5) OFF:

The result type is not used.

(6) View:

The measurement result is frozen on the display, i.e. it is not

recalculated in further measurements.

resultType2

Variable Type       ViInt32

Selects the result type 2, to be measured in parallel by selected

demodulation.

Valid Range:

0 - Current Meas Value

1 - Average

2 - Maximum

3 - Minimum

4 - OFF (Default Value)

5 - View

Notes:

(1) Current Meas Value:

The current measured values are recorded.

(2) Average:

The measurement results are averaged over the given number of

measurements.

(3) Maximum:

The maximum result over the given number of measurements is obtained.

(4) Minimum:

The minimum result over the given number of measurements is obtained.

(5) OFF:

The result type is not used.

(6) View:

The measurement result is frozen on the display, i.e. it is not

recalculated in further measurements.

resultType3

Variable Type       ViInt32

Selects the result type 2, to be measured in parallel by selected

demodulation.

Valid Range:

0 - Current Meas Value

1 - Average

2 - Maximum

3 - Minimum

4 - OFF (Default Value)

5 - View

Notes:

(1) Current Meas Value:

The current measured values are recorded.

(2) Average:

The measurement results are averaged over the given number of

measurements.

(3) Maximum:

The maximum result over the given number of measurements is obtained.

(4) Minimum:

The minimum result over the given number of measurements is obtained.

(5) OFF:

The result type is not used.

(6) View:

The measurement result is frozen on the display, i.e. it is not

recalculated in further measurements.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Demodulation Type) out of range.

BFFC0004  Parameter 4 (Result Type 1) out of range.

BFFC0005  Parameter 5 (Result Type 2) out of range.

BFFC0006  Parameter 6 (Result Type 3) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMADemodZeroPhase

ViStatus rsfsu_confVSAMADemodZeroPhase (ViSession instrumentHandle,

ViInt32 window,

ViReal64 referencePoint);

Purpose

This function activates the input of the reference point, where the phase

is set to 0 rad.

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

[SENSe1:]ADEMod:PM:RPO:X

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

referencePoint

Variable Type       ViReal64

The reference point, where the phase is set to 0 rad.

Valid Range: none

Default Vaule: 10.0e-6 s

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Reference Point) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDemodBW

ViStatus rsfsu_confVSAMDemodBW (ViSession instrumentHandle,

ViInt32 window,

ViInt32 demodulationBandwidth);

Purpose

This function sets demodulation bandwidth of FM demodulation.

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

SENSe1:ADEMod:BWIDth:DEMod

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

This control determines the display screen that will be affected.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Note:

If split screen mode has not been configured, '2' will produce an

execution error.

demodulationBandwidth

Variable Type       ViInt32

This control sets the bandwidth for analog demodulation. Depending on

the selected demodulation bandwidth, the instrument selects the

required sampling rate. The available values of the demodulation

bandwidths are determined by the sampling rates.

Valid Range:

0 - 10 MHz

1 - 8 MHz

2 - 5 MHz   (Default Value)

3 - 3 MHz

4 - 1.6 MHz

5 - 800 kHz

6 - 400 kHz

7 - 200 kHz

8 - 100 kHz

9 - 50 kHz

10 - 25 kHz

11 - 12.5 kHz

12 - 6.4 kHz

13 - 3.2 kHz

14 - 1.6 kHz

15 - 800 Hz

16 - 400 Hz

17 - 200 Hz

18 - 100 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Demodulation Bandwidth) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDemodFilt

ViStatus rsfsu_confVSAMDemodFilt (ViSession instrumentHandle,

ViInt32 filterType,

ViBoolean filterState,

ViInt32 filterFrequency,

ViInt32 AMFMDeemphasisValue);

Purpose

This function selects AF filters. The bandwidth of the demodulated signal

can be reduced by high pass or low pass filters and also a de-emphasis

can be switched on.

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

[SENSe:]FILTer:HPASs[:STATe] ON | OFF

[SENSe:]FILTer:HPASs:FREQuency 50 Hz | 300 Hz

[SENSe:]FILTer:LPASs[:STATe] ON | OFF

[SENSe:]FILTer:LPASs:FREQuency 3 kHz | 15 kHz | 150 kHz

[SENSe:]FILTer:DEMPhasis[:STATe] ON | OFF

[SENSe:]FILTer:DEMPhasis:TCONstant 25 us | 50 us | 75 us | 750 us

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

filterType

Variable Type       ViInt32

This control selects the filter type.

Valid Range:

0 - High Pass AF Filter (Default Value)

1 - Low Pass AF Filter

2 - AM/FM Deemphasis

Notes:

(1) High Pass AF Filter:

Activates the high-pass filter in the AF-branch for analog

demodulation.

(2) Low Pass AF Filter:

Activates the low-pass filter in the AF-branch with analog

demodulation.

(3) AM/FM Deemphasis:

Activates the selected de-emphasis for analog demodulation.

filterState

Variable Type       ViBoolean

This control activates the selected filter.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

filterFrequency

Variable Type       ViInt32

This control sets the filter frequency.

Valid Values:

for High Pass filter:  50 Hz,   300 Hz

for Low Pass filter: 3000 Hz, 15000 Hz, 150000 Hz

Default Value: 50 Hz

Note(s):

(1) High Pass:

50 Hz for 200 Hz <= demodulation bandwidth <=  4 MHz

300 Hz     800 Hz <= demodulation bandwidth <= 16 MHz

(2) Low Pass:

3000 Hz for 6.4 kHz <= demodulation bandwidth <=  4 MHz

15000 Hz      50 kHz <= demodulation bandwidth <= 16 MHz

150000 Hz     400 kHz <= demodulation bandwidth <= 16 MHz

AMFMDeemphasisValue

Variable Type       ViInt32

This control sets the time constant of the de-emphasis for analog

demodulation.

Valid Range:

0 - 25 us (Default Value)

1 - 50 us

2 - 75 us

3 - 750 us

Note(s):

(1) The de-emphasis is active in the following demodulation bandwidth

range:

25 us for  25 kHz <= demodulation bandwidth <= 30 MHz

50 us     6.4 kHz <= demodulation bandwidth <= 18 MHz

75 us     6.4 kHz <= demodulation bandwidth <= 18 MHz

750 us     800  Hz <= demodulation bandwidth <=  4 MHz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Filter Type) out of range.

BFFC0003  Parameter 3 (Filter State) out of range.

BFFC0004  Parameter 4 (Filter Frequency) out of range.

BFFC0005  Parameter 5 (AM/FM Deemphasis Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDemodMode

ViStatus rsfsu_confVSAMDemodMode (ViSession instrumentHandle,

ViInt32 window,

ViInt32 operatingMode);

Purpose

This function selects demodulation mode of the signal analysis.

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

CALCulate<1|2>FEED 'XTIM:AM'

CALCulate<1|2>FEED 'XTIM:FM'

CALCulate<1|2>FEED 'XTIM:SPEC'

CALCulate<1|2>FEED 'XTIM:AMS<1|2|3>'

CALCulate<1|2>FEED 'XTIM:FMS<1|2|3>'

CALCulate<1|2>FEED 'XPOW:CDP'

CALCulate<1|2>FEED 'XPOW:CDP:RAT'

CALCulate<1|2>FEED 'XPOW:CDP:OVER'

CALCulate<1|2>FEED 'XTIM:CDP:ERR

CALCulate<1|2>FEED 'XTIM:CDP:ERR:PHAS'

CALCulate<1|2>FEED 'XTIM:CDP:ERR:SUMM'

CALCulate<1|2>FEED 'XTIM:CDP:ERR:CTAB'

CALCulate<1|2>FEED 'XTIM:CDP:ERR:PCD'

CALCulate<1|2>FEED 'XTIM:CDP:MACC'

CALCulate<1|2>FEED 'XTIM:CDP:PVSL'

CALCulate<1|2>FEED 'XTIM:CDP:PVS'

CALCulate<1|2>FEED 'XTIM:CDP:BSTR'

CALCulate<1|2>FEED 'XTIM:CDP:SYMB:CONS'

CALCulate<1|2>FEED 'XTIM:CDP:SYMB:EVM'

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

This control determines the display screen that will be affected.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Note:

If split screen mode has not been configured, '2' will produce an

execution error.

operatingMode

Variable Type       ViInt32

This control selects the operating mode in vector signal analysis.

Valid Range:

0 - Demodulated AM signal in level display

1 - Demodulated FM signal

2 - RF spectrum of FM signal

3 - AM-Summary, trace 1

4 - AM-Summary, trace 2

5 - AM-Summary, trace 3

6 - FM-Summary, trace 1

7 - FM-Summary, trace 2

8 - FM-Summary, trace 3

9 - CDP Power (abs)

10 - CDP Power (rel)

11 - CDP Overview

12 - CDP timing error

13 - CDP phase error

14 - CDP errors in tabular form

15 - CDP error in channel table

16 - CDP peak code domain error

17 - CDP modulation accuracy

18 - CDP power versus slot

19 - CDP power versus symbol

20 - CDP bitstream

21 - CDP symbol constellation

22 - CDP vector magnitude error

Notes:

(0) Demodulated AM signal in level display.

(1) Demodulated FM signal.

(2) RF spectrum of FM signal determined from the measured data by

means of FFT.

(3) AM-Summary, trace 1 (AM summary marker, referred to trace 1)

(4) AM-Summary, trace 2 (AM summary marker, referred to trace 2)

(5) AM-Summary, trace 3 (AM summary marker, referred to trace 3)

(6) FM-Summary, trace 1 (FM summary marker, referred to trace 1)

(7) FM-Summary, trace 2 (FM summary marker, referred to trace 2)

(8) FM-Summary, trace 3 (FM summary marker, referred to trace 3)

(9) The selections 0 - 8 are available only with option FS-K7 (FM

Demodulator).

(10) CDP Power (abs) (Bargraph result display of code domain power

(absolute) - only window 1. The window setting has no effect).

(11) CDP Power (rel) (Bargraph result display of code domain power

ratio  (relative), only window 1. The window setting has no effect).

(12) CDP Overview (Overview, screen A displays CDP Rel I,   screen B

CDP Q (option FS-K73 only)).

(13) CDP timing error (Bargraph result display of timing error)

(14) CDP phase error (Bargraph result display of  phase error)

(15) CDP errors in tabular form (Display of results in tabular form,

only window 2. The window setting has no effect).

(16) CDP error in channel table (Display of channel table, only

window 1. The window setting has no effect).

(17) CDP peak code domain error (Result display of peak code domain

error, only window 2. The window setting has no effect).

(18) CDP modulation accuracy (Result display of modulation accuracy,

only window 2. The window setting has no effect).

(19) CDP power versus slot  (Result display of power versus slot,

only window 2. The window setting has no effect).

(20) CDP power versus symbol (Result display of power versus symbol,

only window 2. The window setting has no effect).

(21) CDP bitstream (Result display of bitstream, only window 2. The

window setting has no effect).

(22) CDP symbol constellation (Result display of symbol

constellation, only window 2. The window setting has no effect).

(23) CDP vector magnitude error (Result display of vector magnitude

error, only window 2. The window setting has no effect).

(24) The selections 9 - 22 are available only with option FS-K72/K73

(WCDMA 3GPP FDD BTS/MS).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Operating Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDemodZoom

ViStatus rsfsu_confVSAMDemodZoom (ViSession instrumentHandle,

ViInt32 window,

ViInt32 demodulationParameter,

ViBoolean zoomState,

ViReal64 startTime);

Purpose

This function sets analog demodulation parameters such as zoom state,

start time.

Note:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

SENSe1:ADEMod:ZOOM:STATe ON | OFF

SENSe1:ADEMod:ZOOM:STARt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

This control determines the display screen that will be affected.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Note:

If split screen mode has not been configured, '2' will produce an

execution error.

demodulationParameter

Variable Type       ViInt32

This parameter selects functions for zoom setting.

Valid Range:

0 - Zoom State (Default Value)

1 - Start Time

Notes:

(1) Only the specified parameter is configured at the function call.

zoomState

Variable Type       ViBoolean

Turns on or off the zoom function for the trace data of the FM

demodulator.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Notes:

(1) This control is active only if Demodulation Parameter is set to

Zoom State.

startTime

Variable Type       ViReal64

Selects the start time for the display of the single measured values

of the FM demodulator.

Valid Range:

Start Time: 0.0 to Max s

Notes:

(1) Max depends on the measuring time set in the instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Demodulation Parameter) out of range.

BFFC0004  Parameter 4 (Zoom State) out of range.

BFFC0005  Parameter 5 (Start Time) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDigiDemodFilter

ViStatus rsfsu_confVSAMDigiDemodFilter (ViSession instrumentHandle,

ViInt32 window,

ViInt32 measurementFilter);

Purpose

This function configures the digital demodulation signal filter

Remote-control command(s):

SENSe<1|2>:DDEMod:FILTer:MEASurement OFF | BTOoth

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

measurementFilter

Variable Type       ViInt32

This control selects the input filter for the test signal.

Valid Range:

0 - Off (Default Value)

1 - Raised Cosine

2 - Root Raised Cosine

3 - Gaussian

4 - BlueTooth

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Measurement Filter) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDigiSearchTime

ViStatus rsfsu_confVSAMDigiSearchTime (ViSession instrumentHandle,

ViInt32 window,

ViBoolean autoTime,

ViReal64 value);

Purpose

This function sets the search time. In this time the sync. word or pulse

should be found.

Note:

(1) This function is available only in conjunction with option FS-K8.

Remote-control command(s):

SENSe<1|2>:DDEMod:SEARch:TIME:AUTO ON | OFF

SENSe<1|2>:DDEMod:SEARch:TIME

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

autoTime

Variable Type       ViBoolean

Sets the mode for search time.

Valid Values:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

value

Variable Type       ViReal64

Sets the time value.

Valid Range:

100.0E-6 s to 130560.0E-6 s

Default Value:

1875.0E-6 s

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Auto Time) out of range.

BFFC0004  Parameter 4 (Value) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDigiSeqOffset

ViStatus rsfsu_confVSAMDigiSeqOffset (ViSession instrumentHandle,

ViInt32 window,

ViInt32 syncOffset);

Purpose

This function defines the offset of the display with reference to the

synchronization sequence.

Remote-control command(s):

SENSe<1|2>:DDEMod:SEARch:SYNC:OFFSet

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

syncOffset

Variable Type       ViInt32

This control defines the offset of the display with reference to the

synchronization sequence.

Valid Range: -10000 to 10000

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Sync Offset) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDigiSeqPulseOffset

ViStatus rsfsu_confVSAMDigiSeqPulseOffset (ViSession instrumentHandle,

ViInt32 window,

ViReal64 syncOffset);

Purpose

This function defines the pulse offset.

Remote-control command(s):

SENSe<1|2>:DDEMod:SEARch:PULSe:OFFSet

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

syncOffset

Variable Type       ViReal64

This control defines the pulse offset (time).

Valid Range: -10.0E-3 to 10.0E-3

Default Value: 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Sync Offset) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDigiSetSyncLap

ViStatus rsfsu_confVSAMDigiSetSyncLap (ViSession instrumentHandle,

ViInt32 window, ViInt32 LAP);

Purpose

This function sets the lowest 24 bits of Bluetooth Device Address. These

bits are used to determine the begin of the signal packet.

Note:

(1) This function is available only with Bluetooth Analyzer  option

FS-K8.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

LAP

Variable Type       ViInt32

Sets the lowest 24 bits (LAP).

Valid Range:

0 to 16777215

Default Value:

0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (LAP) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDigiTrg

ViStatus rsfsu_confVSAMDigiTrg (ViSession instrumentHandle,

ViInt32 window, ViBoolean findBurst,

ViBoolean findSync);

Purpose

This function controls digital trigger parameters (switches the signal

burst search and search for a synchronization sequence on or off) in

signal analysis operation.

Remote-control command(s):

SENSe<1|2>:DDEMod:SEARch:PULSe:STATe ON | OFF

SENSe<1|2>:DDEMod:SEARch:SYNC:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

findBurst

Variable Type       ViBoolean

This control switches the signal burst search on or off.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

findSync

Variable Type       ViBoolean

This control switches the search for a synchronization sequence on or

off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_FALSE (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Find Burst) out of range.

BFFC0004  Parameter 4 (Find Sync) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confVSAMDisp

ViStatus rsfsu_confVSAMDisp (ViSession instrumentHandle, ViInt32 window,

ViInt32 refValueOperation,

ViReal64 refValueX, ViReal64 refValueY,

ViReal64 refValuePosition,

ViReal64 y_perDivision);

Purpose

This function configures the reference value position on the display.

Note:

(1)This function is valid only  with option Tracking Generator/Ext.

Generator Control (FS-B9/B10) and normalizing switched on in NETWORK mode

as well as with the option FM Demodulator (FS-K7) and activated Result

Display FM or with option (FS-K8) Bluetooth.

Remote-control command(s):

DISPlay:WINDow<1|2>:TRACe:Y:RVALue

DISPlay:WINDow<1|2>:TRACe:Y:RPOSition

DISPlay:WINDow<1|2>:TRACe:Y:PDIVision

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

refValueOperation

Variable Type       ViInt32

Specifies the operation to be performed.

Valid Range:

0 - Reserved (Set Ref Value X)

1 - Set Ref Value Y (Default Value)

2 - Reserved (Set Ref Value X And Y)

3 - Set Y per Div

refValueX

Variable Type       ViReal64

This control defines the reference value for the x-axis of the

measurement diagram. This defines the upper limit of the display

range, the corresponding parameters of the manual control

vary according to the measurement mode.

Valid Range: -1.0E12 to 1.0E12

Default Value: 0.0 dB

Note:

(1) Not used in this time

refValueY

Variable Type       ViReal64

This control defines the reference value for the y-axis of the

measurement diagram. This defines the upper limit of the display

range, the corresponding parameters of the manual control

vary according to the measurement mode.

Valid Range: -1.0E12 to 1.0E12

Default Value: 0 dB (mode NETWORK)

0 Hz (mode FM Demodulator or Bluetooth)

refValuePosition

Variable Type       ViReal64

This function defines the position of the y-reference value.

Valid Range: 0.0 % to 100.0 %

Default Value: 100.0 %

y_perDivision

Variable Type       ViReal64

This control defines the scaling of the Y-axis in the current unit.

Valid Range: -1.0E12 to 1.0E12

Default Value: 5.0

Note:

(1) It is only available with option FS-K7 (FM Demodulator) or with

option FS-K8 (Bluetooth).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Ref Value Operation) out of range.

BFFC0004  Parameter 4 (Ref Value X) out of range.

BFFC0005  Parameter 5 (Ref Value Y) out of range.

BFFC0006  Parameter 6 (Ref Value Position [%]) out of range.

BFFC0007  Parameter 7 (Y per Division) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPChTable

ViStatus rsfsu_confWCDPChTable (ViSession instrumentHandle,

ViBoolean channelTable);

Purpose

This function switches the channel table on or off in the code domain

power measurement mode.

On switching on, the measured channel table is stored under the name

RECENT and switched on. After the RECENT channel table is switched on,

another channel table can be selected with the function WCDP Channel

Table File (confWCDPChTableFile). The RECENT channel table must always be

switched on first with this function and then the required channel table

can be selected with the function WCDP Channel Table File

(confWCDPChTableFile).

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K74.

Remote-control command(s):

CONFigure:WCDPower:BTS:CTABle:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTable

Variable Type       ViBoolean

This control switches the channel table on or off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPChTableCatalog

ViStatus rsfsu_confWCDPChTableCatalog (ViSession instrumentHandle,

ViPInt32 number_ofChannelTables,

ViInt32 bufferSize,

ViChar _VI_FAR channelTablesList[]);

Purpose

This function reads out the names of all channel tables stored on the

harddisk in the code domain power measurement mode.

Notes:

(1) Driver returns an instrument specific error, when catalog list is

empty.

(2) This function is available only in conjunction with option FS-K72 or

FS-K74.

Remote-control command(s):

CONFigure:WCDPower:BTS:CTABle:CATalog?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

number_ofChannelTables

Variable Type       ViInt32 (passed by reference)

This control returns number of channel tables, stored on harddisk.

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the string you specify for the catalog

list parameter.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the value contains more

bytes than you specify with the Buffer Size parameter, the function

copies Buffer Size - 1 bytes into the buffer and returns the buffer

size you must pass to get the entire value. For example, if the value

is "file1,file2" and the Buffer Size is 6, the function places

"file1" into the buffer and returns 12.

channelTablesList

Variable Type       ViChar[]

This control returns array of channel tables, stored on harddisk.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the value contains more

bytes than you specify with the Buffer Size parameter, the function

copies Buffer Size - 1 bytes into the buffer and returns the buffer

size you must pass to get the entire value. For example, if the value

is "file1,file2" and the Buffer Size is 6, the function places

"file1" into the buffer and returns 12.

if you pass 0 for buffer Size, you can use VI_NULL for this

parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Number of Channel Tables) is null pointer.

BFFC0003  Parameter 3 (Buffer Size) out of range.

BFFC0004  Parameter 4 (Channel Tables List) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPChTableComment

ViStatus rsfsu_confWCDPChTableComment (ViSession instrumentHandle,

ViString channelTableComment);

Purpose

This function defines a comment for the selected channel table in the

code domain power measurement mode.

Prior to this function, the name of the channel table has to be defined

with function WCDP Channel Table Name (confWCDPChTableName) and the

values of the table have to be defined with function WCDP Channel Table

Data (confWCDPChTableData).

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K74.

Remote-control command(s):

CONFigure:WCDPower:BTS:CTABle:COMMent

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableComment

Variable Type       ViString

This control defines a comment for the selected channel table.

Valid Range: up to 128 characters

Default Value: "Comment for table 1"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Comment) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPChTableCopy

ViStatus rsfsu_confWCDPChTableCopy (ViSession instrumentHandle,

ViString channelTableName);

Purpose

This function copies one channel table onto another one in the code

domain power measurement mode.

The channel table to be copied is selected with function WCDP Channel

Table Name (confWCDPChTableName).

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K74.

Remote-control command(s):

CONFigure:WCDPower:BTS:CTABle:COPY

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableName

Variable Type       ViString

This control defines channel table onto which is copied channel table

selected with the function WCDP Channel Table Name

(confWCDPChTableName).

Valid Range: up to 8 characters

Default Value: "CTAB_2"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPChTableData

ViStatus rsfsu_confWCDPChTableData (ViSession instrumentHandle,

ViInt32 _VI_FAR codeClass[],

ViInt32 _VI_FAR codeNumber[],

ViInt32 _VI_FAR useTFCI[],

ViInt32 _VI_FAR timingOffset[],

ViInt32 _VI_FAR pilotLength[],

ViInt32 _VI_FAR channelType[],

ViInt32 _VI_FAR status[],

ViReal64 _VI_FAR CDPRelative[],

ViInt32 num_ofElements);

Purpose

This function defines the values of the selected channel table in the

code domain power measurement mode. Each line of the table consists of 8

values (each value is represented by array of 3 elements):

code class

code number

use TFCI

timing offset

pilot length

channel type (former pitch flag)

status

CDP relative [dB]

Channels PICH, CPICH and PCCPCH may only be defined once. If channel

CPICH or PCCPCH is missing in the function, it is automatically added at

the end of the table.

Prior to this function, the name of the channel table has to be defined

with function WCDP Channel Table Name (confWCDPChTableName).

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K74.

Remote-control command(s):

CONFigure:WCDPower:BTS:CTABle:DATA

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

codeClass

Variable Type       ViInt32[]

Sets code class.

Valid Range: 2 to 9

Default Value: none

codeNumber

Variable Type       ViInt32[]

Sets code number.

Valid Range: 0 to 511

Default Value: none

useTFCI

Variable Type       ViInt32[]

Sets use of TFCI.

Valid Range:

0: not used

1: used

Default Value: none

timingOffset

Variable Type       ViInt32[]

Sets timing offset. For code class 9, the step width is 512, else

256. For AUTO (-1), the timing offset is calculated by the

instrument.

Valid Range: 0 to 38400 for AUTO mode pass -1

Default Value: none

pilotLength

Variable Type       ViInt32[]

Sets pilot length.

Valid Range:

code class 9: 4

code class 8: 2, 4, 8

code class 7: 4, 8

code class 5/6: 8

code class 2/3/4: 16

Default Value: none

channelType

Variable Type       ViInt32[]

Sets channel type (former pitch flag).

Valid Range:

0: DPCH Dedicated Physical Channel of a standard WCDMA Frame

1: PICH Paging Indication Channel

2: SCCPCH Secondary Common Control Physical Channel

3: HS_SCCH HSDPA: High Speed Shared Control Channel

4: HS_PDSCH HSDPA: High Speed Physical Downlink Shared Channel

5: CHAN any other QPSK modulated channel without pilot symbols

Default Value: none

status

Variable Type       ViInt32[]

Sets status.

Valid Range:

0: not active

1: active

Default Value: none

CDPRelative

Variable Type       ViReal64[]

Sets CDP relative.

Valid Range: any value

Default Value: none

num_ofElements

Variable Type       ViInt32

Sets number of elements for parameter arrays.

Valid Range: 1 to 3

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Code Class) out of range.

BFFC0003  Parameter 3 (Code Number) out of range.

BFFC0004  Parameter 4 (Use TFCI) out of range.

BFFC0005  Parameter 5 (Timing Offset) out of range.

BFFC0006  Parameter 6 (Pilot Length) out of range.

BFFC0007  Parameter 7 (Channel Type) out of range.

BFFC0008  Parameter 8 (Status) out of range.

BFFC0009  Parameter 9 (CDP Relative) out of range.

BFFC000A  Parameter 10 (Num of Elements) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPChTableDelete

ViStatus rsfsu_confWCDPChTableDelete (ViSession instrumentHandle);

Purpose

This function deletes the selected channel table in the code domain power

measurement mode.

The channel table to be deleted is selected with function WCDP Channel

Table Name (confWCDPChTableName).

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K74.

Remote-control command(s):

CONFigure:WCDPower:BTS:CTABle:DELete

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPChTableFile

ViStatus rsfsu_confWCDPChTableFile (ViSession instrumentHandle,

ViString channelTableFile);

Purpose

This function selects a predefined channel table file in the code domain

power measurement mode.

Prior to use this function, the RECENT channel table must be switched on

first with the function WCDP Channel Table (confWCDPChTable).

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K74.

Remote-control command(s):

CONFigure:WCDPower:BTS:CTABle:SELect

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableFile

Variable Type       ViString

This control selects a predefined channel table file.

Valid Range: up to 8 characters

Default Value: "RECENT"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table File) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPChTableName

ViStatus rsfsu_confWCDPChTableName (ViSession instrumentHandle,

ViString channelTableName);

Purpose

This function selects an existing channel table or creates the name of a

new channel table in the code domain power measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K74.

Remote-control command(s):

CONFigure:WCDPower:BTS:CTABle:NAME

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableName

Variable Type       ViString

This control selects an existing channel table or creates the name of

a new channel table.

Valid Range: up to 8 characters

Default Value: "NEW_TAB"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMeas

ViStatus rsfsu_confWCDPMeas (ViSession instrumentHandle,

ViInt32 WCDPMeasurement);

Purpose

This function selects the WCDMA base station measurement tests (power,

ACLR, occupied bandwidth, code domain power, frequency domain, time

domain, emission spectrum mask or complementary cumulative distribution

function).

Note:

(1) This function is available only in conjunction with option FS-K72 or

FS-K74.

Remote-control command(s):

CONFigure:WCDPower:BTS:MEASurement POWer | ACLR | ESPectrum | OBWidth |

WCDPower | FDOMain | TDOMain | CCDF | RFC

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

WCDPMeasurement

Variable Type       ViInt32

This control selects WCDP Measurement mode.

Valid Range:

0 - POWer (Default Value)

1 - ACLR

2 - OBWidth

3 - WCDPower

4 - FDOMain

5 - TDOMain

6 - ESPectrum

7 - CCDF

8 - RFCombi

Notes:

(1) POWer:

Channel power measurement (standard 3GPP WCDMA Forward) with

predefined settings.

(2) ACLR:

Adjacent channel power measurement (standard 3GPP WCDMA Forward) with

predefined settings.

(3) OBWidth:

Measurement of occupied power bandwidth.

(4) WCDPower:

Code domain power measurement.

(5) FDOMain:

Overview measurement in the frequency domain with predefined

settings.

(6) TDOMain:

Measurement of crest factor in the time domain with predefined

settings.

(7) ESPectrum:

Measurement of spectrum emission mask.

(8) CCDF:

CCDF Measurement of complementary cumulative distribution function.

(9) RFCombi:

RF combination measurement of adjacent channel leackage error (ACLR),

spectrum emission mask (SEM) and occupied bandwidth (OBW).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (WCDP Measurement) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMeasMode

ViStatus rsfsu_confWCDPMeasMode (ViSession instrumentHandle,

ViInt32 measurementMode);

Purpose

This function selects the WCDP measurement mode (code domain power, code

domain power ratio, CDP error summary, CDP modulation accuracy, CDP power

versus slot, CDP bit stream, CDP channel assignment table, peak code

domain error, symbol constellation, symbol error vector magnitude, code

power overview (option FS-K73 only)), code domain power absolute scaling,

CDP Error Power, CDP Power Versus Symbol.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73 or FS-K74.

Remote-control command(s):

CALCulate1:FEED 'XPOW:CDP'

CALCulate1:FEED 'XPOW:CDP:RAT'

CALCulate2:FEED 'XTIM:CDP:ERR:SUMM'

CALCulate2:FEED 'XTIM:CDP:MACC'

CALCulate2:FEED 'XTIM:CDP:PVSL'

CALCulate2:FEED 'XTIM:CDP:BSTR'

CALCulate1:FEED 'XTIM:CDP:ERR:CTAB'

CALCulate2:FEED 'XTIM:CDP:ERR:PCD'

CALCulate2:FEED 'XTIM:CDP:SYMB:CONS'

CALCulate2:FEED 'XTIM:CDP:SYMB:EVM'

CALCulate1:FEED 'XPOW:CDP:OVER'

CALCulate1:FEED 'XPOW:CDP:ABS'

CALCulate1:FEED 'XPOW:CDEP'

CALCulate2:FEED 'XTIM:CDP:PVSYmbol'

CALCulate1:FEED 'XTIM:CDP:PVSL:ABS'

CALCulate1:FEED 'XTIM:CDP:PVSL:RAT'

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measurementMode

Variable Type       ViInt32

This control selects the operating mode in vector signal analysis.

Valid Range:

0 - Code Domain Power (Default Value)

1 - Code Domain Power Ratio

2 - CDP Error Summary

3 - CDP Modulation Accuracy

4 - CDP Power Versus Slot

5 - CDP Bit Stream

6 - CDP Channel Assignment Table

7 - Peak Code Domain Error

8 - Symbol Constellation

9 - Symbol Error Vector Magnitude

10 - Code Power Overview (option FS-K73 only)

11 - Code Domain Power Absolute Scaling

12 - CDP Error Power

13 - CDP Power Versus Symbol

14 - Code Power Absolute

15 - Code Power Relative

Notes:

(1) Code Domain Power:

Result display of code domain power as bargraph.

(2) Code Domain Power Ratio:

Result display of code domain power ratio.

(3) CDP Tabular Form of Results:

Result display in tabular form.

(4) CDP Modulation Accuracy:

Result display of modulation accuracy.

(5) CDP Power Versus Slot:

Result display of power versus slot.

(6) CDP Bit Stream:

Result display of bit stream.

(7) CDP Channel Assignment Table:

Result disp of channel assignment table.

(8) Peak Code Domain Error:

Result display of peak code domain error.

(9) Symbol Constellation:

Result disp of symbol constellation.

(10) Symbol Error Vector Magnitude:

Result disp of symbol error vector magnitude.

(11) Code Power Overview:

Screen A displays CDP Rel I, screen B CDP Q (option FSIQK73 only).

(12) Code Domain Power Absolute Scaling:

Result display of code domain power as bargraph absolute scaling

(13) CDP Error Power:

Result display of code domain error power as bargraph.

(14) CDP Power Versus Symbol:

Result display of power versus symbol.

(15) Code Power Absolute:

If the power versus slot measurement is active, the slot power is

given in absolute values.

(16) Code Power Relative:

If the power versus slot measurement is active, the slot power is

given in relative values.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Measurement Mode) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSChEvaluation

ViStatus rsfsu_confWCDPMSChEvaluation (ViSession instrumentHandle,

ViInt32 base_ofMeasurement);

Purpose

This function determines the base of the WCDPower Mobile station

measurement.

Note:

(1) This function is available only in conjunction with option  FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:EVALuation:LENGth SLOT | FRAMe

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

base_ofMeasurement

Variable Type       ViInt32

Selects the base of the measurments.

Valid Range:

0 - Slot (Default Value)

1 - Frame

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Base of Measurement) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSChHSDPCCH

ViStatus rsfsu_confWCDPMSChHSDPCCH (ViSession instrumentHandle,

ViBoolean state);

Purpose

This function activates the HS-DPCCH channel in an user defined channel

table.

Note:

(1) This function is available only in conjunction with option  FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:CTABle:DATA:HSDPcch ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

state

Variable Type       ViBoolean

Activates [ON] or inactivates [OFF] the HS-DPCCH entry in a

predefined channel table.

Valid Values:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSChTable

ViStatus rsfsu_confWCDPMSChTable (ViSession instrumentHandle,

ViBoolean channelTable);

Purpose

This function switches the channel table on or off in the code domain

power measurement mode.

On switching on, the measured channel table is requested by the DSPs,

stored under the name RECENT and switched on. After the RECENT channel

table is switched on, another channel table can be selected with the

function WCDP MS Channel Table File (confWCDPMSChTableFile). The RECENT

channel table must always be switched on first with this function and

then the required channel table can be selected with the function WCDP MS

Channel Table File (confWCDPMSChTableFile).

Note:

(1) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:CTABle:STATe ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTable

Variable Type       ViBoolean

This control switches the channel table on or off.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSChTableCatalog

ViStatus rsfsu_confWCDPMSChTableCatalog (ViSession instrumentHandle,

ViPInt32 number_ofChannelTables,

ViInt32 bufferSize,

ViChar _VI_FAR channelTablesList[]);

Purpose

This function reads out the names of all channel tables stored on the

harddisk in the code domain power measurement mode.

Notes:

(1) Driver returns an instrument specific error, when catalog list is

empty.

(2) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:CTABle:CATalog?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

number_ofChannelTables

Variable Type       ViInt32 (passed by reference)

This control returns number of channel tables, stored on harddisk.

bufferSize

Variable Type       ViInt32

Pass the number of bytes in the string you specify for the catalog

list parameter.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the value contains more

bytes than you specify with the Buffer Size parameter, the function

copies Buffer Size - 1 bytes into the buffer and returns the buffer

size you must pass to get the entire value. For example, if the value

is "file1,file2" and the Buffer Size is 6, the function places

"file1" into the buffer and returns 12.

channelTablesList

Variable Type       ViChar[]

This control returns array of channel tables, stored on harddisk.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the value contains more

bytes than you specify with the Buffer Size parameter, the function

copies Buffer Size - 1 bytes into the buffer and returns the buffer

size you must pass to get the entire value. For example, if the value

is "file1,file2" and the Buffer Size is 6, the function places

"file1" into the buffer and returns 12.

if you pass 0 for buffer Size, you can use VI_NULL for this

parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Number of Channel Tables) is null pointer.

BFFC0003  Parameter 3 (Buffer Size) out of range.

BFFC0004  Parameter 4 (Channel Tables List) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSChTableComment

ViStatus rsfsu_confWCDPMSChTableComment (ViSession instrumentHandle,

ViString channelTableComment);

Purpose

This function defines a comment for the selected channel table in the

code domain power measurement mode.

Prior to this function, the name of the channel table has to be defined

with function WCDP MS Channel Table Name (confWCDPMSChTableName) and the

values of the table have to be defined with function WCDP MS Channel

Table Data (confWCDPMSChTableData).

Note:

(1) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:CTABle:COMMent

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableComment

Variable Type       ViString

This control defines a comment for the selected channel table.

Valid Range: up to 128 characters

Default Value: "Comment for table 1"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Comment) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSChTableCopy

ViStatus rsfsu_confWCDPMSChTableCopy (ViSession instrumentHandle,

ViString channelTableName);

Purpose

This function copies one channel table onto another one in the code

domain power measurement mode.

The channel table to be copied is selected with function WCDP MS Channel

Table Name (confWCDPMSChTableName).

Note:

(1) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:CTABle:COPY

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableName

Variable Type       ViString

This control defines channel table onto which is copied channel table

selected with the function WCDP MS Channel Table Name

(confWCDPMSChTableName).

Valid Range: up to 8 characters

Default Value: "CTAB_2"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSChTableData

ViStatus rsfsu_confWCDPMSChTableData (ViSession instrumentHandle,

ViInt32 codeClass,

ViInt32 pilotLength,

ViInt32 number_ofActiveChannels);

Purpose

This function defines the values of the selected channel table in the

code domain power measurement mode:

Pilot Length

Code Class

Number of Active Channels

Prior to this function, the name of the channel table has to be defined

with function WCDP MS Channel Table Name (confWCDPMSChTableName).

Note:

(1) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:CTABle:DATA

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

codeClass

Variable Type       ViInt32

Sets code class of channel 1. I-mapped.

Valid Range: not checked

Default Value: 8

pilotLength

Variable Type       ViInt32

Sets pilot length of channel DPCCH.

Valid Range: not checked

Default Value: 4

number_ofActiveChannels

Variable Type       ViInt32

Sets number of active channels.

Valid Range: 0 to 6

Default Value: 1

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Code Class) out of range.

BFFC0003  Parameter 3 (Pilot Length) out of range.

BFFC0004  Parameter 4 (Number of Active Channels) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSChTableDelete

ViStatus rsfsu_confWCDPMSChTableDelete (ViSession instrumentHandle);

Purpose

This function deletes the selected channel table in the code domain power

measurement mode.

The channel table to be deleted is selected with function WCDP MS Channel

Table Name (confWCDPMSChTableName).

Note:

(1) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:CTABle:DELete

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSChTableFile

ViStatus rsfsu_confWCDPMSChTableFile (ViSession instrumentHandle,

ViString channelTableFile);

Purpose

This function selects a predefined channel table file in the code domain

power measurement mode.

Prior to use this function, the RECENT channel table must be switched on

first with the function WCDP MS Channel Table (confWCDPMSChTable).

Note:

(1) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:CTABle:SELect

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableFile

Variable Type       ViString

This control selects a predefined channel table file.

Valid Range: up to 8 characters

Default Value: "RECENT"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table File) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSChTableName

ViStatus rsfsu_confWCDPMSChTableName (ViSession instrumentHandle,

ViString channelTableName);

Purpose

This function selects an existing channel table or creates the name of a

new channel table in the code domain power measurement mode.

Note:

(1) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:CTABle:NAME

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

channelTableName

Variable Type       ViString

This control selects an existing channel table or creates the name of

a new channel table.

Valid Range: up to 8 characters

Default Value: "NEW_TAB"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Channel Table Name) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_confWCDPMSMeas

ViStatus rsfsu_confWCDPMSMeas (ViSession instrumentHandle,

ViInt32 WCDPMeasurement);

Purpose

This function selects the WCDMA mobile station measurement tests (power,

ACLR, occupied bandwidth, code domain power, frequency domain, time

domain, emission spectrum mask or complementary cumulative distribution

function).

Note:

(1) This function is available only in conjunction with option FS-K73.

Remote-control command(s):

CONFigure:WCDPower:MS:MEASurement POWer | ACLR | ESPectrum | OBWidth |

WCDPower | FDOMain | TDOMain | CCDF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

WCDPMeasurement

Variable Type       ViInt32

This control selects WCDP Measurement mode.

Valid Range:

0 - POWer (Default Value)

1 - ACLR

2 - OBWidth

3 - WCDPower

4 - FDOMain

5 - TDOMain

6 - ESPectrum

7 - CCDF

Notes:

(1) POWer:

Channel power measurement (standard 3GPP WCDMA Forward) with

predefined settings.

(2) ACLR:

Adjacent channel power measurement (standard 3GPP WCDMA Forward) with

predefined settings.

(3) OBWidth:

Measurement of occupied power bandwidth.

(4) WCDPower:

Code domain power measurement.

(5) FDOMain:

Overview measurement in the frequency domain with predefined

settings.

(6) TDOMain:

Measurement of crest factor in the time domain with predefined

settings.

(7) ESPectrum:

ESPectrum Measurement of spectrum emission mask.

(8) CCDF:

CCDF Measurement of complementary cumulative distribution function.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (WCDP Measurement) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataFetchBurst

ViStatus rsfsu_dataFetchBurst (ViSession instrumentHandle,

ViBoolean phaseFrequencyError,

ViBoolean resultType,

ViBoolean measResult,

ViPReal64 measuredValue);

Purpose

This function is used for reading out results of measurements in the GSM

MS (option FS-K5) Analyzer mode, which are performed on individual bursts

(Phase/Frequency Error) without starting the measurement by themselves.

Notes:

(1) This function is available only in conjunction with option FS-K5, and

if measurement of the phase/frequency error is selected.

(2) If no measurement has been performed yet, a query error will occur.

Remote-control command(s):

*CLS

FETCh:BURSt:PERRor:RMS:AVERage? | MAXimum?

FETCh:BURSt:PERRor:PEAK:AVERage? | MAXimum?

FETCh:BURSt:FERRor:AVERage? | MAXimum?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

phaseFrequencyError

Variable Type       ViBoolean

This control selects phase or frequency measurement.

Valid Range:

VI_FALSE (0) - Phase (Default Value)

VI_TRUE  (1) - Frequency

resultType

Variable Type       ViBoolean

This control selects result type.

Valid Range:

VI_FALSE (0) - RMS (Default Value)

VI_TRUE  (1) - Peak

Note:

(1) Control is ignored for Frequency Error.

measResult

Variable Type       ViBoolean

This control selects measurement result.

Valid Range:

VI_FALSE (0) - Average (Default Value)

VI_TRUE  (1) - Maximum

measuredValue

Variable Type       ViReal64 (passed by reference)

This control outputs measured value for the selected parameters.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Phase/Frequency Error) out of range.

BFFC0003  Parameter 3 (Result Type) out of range.

BFFC0004  Parameter 4 (Meas Result) out of range.

BFFC0005  Parameter 5 (Measured Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataFetchBurstIQ

ViStatus rsfsu_dataFetchBurstIQ (ViSession instrumentHandle,

ViInt32 IQMeasurement, ViPInt32 status,

ViPReal64 average, ViPReal64 maximum);

Purpose

This function reads out the IQ DC offset and the IQ imbalance measurement

in % for the selected number of bursts.

Note:

(1) It is only available in GSM/EDGE mode (FS-K5).

(2) If no measurement has been performed yet, a query error will occur.

(3) It is available only if phase/frequency error (PFE) measurement has

been selected.

Remote-control command(s):

*CLS

FETCh:BURSt:IQOFfset:AVERage? | MAXimum?

FETCh:BURSt:IQIMbalance:AVERage? | MAXimum?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

IQMeasurement

Variable Type       ViInt32

This control selects modulation accuracy values to be measured.

Valid Range:

0 - IQ Offset (Default Value)

1 - IQ Imbalance

status

Variable Type       ViInt32 (passed by reference)

This control returns the status of the selected measurement.

(Results = 0: failed, 1: passed).

You can use VI_NULL for this parameter.

Note:

(1) This parameter is for future use. It always returns 1.

average

Variable Type       ViReal64 (passed by reference)

This control returns the average of the selected measurement.

You can use VI_NULL for this parameter.

maximum

Variable Type       ViReal64 (passed by reference)

This control returns the maximum of the selected measurement.

You can use VI_NULL for this parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (IQ Measurement) out of range.

BFFC0003  Parameter 3 (Status) is null pointer.

BFFC0004  Parameter 4 (Average) is null pointer.

BFFC0005  Parameter 5 (Maximum) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataFetchBurstMACC

ViStatus rsfsu_dataFetchBurstMACC (ViSession instrumentHandle,

ViInt32 MACCMeasurement,

ViPInt32 RESERVED, ViPReal64 average,

ViPReal64 maximum);

Purpose

This function reads the selected measurement of the modulation accuracy

of the base station or mobile for the selected number of bursts.

Note:

(1) It is available only in conjunction with option GSM/EDGE Analyzer,

FS-K5.

(2) If no measurement has been performed yet, a query error will occur.

(3) It is available only if measurement of the modulation accuracy is

selected (see GSM Modulation Accuracy (actGSMModAccuracy)).

Remote-control command(s):

*CLS

FETCh:BURSt:MACCuracy:RMS:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:PEAK:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:OSUPpress:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:PERCentile:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:FREQuency:AVERage? | MAXimum?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

MACCMeasurement

Variable Type       ViInt32

This control selects modulation accuracy values to be measured.

Valid Range:

0 - RMS (Default Value)

1 - Peak

2 - Osuppress

3 - Percentile

4 - Frequency

RESERVED

Variable Type       ViInt32 (passed by reference)

This parameter is not used by the function.

You must pass VI_NULL for this parameter.

average

Variable Type       ViReal64 (passed by reference)

This control returns the average of the selected measurement of the

modulation accuracy taken over the selected number of bursts.

You can use VI_NULL for this parameter.

maximum

Variable Type       ViReal64 (passed by reference)

This control returns the maximum of the selected measurement of the

modulation accuracy taken over the selected number of bursts.

You can use VI_NULL for this parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (MACC Measurement) out of range.

BFFC0003  Parameter 3 (RESERVED) is null pointer.

BFFC0004  Parameter 4 (Average) is null pointer.

BFFC0005  Parameter 5 (Maximum) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataFetchModRef

ViStatus rsfsu_dataFetchModRef (ViSession instrumentHandle,

ViPReal64 measuredLevel,

ViPReal64 correctedLevel,

ViPReal64 bandwidth);

Purpose

This function reads out the results of the premeasurement of modulation

due to spectrum measurement.

Note:

This function is only available with GSM/EDGE Analyzer FS-K5 option and

when measurement modulation due to spectrum is selected.

Remote-control command(s):

*CLS

FETCh:SPECtrum:MODulation:REFerence?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measuredLevel

Variable Type       ViReal64 (passed by reference)

This control returns the measured level value.

correctedLevel

Variable Type       ViReal64 (passed by reference)

This control returns the corrected level value. This value is

corrected by means of the bandwidth.

bandwidth

Variable Type       ViReal64 (passed by reference)

This control returns the bandwidth value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Measured Level) is null pointer.

BFFC0003  Parameter 3 (Corrected Level) is null pointer.

BFFC0004  Parameter 4 (Bandwidth) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataFetchNoiseMeasArray

ViStatus rsfsu_dataFetchNoiseMeasArray (ViSession instrumentHandle,

ViInt32 measType,

ViInt32 arraySize,

ViReal64 _VI_FAR measResult[]);

Purpose

This function requests the R&S FS-K30 option to return the last recorded

noise measurement results. The results will be returned as an array of up

to 100 elements, either of noise figure, noise temperature or noise gain

results depending on the argument used.

This function will only return a result array if the y-axis measurement

for the passed argument is currently active.

Noise gain measurements will be active if the remote control command

DISP:WIND:TRAC:Y:SCAL:GAIN ON was previously entered.

Noise figure or temperature measurements will be active if the remote

control command DISP:WIND:TRAC:Y:SCAL:NFIG ON or

DISP:WIND:TRAC:Y:SCAL:NTEM ON was previously entered.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

FETCh:ARRAy:NOISE:FIGure|TEMPerature|GAIN?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measType

Variable Type       ViInt32

This parameter selects particular measurement type.

Valid Range:

0 - Noise Figure

1 - Noise Temperature

2 - Gain

Default Value: 0

arraySize

Variable Type       ViInt32

measResult

Variable Type       ViReal64[]

Returns the last recorded noise measurement results. The results will

be returned as an array of up to 100 elements.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Meas Type) out of range.

BFFC0003  Parameter 3 (Array Size) out of range.

BFFC0004  Parameter 4 (Meas Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataFetchNoiseMeasScalar

ViStatus rsfsu_dataFetchNoiseMeasScalar (ViSession instrumentHandle,

ViInt32 measType,

ViPReal64 measResult);

Purpose

This function requests the R&S FS-K30 option to return the last recorded

noise measurement result for a single frequency measurement. The results

will be a single noise figure, noise temperature or noise gain results

depending on the argument used.

This function will only return a result if the y-axis measurement for the

passed argument is currently active.

Noise gain measurements will be active if the remote control command

DISP:WIND:TRAC:Y:SCAL:GAIN ON was previously entered.

Noise figure or temperature measurements will be active if the remote

control command DISP:WIND:TRAC:Y:SCAL:NFIG ON or

DISP:WIND:TRAC:Y:SCAL:NTEM ON was previously entered.

Note:

(1) This function is available only in conjunction with option FS-K30.

Remote-control command(s):

FETCh:SCALar:NOISE:FIGure|TEMPerature|GAIN?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measType

Variable Type       ViInt32

This parameter selects particular measurement type.

Valid Range:

0 - Noise Figure

1 - Noise Temperature

2 - Gain

Default Value: 0

measResult

Variable Type       ViReal64 (passed by reference)

The last recorded noise measurement result for a single frequency

measurement.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Meas Type) out of range.

BFFC0003  Parameter 3 (Meas Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataFetchPMetResult

ViStatus rsfsu_dataFetchPMetResult (ViSession instrumentHandle,

ViInt32 window,

ViPReal64 measurementResult);

Purpose

This function outputs the result of the power sensor. If no measurement

has been carried out yet, a query error is triggered.

Note:

(1) This function is available only in conjunction with option FS-K9.

Remote-control command(s):

FETCh<1|2>:PMETer?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

measurementResult

Variable Type       ViReal64 (passed by reference)

Outputs the result of the power sensor.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Measurement Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataFetchPtempRef

ViStatus rsfsu_dataFetchPtempRef (ViSession instrumentHandle,

ViPReal64 measuredLevel,

ViPReal64 correctedLevel,

ViPReal64 bandwidth);

Purpose

This function reads out the results of the premeasurement of power vs.

time.

Note:

This function is only available with GSM/EDGE Analyzer FS-K5 option and

when measurement of the power vs. time is selected.

Remote-control command(s):

*CLS

FETCh:BURSt:PTEMplate:REFerence?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measuredLevel

Variable Type       ViReal64 (passed by reference)

This control returns the measured level value.

correctedLevel

Variable Type       ViReal64 (passed by reference)

This control returns the corrected level value. This value is

corrected by means of the bandwidth.

bandwidth

Variable Type       ViReal64 (passed by reference)

This control returns the bandwidth value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Measured Level) is null pointer.

BFFC0003  Parameter 3 (Corrected Level) is null pointer.

BFFC0004  Parameter 4 (Bandwidth) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataFetchSpectrum

ViStatus rsfsu_dataFetchSpectrum (ViSession instrumentHandle,

ViBoolean spectrumMeasMode,

ViInt32 modulationRange,

ViInt32 _VI_FAR index[],

ViReal64 _VI_FAR startFrequencies[],

ViReal64 _VI_FAR stopFrequencies[],

ViReal64 _VI_FAR levels[],

ViReal64 _VI_FAR limits[],

ViInt32 _VI_FAR absRel[],

ViInt32 _VI_FAR status[],

ViPInt32 num_ofResults);

Purpose

This function reads out results of measurements in the GSM/EDGE (FS-K5)

Analyzer mode, used to measure the power of the spectral contributions

due to modulation and switching (modulation spectrum, transient spectrum)

without first restarting a new measurement.

Notes:

(1) This function is available only in conjunction with option FS-K5, and

when measurement of the modulation spectrum is selected.

(2) If no measurement has been performed yet, a query error will occur.

Remote-control command(s):

*CLS

FETCh:SPECtrum:SWITching?

FETCh:SPECtrum:MODulation? ARFCn | TXBand | RXBand| COMBined | DCSRx1800

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

spectrumMeasMode

Variable Type       ViBoolean

This control sets spectrum measurement mode.

Valid Range:

VI_FALSE (0) - Switching

VI_TRUE  (1) - Modulation (Defaul Value)

Notes:

(1) Switching:

This command reads out the result of the measurement of the transient

spectrum of the base station or mobile.

(2) Modulation:

This command reads out the result of the measurement of the

modulation spectrum of the base station or mobile.

modulationRange

Variable Type       ViInt32

This control selects the frequency range for the modulation

measurement result.

Valid Range:

0 - ARFCN � 1.8 MHz (Default Value)

1 - TX Band

2 - RX Band

3 - ARFCN � 1.8 MHz / TX Band

4 - DCSRx1800 RX band DCS 1800 (option FSE-K10 only)

index

Variable Type       ViInt32[]

Returns indexed array of measurement range (last number indicates sum

of measurements).

startFrequencies

Variable Type       ViReal64[]

Returns start frequencies of the measurement range or frequencies

where the limit is exceeded.

stopFrequencies

Variable Type       ViReal64[]

Returns stop frequencies of the measurement range or frequencies

where the measured range is exceeded. This values are equal to the

values of start frequencies, if either the measurement is performed

in the time domain or the result value contains a limit excess.

levels

Variable Type       ViReal64[]

Returns array of measured maximum levels of the partial range or

measured levels at the test point.

limits

Variable Type       ViReal64[]

Returns limits in the partial range or measured levels at the test

point.

absRel

Variable Type       ViInt32[]

Returns array of ABS/REL unit indicators. Meaning of numbers is

following:

ABS = 0

REL = 1

status

Variable Type       ViInt32[]

Limit checks array of results.

PASSED   ... no limits exceeded

FAILED   ... limits exceeded

MARGIN   ... margin exceeded

EXC      ... limited excesses characterized as an exception

Meaning of numbers is:

PASSED = 0

FAILED = 1

MARGIN = 2

EXC = 3

num_ofResults

Variable Type       ViInt32 (passed by reference)

This control returns number of results, received from the instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Spectrum Meas Mode) out of range.

BFFC0003  Parameter 3 (Modulation Range) out of range.

BFFC0004  Parameter 4 (Index) is null pointer.

BFFC0005  Parameter 5 (Start Frequencies) is null pointer.

BFFC0006  Parameter 6 (Stop Frequencies) is null pointer.

BFFC0007  Parameter 7 (Levels) is null pointer.

BFFC0008  Parameter 8 (Limits) is null pointer.

BFFC0009  Parameter 9 (Abs/Rel) is null pointer.

BFFC000A  Parameter 10 (Status) is null pointer.

BFFC000B  Parameter 11 (Num of Results) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadBurst

ViStatus rsfsu_dataReadBurst (ViSession instrumentHandle,

ViInt32 phaseFrequencyError,

ViBoolean resultType,

ViBoolean measResult,

ViPReal64 measuredValue);

Purpose

This function starts the measurement of the phase and frequency error of

the base station or mobile and reads out the average of the

RMS-measurement of the phase error taken over the selected number of

bursts.

Notes:

(1) This function is available only in conjunction with option FS-K5, and

if measurement of the phase/frequency error is selected.

(2) When the measurement is started the instrument automatically assumes

the SINGLE mode.

(3) An ongoing measurement can be aborted via the command ABORt.

(4) Further results of the phase/frequency error measurement can be then

queried without restart of the measurement via the FETCh:BURSt-subsystem.

Remote-control command(s):

*CLS

READ:BURSt:PERRor:RMS:AVERage? | MAXimum?

READ:BURSt:PERRor:PEAK:AVERage? | MAXimum?

READ:BURSt:FERRor:AVERage? | MAXimum?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

phaseFrequencyError

Variable Type       ViInt32

This control selects phase or frequency measurement.

Valid Range:

0 - Phase (Default Value)

1 - Frequency

2 - IQ DC-Offset

3 - IQ Imbalance

resultType

Variable Type       ViBoolean

This control selects result type.

Valid Range:

VI_FALSE (0) - RMS (Default Value)

VI_TRUE  (1) - Peak

For Frequency Error measurement this control is not used.

measResult

Variable Type       ViBoolean

This control selects measurement result.

Valid Range:

VI_FALSE (0) - Average (Default Value)

VI_TRUE  (1) - Maximum

measuredValue

Variable Type       ViReal64 (passed by reference)

This control outputs measured value for the selected parameters.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Phase/Frequency Error) out of range.

BFFC0003  Parameter 3 (Result Type) out of range.

BFFC0004  Parameter 4 (Meas Result) out of range.

BFFC0005  Parameter 5 (Measured Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadBurstIQ

ViStatus rsfsu_dataReadBurstIQ (ViSession instrumentHandle,

ViInt32 IQMeasurement, ViPInt32 status,

ViPReal64 average, ViPReal64 maximum);

Purpose

This function triggers the measurement of the phase/frequency error (PFE)

of the base station or of the mobile and reads out the IQ DC offset

measurement or IQ imbalance measurement for the selected number of

bursts.

Note:

(1) It is only available GSM/EDGE mode (FS-K5).

(2) If no measurement has been performed yet, a query error will occur.

(3) It is available only if phase/frequency error (PFE) measurement has

been selected.

The start of the measurement automatically activates the single sweep

mode.

An ongoing measurement can be aborted by means of the ABORt command.

Further results of the PFE measurement can then be queried via the

function Fetch Burst IQ Values (dataFetchBurstIQ) without restarting the

measurement

Remote-control command(s):

*CLS

READ:BURSt:IQOFfset:AVERage? | MAXimum?

READ:BURSt:IQIMbalance:AVERage? | MAXimum?

FETCh:BURSt:IQOFfset:MAXimum?

FETCh:BURSt:IQIMbalance:MAXimum?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

IQMeasurement

Variable Type       ViInt32

This control selects modulation accuracy values to be measured.

Valid Range:

0 - IQ Offset (Default Value)

1 - IQ Imbalance

status

Variable Type       ViInt32 (passed by reference)

This control returns the status of the selected measurement.

(Results = 0: failed, 1: passed).

You can use VI_NULL for this parameter.

Notes:

(1) This value is not used no. It always returns 1.

average

Variable Type       ViReal64 (passed by reference)

This control returns the average of the selected measurement.

You can use VI_NULL for this parameter.

maximum

Variable Type       ViReal64 (passed by reference)

This control returns the maximum of the selected measurement.

You can use VI_NULL for this parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (IQ Measurement) out of range.

BFFC0003  Parameter 3 (Status) is null pointer.

BFFC0004  Parameter 4 (Average) is null pointer.

BFFC0005  Parameter 5 (Maximum) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadBurstMACC

ViStatus rsfsu_dataReadBurstMACC (ViSession instrumentHandle,

ViInt32 MACCMeasurement,

ViPInt32 RESERVED, ViPReal64 average,

ViPReal64 maximum);

Purpose

This function starts the selected measurement of the modulation accuracy

of the base station or mobile and reads out the result of the measurement

for the selected number of bursts.

Note:

(1) It is available only in conjunction with option GSM/EDGE Analyzer,

FS-K5.

(2) When the measurement is started the instrument automatically assumes

the SINGLE mode. An ongoing measurement can be aborted via the command

ABORt. Further results of the modulation accuracy measurement can be then

queried without restart of the

measurement via the Fetch Burst MACC Values (dataFetchBurstMACC).

(3) It is available only if measurement of the modulation accuracy is

selected (see GSM Modulation Accuracy (actGSMModAccuracy)).

Remote-control command(s):

*CLS

READ:BURSt:MACCuracy:RMS:AVERage? | MAXimum?

READ:BURSt:MACCuracy:PEAK:AVERage? | MAXimum?

READ:BURSt:MACCuracy:OSUPpress:AVERage? | MAXimum?

READ:BURSt:MACCuracy:PERCentile:AVERage? | MAXimum?

READ:BURSt:MACCuracy:FREQuency:AVERage? | MAXimum?

READ:BURSt:MACCuracy:IQOFfset:AVERage? | MAXimum?

READ:BURSt:MACCuracy:IQIMbalance:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:RMS:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:PEAK:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:OSUPpress:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:PERCentile:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:FREQuency:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:IQOFfset:AVERage? | MAXimum?

FETCh:BURSt:MACCuracy:IQIMbalance:AVERage? | MAXimum?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

MACCMeasurement

Variable Type       ViInt32

This control selects modulation accuracy values to be measured.

Valid Range:

0 - RMS (Default Value)

1 - Peak

2 - Osuppress

3 - Percentile

4 - Frequency

5 - IQ Offset

6 - IQ Imbalance

RESERVED

Variable Type       ViInt32 (passed by reference)

This parameter is not used by the function.

You must pass VI_NULL for this parameter.

average

Variable Type       ViReal64 (passed by reference)

This control returns the average of the selected measurement of the

modulation accuracy taken over the selected number of bursts.

You can use VI_NULL for this parameter.

maximum

Variable Type       ViReal64 (passed by reference)

This control returns the maximum of the selected measurement of the

modulation accuracy taken over the selected number of bursts.

You can use VI_NULL for this parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (MACC Measurement) out of range.

BFFC0003  Parameter 3 (RESERVED) is null pointer.

BFFC0004  Parameter 4 (Average) is null pointer.

BFFC0005  Parameter 5 (Maximum) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadBurstRef

ViStatus rsfsu_dataReadBurstRef (ViSession instrumentHandle,

ViPReal64 referenceLevel);

Purpose

This function starts the premeasurement and as a result provides the

measured level in dBm in the GSM/EDGE analyzer mode.

Notes:

(1) This function is available only in conjunction with option FS-K5.

(2) This function has been included only for reasons of compatibility

with FSE-K10. Instead of this function, the function Read Burst Ptemplate

Ref (rsfsu_dataReadPtempRef) should be used, whose output format conforms

to that of the premeasurement query commands.

Remote-control command(s):

*CLS

READ:BURSt:REFerence?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

referenceLevel

Variable Type       ViReal64 (passed by reference)

This control returns the measured level in dBm.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Reference Level) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadLevelTime

ViStatus rsfsu_dataReadLevelTime (ViSession instrumentHandle,

ViPBoolean status,

ViPReal64 signalPowerdBm,

ViPReal64 triggerTimes,

ViPReal64 triggerLeveldBmV,

ViPReal64 reserved);

Purpose

This function starts a measurement sequence which automatically adjust

the level and trigger timing of the input signal to the analyzer.

Notes:

(1) This function is available only in conjunction with option FS-K5, and

when measurement of the modulation spectrum is selected.

Remote-control command(s):

READ:AUTO:LEVTime?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

status

Variable Type       ViBoolean (passed by reference)

This control returns the measurement successuf state

0=PASSED - measurement successful, result are valid

1=FAILED - measurement failed, results figures are invalid

signalPowerdBm

Variable Type       ViReal64 (passed by reference)

This control returns the corrected level signal power (reference

level is automatically set 3 dB above the signal power)

triggerTimes

Variable Type       ViReal64 (passed by reference)

This control returns the correction value for the time offset between

the currently active GSM trigger and beginning of the slot.

triggerLeveldBmV

Variable Type       ViReal64 (passed by reference)

This control returns trigger level of the current active GSM trigger.

For IF/RF the units are dBm

for the external trigger the unit is Volts

reserved

Variable Type       ViReal64 (passed by reference)

For future use, always 0.0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Status) is null pointer.

BFFC0003  Parameter 3 (Signal Power [dBm]) is null pointer.

BFFC0004  Parameter 4 (Trigger Time [s]) is null pointer.

BFFC0005  Parameter 5 (Trigger Level [dBm/V]) is null pointer.

BFFC0006  Parameter 6 (Reserved) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadMemoryIQ

ViStatus rsfsu_dataReadMemoryIQ (ViSession instrumentHandle,

ViInt32 offsetSamples,

ViInt32 no_ofSamples,

ViInt32 bufferSize,

ViPInt32 no_ofPoints,

ViReal64 _VI_FAR realParts_I[],

ViReal64 _VI_FAR imaginaryParts_Q[]);

Purpose

This command permits the readout of previously acquired (and frequency

response corrected) I/Q data from the memory, with indication of the

offset related to the start of measurement and with indication of the

number of measurement values. Therefore a previously acquired data set

can be read out in smaller portions. The result values are scaled linear

in unit Volt and correspond to the voltage at the RF input of the

instrument.

Notes:

(1) For large I/Q trace data is recommended to increase default timeout

value to the appropriate value which allows instrument handle all pending

tasks. To change default timeout value call following VISA function:

viSetAttribute (instrSession, VI_ATTR_TMO_VALUE, 60000);

This attribute specifies the minimum timeout value to use (in

milliseconds) when accessing the device associated with the given

session.

(2) If there are not I/Q data available in memory because the

corresponding measurement had not been started, the function will cause a

Query Error.

Example:

(1) Switch on I/Q data acquisition:

rsfsu_confTraceIQState (instrSession, VI_TRUE);

(2) Choose configuration:

rsfsu_confTraceIQ (instrSession, 32.0e6, 0, VI_TRUE,

1.0e6, 0, 1024);

(3) Start measurement and read results:

rsfsu_dataReadTraceIQ (instrSession, bufferSize, &points,

I, Q);

(4) Read memory results:

rsfsu_dataReadMemoryIQ (instrSession, offset, noSamples,

bufferSize, &points, I, Q);

(5) Switch off I/Q data acquisition:

rsfsu_confTraceIQState (instrSession, VI_FALSE)

Remote-control command(s):

FORMat REAL,32

TRACe:IQ:DATA:MEMory?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

offsetSamples

Variable Type       ViInt32

This control sets offset of the values to be read related to the

start of the acquired data.

Valid Range: 0 to <# of samples> - 1

Default Value: 0

Note:

Where <# of samples> is the value indicated with function Trace I/Q

Set (rsfsu_confTraceIQ).

no_ofSamples

Variable Type       ViInt32

This control sets the number of measurement values to be read.

Valid Range: 1 to <# of samples> - <offset samples>

Default Value: 1

Note:

Where <# of samples> is the value indicated with function Trace I/Q

Set (rsfsu_confTraceIQ).

bufferSize

Variable Type       ViInt32

Pass the number of doubles in the array you specify for the real

parts and imaginary parameters.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the data contains more

elemens than you specify with the Buffer Size parameter, the function

returns Buffer Size elements into the arrays and returns the buffer

size you must pass to get the all data.  For example, if the data

contains 16000 elements for real parts and 16000 elements for

imaginary part and the Buffer Size is 10000, the function places

10000 elements into the arrays and returns 16000.

Notes:

(1) If VI_ATTR_INTF_TYPE is equal to VI_INTF_ASRL, then use this

parameter to pass the expected number of points to be read from the

instrument. Data points will be read until instruments output buffer

is empty. If Buffer Size is lower than actual number of results

available in the instruments output buffer, then this function

returns an error code (BFFC0002  Parameter 2 (Buffer Size) out of

range.) and measurement results will not be valid. This is due to

limitation if the unit is remote-controlled via the RS-232-C

interface.

(2) This function requires that all response data are read out

completely before the instrument accepts further commands.

no_ofPoints

Variable Type       ViInt32 (passed by reference)

Returns the number of I/Q trace data points.

realParts_I

Variable Type       ViReal64[]

Returns an array of trace data (I-value x, real part) point. The

buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter.

imaginaryParts_Q

Variable Type       ViReal64[]

Returns an array of trace data (Q-value y, imaginary part) point. The

buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Offset Samples) out of range.

BFFC0003  Parameter 3 (No of Samples) out of range.

BFFC0004  Parameter 4 (Buffer Size) out of range.

BFFC0005  Parameter 5 (No of Points) is null pointer.

BFFC0006  Parameter 6 (Real Parts (I)) is null pointer.

BFFC0007  Parameter 7 (Imaginary Parts (Q)) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadModRef

ViStatus rsfsu_dataReadModRef (ViSession instrumentHandle,

ViPReal64 measuredLevel,

ViPReal64 correctedLevel,

ViPReal64 bandwidth);

Purpose

This function starts the premeasurement of the modulation due to spectrum

measurement and reads out the result.

Note:

This function is only available with GSM/EDGE Analyzer FS-K5 option and

when measurement modulation due to spectrum is selected.

Remote-control command(s):

*CLS

READ:SPECtrum:MODulation:REFerence:IMMediate?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measuredLevel

Variable Type       ViReal64 (passed by reference)

This control returns the measured level value.

correctedLevel

Variable Type       ViReal64 (passed by reference)

This control returns the corrected level value. This value is

corrected by means of the bandwidth.

bandwidth

Variable Type       ViReal64 (passed by reference)

This control returns the bandwidth value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Measured Level) is null pointer.

BFFC0003  Parameter 3 (Corrected Level) is null pointer.

BFFC0004  Parameter 4 (Bandwidth) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadPMetResult

ViStatus rsfsu_dataReadPMetResult (ViSession instrumentHandle,

ViInt32 window,

ViPReal64 measurementResult);

Purpose

This function triggers a measurement with the power sensor and then

outputs the result.

Note:

(1) This function is available only in conjunction with option FS-K9.

Remote-control command(s):

READ<1|2>:PMETer?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

measurementResult

Variable Type       ViReal64 (passed by reference)

Outputs the result of the power sensor.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Measurement Result) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadPtempRef

ViStatus rsfsu_dataReadPtempRef (ViSession instrumentHandle,

ViPReal64 measuredLevel,

ViPReal64 correctedLevel,

ViPReal64 bandwidth);

Purpose

This function starts the premeasurement of power vs. time and read out

the results.

Note:

This function is only available with GSM/EDGE Analyzer FS-K5 option and

when measurement of the power vs. time is selected.

Remote-control command(s):

*CLS

READ:BURSt:PTEMplate:REFerence:IMMediate?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

measuredLevel

Variable Type       ViReal64 (passed by reference)

This control returns the measured level value.

correctedLevel

Variable Type       ViReal64 (passed by reference)

This control returns the corrected level value. This value is

corrected by means of the bandwidth.

bandwidth

Variable Type       ViReal64 (passed by reference)

This control returns the bandwidth value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Measured Level) is null pointer.

BFFC0003  Parameter 3 (Corrected Level) is null pointer.

BFFC0004  Parameter 4 (Bandwidth) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadSpectrum

ViStatus rsfsu_dataReadSpectrum (ViSession instrumentHandle,

ViBoolean spectrumMeasMode,

ViInt32 _VI_FAR index[],

ViReal64 _VI_FAR startFrequencies[],

ViReal64 _VI_FAR stopFrequencies[],

ViReal64 _VI_FAR levels[],

ViReal64 _VI_FAR limits[],

ViInt32 _VI_FAR absRel[],

ViInt32 _VI_FAR status[],

ViPInt32 num_ofResults);

Purpose

This function provides the commands for starting measurements in the GSM

MS (option FS-K5) Analyzer mode, which are used to measure the power of

the spectral components due to modulation and switching (modulation

spectrum, transient spectrum), and for reading out the results

subsequently.

Notes:

(1) This function is available only in conjunction with option FS-K5, and

when measurement of the modulation spectrum is selected.

(2) An ongoing measurement can be aborted with the command ABORT.

Remote-control command(s):

*CLS

READ:SPECtrum:SWITching?

READ:SPECtrum:MODulation?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

spectrumMeasMode

Variable Type       ViBoolean

This control sets spectrum measurement mode.

Valid Range:

VI_FALSE (0) - Switching

VI_TRUE  (1) - Modulation (Defaul Value)

Notes:

(1) Switching:

This command starts the measurement of the transient spectrum of the

base station or mobile and reads out the result.

(2) Modulation:

This command starts the measurement of the modulation spectrum of the

base station or mobile and reads out the result. The measurement is

performed in the currently set frequency range.

index

Variable Type       ViInt32[]

Returns indexed array of measurement range (last number indicates sum

of measurements).

startFrequencies

Variable Type       ViReal64[]

Returns start frequencies of the measurement range or frequencies

where the limit is exceeded.

stopFrequencies

Variable Type       ViReal64[]

Returns stop frequencies of the measurement range or frequencies

where the measured range is exceeded. This values are equal to the

values of start frequencies, if either the measurement is performed

in the time domain or the result value contains a limit excess.

levels

Variable Type       ViReal64[]

Returns array of measured maximum levels of the partial range or

measured levels at the test point.

limits

Variable Type       ViReal64[]

Returns limits in the partial range or measured levels at the test

point.

absRel

Variable Type       ViInt32[]

Returns array of ABS/REL unit indicators. Meaning of numbers is

following:

ABS = 0

REL = 1

status

Variable Type       ViInt32[]

Limit checks array of results.

PASSED   ... no limits exceeded

FAILED   ... limits exceeded

MARGIN   ... margin exceeded

EXC      ... limited excesses characterized as an exception

Meaning of numbers is:

PASSED = 0

FAILED = 1

MARGIN = 2

EXC = 3

num_ofResults

Variable Type       ViInt32 (passed by reference)

This control returns number of results, received from the instrument.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Spectrum Meas Mode) out of range.

BFFC0003  Parameter 3 (Index) is null pointer.

BFFC0004  Parameter 4 (Start Frequencies) is null pointer.

BFFC0005  Parameter 5 (Stop Frequencies) is null pointer.

BFFC0006  Parameter 6 (Levels) is null pointer.

BFFC0007  Parameter 7 (Limits) is null pointer.

BFFC0008  Parameter 8 (Abs/Rel) is null pointer.

BFFC0009  Parameter 9 (Status) is null pointer.

BFFC000A  Parameter 10 (Num of Results) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadTrace

ViStatus rsfsu_dataReadTrace (ViSession instrumentHandle,

ViInt32 sourceTrace, ViPInt32 no_ofPoints,

ViReal64 _VI_FAR traceData[]);

Purpose

This function reads out trace data from the instrument.

Note:

The associated measurement window is selected with the function Set

Active Window (rsfsu_confSetActiveWindow).

Remote-control command(s):

FORMat ASCII

TRACe<1|2|3>? TRACE1 | TRACE2 | TRACE3

FORMat REAL,32

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

sourceTrace

Variable Type       ViInt32

This control sets the source trace data to be read from instrument.

Valid Range:

0 - Trace 1 (Default Value)

1 - Trace 2

2 - Trace 3

no_ofPoints

Variable Type       ViInt32 (passed by reference)

Returns the number of trace data points.

Note:

If VI_ATTR_INTF_TYPE is equal to VI_INTF_ASRL, then use this

parameter to pass the expected number of points to be read from the

instrument. Data points will be read until either number of data

points are consumed or instruments output buffer is empty. This is

due to limitation if the unit is remote-controlled via the RS-232-C

interface.

traceData

Variable Type       ViReal64[]

Trace data array.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Source Trace) out of range.

BFFC0003  Parameter 3 (No. of Points) is null pointer.

BFFC0004  Parameter 4 (Trace Data) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadTraceC2k

ViStatus rsfsu_dataReadTraceC2k (ViSession instrumentHandle,

ViInt32 sourceTrace,

ViPInt32 no_ofValues,

ViReal64 _VI_FAR traceData[]);

Purpose

This function reads trace data out of the instrument. Only TRACE1 or

TRACE2 can be queried depending on the display mode.

Notes:

(1) This function is available for option FS-K82, FS-K83, FS-K84 or

FS-K85 only.

(2) The associated measurement window is selected with the function Set

Active Window (rsfsu_confSetActiveWindow).

Remote-control command(s):

TRACe[:DATA] TRACE1 | TRACE2

FORMat ASCII

FORMat REAL,32

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

sourceTrace

Variable Type       ViInt32

This parameter indicates which trace will be read out.

Valid Range:

0 - Trace 1 (Default Value)

1 - Trace 2

no_ofValues

Variable Type       ViInt32 (passed by reference)

Returns the number of trace data values.

traceData

Variable Type       ViReal64[]

Returns an array of trace data. Only TRACE1 or TRACE2 can be queried

depending on the display mode. The following measured values are

transferred depending on the display mode.

For more dedails please refer also to the user manual.

---------- Option FS-K82 ----------

CODE DOMAIN POWER ABSOLUTE/CODE DOMAIN POWER RELATIVE (TRACE1):

The following is output for each channel:

Code class:

Code class of the channel. With Hadamard order, normally the

code class corresponding to the base spreading factor. With

BitReverse order, values between 2 and 7.

Code number:

Code number of the channel, values between 0 and 127.

Level:

- For CODE DOMAIN POWER ABSOLUTE in dBm

- For CODE DOMAIN POWER RELATIVE in dB

(referred to the total or pilot power)

In Hadamard order, power values of the different codes are

normally specified, in BitReverse order the consolidated

channel power is returned.

Power detection:

0 - inactive channel

1 - power of own antenna

2 - alias power of own antenna

3 - alias power of other antenna

4 - alias power of own and other antenna.

Alias power can be avoided by analysis when the base

spreading factor is selected correctly.

Thus four values are transferred for all channels:

<Code class>,<Code number>,<Signal level>,<Power ID>,...

For sorting the channels and consolidation, Hadamard order or

BitReverse order is important.

With Hadamard, the different codes are output in ascending order with

their code power. The number of output codes corresponds to the base

spreading factor.

With BitReverse, codes belonging to a channel are next to each other

and are therefore output in the class of the channel with output of

the consolidated channel power. The maximum number of output codes or

channels cannot be higher than the base spreading factor, but

decreases with every concentrated channel.

CODE DOMAIN ERROR POWER (TRACE1):

The following is output for each channel:

Code class:

Code class of the channel normally corresponding to the class

for the CDEP that belongs to the base spreading factor.

Code number:

Code number of the channel, values between 0 and 127.

Error power:

In dB

No difference of power between the Hadamard and BitReverse

orders.

Power ID:

0 - inactive channel

1 - power of own antenna

2 - alias power of own antenna

3 - alias power of other antenna

4 - alias power of own and other antenna.

Alias power can be avoided by analysis when the base

spreading factor is selected correctly.

Thus four values are transferred for all channels:

<Code class>,<code number>,<signal level>,<power ID>,...

For sorting the channels, the Hadamard order or BitReverse order is

important.

In Hadamard order, the different codes are sorted in ascending order

before being output. With BitReverse order, codes belonging to any

one channel are next to each other. Because an error power is output

with the code domain error power, a consolidation of the power values

is not sensible.

The number of output codes therefore generally corresponds to the

base spreading factor.

CHANNEL TABLE (TRACE1):

The following is output for each channel:

Channel type:

The channel type is coded with numbers as follows:

0 = PICH

1 = SYNC

2 = PCH

3 = TDPICH

4 = APICH

5 = ATDPICH

6 = BCH

7 = CPCCH

8 = CACH

9 = CCCH

10 = CHAN

11 = INACTIVE

Code class:

Code class of the channel, values between 2 and 7.

Code number:

Code number of the channel, values between 0 and 127.

Radio configuration:

0 for special channels, 1-5 for CHAN (with CODE CHAN

AUTOSEARCH '1' is returned for RCs 1&2 and '3' for RCs 3 to

5)

Absolute level:

In dBm

Relative level:

In dB, referred to the total or pilot power.

Timing offset:

Referred to the pilot in seconds.

Phase offset:

Referred to the pilot in rad. If the evaluation of the timing

and phase offset is not active or more than 50 active

channels are in the signal, the value 9 is returned. For

inactive channels, the value 9 is returned.

The class specifies the spreading factor:

Class 7 corresponds to the highest spreading factor (128, symbol rate

9.6 ksps), and class 2 to the lowest allowed spreading factor (4,

symbol rate 307.2 ksps).

Thus eight values are transferred for all channels:

<Channel type>,<Code class>,<Code number>,<Radio

configuration>,<Absolute level>,<Relative level>,<Timing

offset>,<Phase offset>,...

All detected special channels are sorted first, followed by the data

channels, in ascending order by code class and within the code class

in ascending order by code number. The unassigned codes appear at the

bottom of the list with the code class of the base spreading factor.

RESULT SUMMARY (TRACE2):

The results of the RESULT SUMMARY are output in the following order:

<SLOT>,<PTOTal>,<PPICh>,<RHO>,<MACCuracy>,<PCDerror>,

<ACTive>,<FERRor>,<FERPpm>,<TFRame>,<CERRor>,<IQOFfset>,

<IQIMbalance>,<SRATe>,<CHANnel>,<SFACtor>,<TOFFset>,

<POFFset>,<CDPRelative>,<CDPabsolute>,<EVMRms>,<EVMPeak>

The results have the following meanings and units:

Global results of the selected PCG:

SLOT        PCG number

PTOTal      Total power in dBm

PPICh       Pilot power in dBm

RHO         RHO

MACCuracy   Composite EVM in %

PCDerror    Peak code domain error in dB

IQOFfset    IQ offset in %

IQIMbalance IQ imbalance in %

Global results of all PCGs:

FERRor      Frequency error in Hz

FERPpm      Frequency error in ppm

CERRor      Chip rate error in ppm

TFRame      Trigger to frame

ACTive      Number of active channels

Channel results:

SRATe       Symbol rate in ksps

CHANnel     Channel number

SFACtor     Spreading factor of the channel

CDPRelative Channel power relative in dB

EVMRms      Error vector magnitude rms in %

TOFFset     Timing Offset in s

POFFset     Phase Offset in rad

CDPabsolute Channel power abs. in dBm

EVMPeak     Error vector mag. peak in %

Hint:

The value trigger to frame (TFRame) delivers a '9' if the trigger is

set to FREE RUN. The values timing/phase offset (TOFFset/POFFset)

deliver a '9' if the timing and phase offset measurement is switched

off (refer to CDP:TPM)or the number of active channel exceeds 50.

POWER VS PCG, PEAK CODE DOMAIN ERR and COMPOSITE EVM (TRACE2):

The number of returned value pairs corresponds to the Capture length.

POWER VS PCG: <Slot number>,<Level value in dB>,<Slot

number>,<Level value in dB>,.....;

PEAK CODE DOMAIN ERROR: <Slot number>,<Level value in dB>,

.....;

COMPOSITE EVM: <Slot number>,<Value in %>, .....;

SYMBOL EVM (TRACE2):

The number of values depends on the spreading factor:

Spreading factor 128:12 values; spreading factor 64:24 values

Spreading factor 32:48 values; spreading factor 16:96 values

Spreading factor 8:192 values; spreading factor 4:384 values

(With transmit diversity on the number of symbols are reduced to the

half.)

<Value in % symbol 0>,<value in % symbol 1>,.....;

POWER VS SYMBOL (TRACE2):

The number of values depends on the spreading factor:

Spreading factor 128:12 values; spreading factor 64:24 values

Spreading factor 32:48 values; spreading factor 16:96 values

Spreading factor 8:192 values; spreading factor 4:384 values

(With transmit diversity on the number of values are reduced to the

half.)

<Value in dBm symbol 0>,<value in dBm symbol 1>,.....;

SYMBOL CONST (TRACE2):

The number of value pairs depends on the spreading factor:

Spreading factor 128:12 pairs; spreading factor 64:24 pairs

Spreading factor 32:48 pairs: spreading factor 16:96 pairs

Spreading factor 8:192 pairs; spreading factor 4:384 pairs

(With transmit diversity on the number of pairs are reduced to the

half.)

Real and imaginary parts are transferred as value pairs.

<re 0>,<im 0>,<re 1>,<im 1>,.....<re n>,<im n>

COMPOSITE CONST (TRACE2):

The number of values corresponds to the chip number of 1536 chips in

a power control group.

<Value in % chip 0>,<value in % chip 1>,.....;

BITSTREAM (TRACE2):

The bitstream of a slot is output. A value is output for each bit

(range 0,1), each symbol consists of 2 bits for QPSK channels and of

one bit for BPSK channels.

The number of values depends on the spreading factor with QPSK:

Spreading factor 128:24 values; spreading factor 64:48 values

Spreading factor 32:96 values; spreading factor 16:192 values

Spreading factor 8:384 values; spreading factor 4:768 values

For BPSK:

Spreading factor 128:12 values; spreading factor 64:24 values

Spreading factor 32:48 values; spreading factor 16:96 values

Spreading factor 8:192 values; spreading factor 4:384 values

(With transmit diversity on the number of values are reduced to the

half.)

Example of bitstream trace: 0, 0, 1, 0, 1, 1, 0 ....

---------- Option FS-K83 ----------

CODE DOMAIN POWER ABSOLUTE/CODE DOMAIN POWER RELATIVE (TRACE1):

The following is output for each channel:

Code class:

Code class of the channel; with the Hadamard order it is the

code class which corresponds to the base spreading factor.

With the BitReverse order, values are between 1..6Code number

Code number of the channel, values are between 0..63.

Level:

- With CODE DOMAIN POWER ABSOLUTE, units are dBm

- With CODE DOMAIN POWER RELATIVE, units are dB

(referred to the total or pilot power, refer to the command

CDPower:PREFerence)

In the Hadamard order, power values of the

individual codes are usually specified; in the BitReverse

order, the consolidated channel power is returned.

Power ID:

0 - Inactive channel

1 - Power of device antenna

2 - Alias power of device antenna

3 - Quasi-inactive channel (on the analyzed branch, the

channel is not assigned, but an active channel exists on

the other branch)

Alias power can be avoided by ensuring that the correct base

spreading factor is selected when performing the analysis.

(Refer to the CDPower:SFACtor command)

Four values are thus transferred for all channels:

<Code class>, <Code number>, <Level>, <Power ID>, ...

The Hadamard or BitReverse order is important for sorting the

channels and consolidation (refer to the command CDPower:ORDer). With

Hadamard, the individual codes are output in ascending order with

their code power. The number of codes which are output corresponds to

the base spreading factor.

With BitReverse, codes which belong to a particular channel are

adjacent to each other and are therefore output in the class of the

channel together with the channel power. The maximum number of codes

or channels that are output cannot be higher than the base spreading

factor, and decreases with each bunched channel.

CODE DOMAIN ERROR POWER (TRACE1):

The following is output for each channel:

Code class:

Code class of the channel, which with the CDEP generally

corresponds to the class which belongs to the base spreading

factor.

Code number:

Code number of the channel, values are between 0..63

Error power:

In dB. No distinction is made for power between the Hadamard

and BitReverse order.

Power ID:

0 - Inactive channel

1 - Power of device antenna

2 - Alias power of device antenna

3 - Quasi-inactive channel (on the analyzed branch, the

channel is not assigned, but an active channel exists on

the other branch)

Alias power can be avoided by ensuring that the correct base

spreading factor is selected when performing the analysis. (Refer to

the CDPower:SFACtor command)

Four values are thus transferred for all channels:

<Code class>, <Code number>, <Level>, <Power ID>, ...

The Hadamard or BitReverse order is important for sorting the

channels (refer to the command CDPower:ORDer). With Hadamard, the

individual codes are output in ascending order. With BitReverse,

codes which belong to a particular channel are adjacent to each

other. Since an error power is output for Code Domain Error Power,

consolidation of the power values is not appropriate. The number of

codes that are output therefore generally corresponds to the base

spreading factor.

CHANNEL TABLE (TRACE1):

The following is output for each channel:

Channel type:

The channel type is numerically coded as follows:

0 = PICH

1 = EACH

2 = CCCH

3 = DCCH

4 = ACKCH

5 = CQICH

6 = FCH

7 = S1CH

8 = S2CH

9 = INACTIVE

Code class:

Code class of the channel, values are between 1..6

Code number:

Code number of the channel, values are between 0..63

Mapping:

0 = I branch

1 = Q branch

Absolute level:

In dBmRelative level In dB, referred to the total or pilot

power (refer to the command CDPower:PREFerence)

Timing offset:

Referred to the pilot in seconds. Phase offset Referred to the

pilot in rad. If the analysis of the timing and phase offset

is not active (refer to CDPower:TPMeas) or there are more than

50 active channels in the signal, a value of 9 is returned. A

value of 9 is generally returned for inactive channels.

The class gives the spreading factor of the channel:

Class 6 corresponds to the highest spreading factor (64,

symbol rate 19.2 ksps), class 1 to the lowest permissible

spreading factor (2, symbol rate 614.4 ksps).

This means that 8 values are transmitted for all active channels:

<Channel type>, <Code class>, <Code number>, <Mapping>,

<Absolute level>, <Relative level>, <Timing offset>, <Phase

offset>, etc.

All detected active channels are output first, followed by the

inactive or quasi-active channels. The channels are sorted in

ascending code number order (with identical code numbers: the I

branch first, followed by the Q branch). The unassigned codes are

displayed together with the code factor of the base spreading factor.

RESULT SUMMARY (TRACE2):

The results of the RESULT SUMMARY are output in the following order:

<SLOT>, <PTOTal>, <PPICh>, <RHO>, <MACCuracy>, <PCDerror>,

<ACTive>, <FERRor>, <FERPpm>, <TFRame>, <CERRor>, <IQOFfset>,

<IQIMbalance>, <SRATe>, <CHANnel>, <SFACtor>, <TOFFset>,

<POFFset>, <CDPRelative>, <CDPabsolute>, <EVMRms>, <EVMPeak>

In this instance the results have the following meaning and units:

Global results of the selected PCG:

SLOT - PCG number

PTOTal - Total power in dBm

PPICh - Pilot power in dBm

RHO - RHO

MACCuracy - Composite EVM in %

PCDerror - Peak code domain error in dB

IQOFfset - IQ offset in %

IQIMbalance - IQ imbalance in %

Global results for all PCGs:

FERRor - Frequency error in Hz

FERPpm - Frequency error in ppm

CERRor - Chip rate error in ppm

TFRame - Trigger to frame

ACTive - Number of active channels

Channel results:

SRATe - Symbol rate in ksps

CHANnel - Channel number

SFACtor - Spreading factor of channel

CDPRelative - Channel power relative in dB (relative to total

or PICH power, refer to command CDP:PREF)

EVMRms - Error vector magnitude RMS in %

TOFFset - Timing offset in s

POFFset - Phase offset in rad

CDPabsolute - Channel power absolute in dBm

EVMPeak - Error vector mag. peak in %

POWER VS PCG, PEAK CODE DOMAIN ERR and COMPOSITE EVM (TRACE2):

The number of returned value pairs corresponds to the IQ capture

length. (Refer to the command CDPower:IQLength).

POWER VS PCG: <Slot number>,<Level value in dB>, <Slot

number>, <Level value in dB>,.....;

PEAK CODE DOMAIN ERROR: <Slot number>, <Level value in dB>,

.....;

COMPOSITE EVM: <Slot number>, <Value in %>, .....;

SYMBOL EVM (TRACE2):

The number of values depends on the spreading factor:

Spreading factor 64 : 24 values Spreading factor 32 : 48 values:

Spreading factor 16 : 96 values Spreading factor 8 : 192 values;

Spreading factor 4 : 384 values Spreading factor 2 : 768 values;

<Value in % symbol 0>, <Value in % symbol 1>,.....;

POWER VS SYMBOL (TRACE2):

The number of values depends on the spreading factor:

Spreading factor 64 : 24 values Spreading factor 32 : 48 values:

Spreading factor 16 : 96 values Spreading factor 8 : 192 values;

Spreading factor 4 : 384 values Spreading factor 2 : 768 values;

<Value in dBm symbol 0>, <Value in dBm symbol 1>,.....;

SYMBOL CONST (TRACE2):

The number of value pairs depends on the spreading factor:

Spreading factor 64 : 24 pairs Spreading factor 32 : 48 pairs:

Spreading factor 16 : 96 pairs Spreading factor 8 : 192 pairs;

Spreading factor 4 : 384 pairs Spreading factor 2 : 768 pairs;

The real and imaginary component are transferred as a value pair.

<re 0>,<im 0>,<re 1>,<im 1>,.....<re n>, <im n>

COMPOSITE CONST (TRACe2):

The number of value pairs corresponds to the number of chips from the

1536 chips in a power control group. Real and imaginary components

are transferred as value pairs.

<re Chip 0>, <in Chip 0>, <re Chip 1>, <in Chip 1>,.....;

BITSTREAM (TRACE2):

The bit stream of a slot is output. A value is read out for each bit

(value range 0,1), each symbol consists of one bit for BPSK channels.

Spreading factor 64 : 24 values Spreading factor 32 : 48 values:

Spreading factor 16 : 96 values Spreading factor 8 : 192 values;

Spreading factor 4 : 384 values Spreading factor 2 : 768 values;

If a channel is detected as being active, the invalid bits in the bit

stream are identified by "9".

Example of bit stream trace: 0, 0, 1, 0, 1, 1, 0 ....

---------- Option FS-K84 ----------

CODE DOMAIN POWER ABSOLUT/CODE DOMAIN POWER RELATIVE (TRACE1/TRACE2):

The following is output for each channel of the selected channel

type:

Code number:

Code number of the channel, values between 0 to 31/15/63

depending on the channel type.

Level:

- For CODE DOMAIN POWER ABSOLUTE in unit dBm

- For CODE DOMAIN POWER RELATIVE in unit dB

Power ID:

0 - inactive channel

1 - active channel

Three values are thus transferred for all channels:

<Code number>, <Level>, <Power ID>, etc.

The number of output codes corresponds to the spreading factor, which

depends on the channel type: PILOT/PREAMBLE: 32, DATA: 16, MAC: 64.

(See table above.) Depending on the mapping settings, the I or Q

branch or the complex signal is returned. With Code Domain Overview,

the I branch in TRACE1 and the Q branch in TRACE2 are always

returned.

CODE DOMAIN ERROR POWER (TRACE1/TRACE2):

The following is output for each channel of the selected channel

type:

Code number:

Code number of the channel, values between 0 to 31/15/63

depending on the channel type.

Error power:

In unit dB

Power ID:

0 - inactive channel

1 - active channel

Three values are thus transferred for all channels:

<Code number>, <Level>, <Power ID>, etc.

The number of output codes corresponds to the spreading factor, which

depends on the channel type: PILOT/PREAMBLE: 32, DATA: 16, MAC: 64.

(See table above.) Depending on the mapping settings, the I or Q

branch or the complex signal is returned. With Code Domain Overview,

the I branch in TRACE1 and the Q branch in TRACE2 are always

returned.

CHANNEL TABLE (TRACE1):

The following is output for each active channel:

Channel type:

The channel type is numerically coded as follows:

0 = PILOT

1 = MAC

2 = PREAMBLE with 64 chip length

3 = PREAMBLE with 128 chip length

4 = PREAMBLE with 256 chip length

5 = PREAMBLE with 512 chip length

6 = PREAMBLE with 1024 chip length

7 = DATA

Code class:

Fixed, depending on channel type: PILOT: 5, MAC: 6, PREAMBLE:

5 and DATA: 4 (spreading factor = 2^code class)

Code number:

Code number of channel, values between 0 and spreading

factor 1.

Modulation:

Modulation type including mapping:

0 = BPSK-I

1 = BPSK-Q

2 = QPSK

3 = 8-PSK

4 = 16-QAM

Modulation types QPSK/8-PSK/16-QAM have complex values.

Absolute level:

In unit dBm

Relative level:

In unit dB, referred to the total power in the channel type.

Timing offset:

Referred to the first active channel in seconds.

Phase offset:

Referred to the first active channel in rad. If the analysis

of the timing and phase offsets is not active (refer to

CDPower:TPMeas) or if there are more than 50 active channels

in the signal, a value of 9 is returned. A value of 9 is

generally returned for inactive channels.

This means that 8 values are transmitted for all active channels:

<Channel type>, <Code class>, <Code number>, <Modulation>,

<Absolute level>, <Relative level>, <Timing offset>, <Phase

offset>, etc.

The channels are listed in the following channel type order: PILOT,

MAC, PREAMBLE, DATA. Within the channel types, the channels are

sorted in ascending code number order.

GENERAL RESULTS/CHANNEL RESULTS (TRACE2):

The results included in GENERAL RESULTS and CHANNEL RESULTS are read

out in the following order:

<FERRor>, <FERPpm>, <CERRor>, <TFRame>, <RHOPilot>, <RHO1>,

<RHO2>, <PPILot>, <PMAC>, <PDATa>, <PPReamble>,

<MACCuracy>,<DMTYpe>, <MACTive>, <DACTive>, <PLENGth>, <RHO>,

<PCDerror> , <IQIMbalance>, <IQOFfset>, <SRATe>, <CHANnel>,

<SFACtor> <TOFFset>, <POFFset>, <CDPRelative>, <CDPabsolute>,

<EVMRms>, <EVMPeak>, <MTYPe>

In this instance the results have the following meaning and units:

Global results for all slots:

FERRor - Frequency error in Hz

FERPpm - Frequency error in ppm

CERRor - Chip rate error in ppm

TFRame - Trigger to Frame

RHOPilot - RHO over all slots for the pilot area

RHO1 - RHOoverall-1 over all slots over all chips with start

of averaging at the half-slot limit

RHO2 - RHOoverall-2 over all slots over all chips with start

of averaging at the quarter-slot limit

Results for the selected slot:

PPILot - Absolute power in the PILOT channel type

PMAC - Absolute power in the MAC channel type

PDATa - Absolute power in the DATA channel type

PPReamble - Absolute power in the PREAMBLE channel type

MACCuracy - Composite EVM in %

DMTYpe - Modulation type in the DATA channel type:

2 = QPSK

3 = 8-PSK

4 = 16-QAM

MACTive - Number of active MAC channels

DACTive - Number of active DATA channels

PLENGth - Length of preamble in chips

RHO - RHO value for the selected channel type/slot

Results for the selected channel type in the selected slot:

PCDerror - Peak code domain error in dB

IQIMbalance - IQ imbalance in %

IQOFfset - IQ offset in %

Channel results:

SRATe - Symbol rate in ksps

CHANnel - Channel number

SFACtor - Spreading factor of the channel

TOFFset - Timing offset in s

POFFset - Phase offset in rad

CDPRelative - Channel power, relative, in dB

CDPabsolute - Channel power, absolute, in dBm

EVMRms - Error vector magnitude RMS in %

EVMPeak - Error vector mag. peak in %

MTYPe - Modulation type:

0 = BPSK-I

1 = BPSK-Q

2 = QPSK

3 = 8-PSK

4 = 16-QAM

5 = 2BPSK (if complex analysis selected for PILOT, PREAMBLE

or MAC)

POWER VS CHIP

The number of returned values is 2048, since one level value is

available per chip.

<Level value in dBm>, <Level value in dBm>, etc.;

PEAK CODE DOMAIN ERR and COMPOSITE EVM (TRACE2):

The number of returned value pairs corresponds to the IQ capture

length. (Refer to the command CDPower:IQLength).

PEAK CODE DOMAIN ERROR:

<Slot number>, <Level value in dB>, .....;

COMPOSITE EVM:

<Slot number>, <Value in %>, .....;

SYMBOL EVM (TRACE2):

The number of values depends on the number of symbols and is between

2 and 100. Refer in this respect to the table above in the column

'Number of symbols per slot and code'.

<Value in % symbol 0>, <Value in % symbol 1>,.....;

POWER VS SYMBOL (TRACE2):

The number of values depends on the number of symbols and is between

2 and 100. Refer in this respect to the table above in the column

'Number of symbols per slot and code'.

<Value in dBm symbol 0>, <Value in dBm symbol 1>,.....;

SYMBOL CONST (TRACE2):

The number of value pairs depends on the number of symbols and is

between 2 and 100. Refer in this respect to the table above in the

column 'Number of symbols per slot and code'. Real and imaginary

parts are transferred as value pairs.

<re 0>,<im 0>,<re 1>,<im 1>,.....<re n>, <im n>

COMPOSITE CONST (TRACe2):

The number of value pairs corresponds to the number of chips for the

channel type. The value is between 64 and 1600 ships per slot. Refer

in this respect to the table above in the column 'Chips per slot'.

<re Chip 0>, <in Chip 0>, <re Chip 1>, <in Chip 1>,.....;

BITSTREAM (TRACE2):

The bit stream of a slot is output. A value is read out for each bit

(range 0,1), each symbol consists of 1 bit for BPSK-I/BPSK-Q, 2 bits

for QPSK and 2BPSK (Complex BPSK-I/BPSK-Q analysis), 3 bits for 8-PSK

and 4 bits for 16-QAM. The bit stream per slot is accordingly of

different lengths. Refer in this respect to the table above in the

column 'Bits per slot and code'. The number of bits is between 2 and

400.

Example of bit stream trace: 0, 0, 1, 0, 1, 1, 0 ....

---------- Option FS-K85 ----------

CODE DOMAIN POWER ABSOLUTE / CODE DOMAIN POWER RELATIVE (TRACE1):

The following is output for each channel:

Code class:

Code class of the channel; with Hadamard order it is usually

code class 4. With BitReverse order, values are between

2 and 4.

Code number:

Code number of channel, values between 0 and 15.

Level:

- For CODE DOMAIN POWER ABSOLUTE, units are dBm

- For CODE DOMAIN POWER RELATIVE, units are dB

(referred to the total or pilot power, see the command

CDPower:PREFerence). Power values of the individual codes

are usually given in Hadamard order; the consolidated

channel power is returned in BitReverse order.

Power ID:

0 - Inactive channel

1 - Active channel

3 - Quasi-inactive channel (on the analyzed branch, the

channel is not occupied, but an active channel exists on

the other branch)

Four values are thus transferred for all channels:

<Code class>, <Code number>, <Level>, <Power ID>

The Hadamard or BitReverse order is important for sorting the

channels and consolidation (see the command CDPower:ORDer). With

Hadamard, the individual codes are output in ascending order with

their code power. The number of codes which are output corresponds to

spreading factor 16. With BitReverse, codes which belong to a

particular channel are adjacent to each other and are therefore

output in the class of the channel together with the channel power.

The maximum number of codes or channels that are output cannot be

higher than spreading factor 16, and decreases with each concentrated

channel.

CODE DOMAIN ERROR POWER (TRACE1):

The following is output for each channel:

Code class:

Code class of the channel is usually 4 since the CDEP is

displayed in base spreading factor 16.

Code number:

Code number of channel, values between 0 and 15.

Error power:

In dB. No difference of power between the Hadamard and

BitReverse order.

Power ID:

0 -   Inactive channel

1 - Active channel

3 - Quasi-inactive channel (on the analyzed branch, the

channel is not occupied, but an active channel exists on

the other branch)

Four values are thus transferred for all channels:

<Code class>, <Code number>, <Level>, <Power ID>

The Hadamard or BitReverse order is important for sorting the

channels (see the CDPower:ORDer command). With Hadamard order, the

individual codes are output in ascending order. With BitReverse

order, codes which belong to a particular channel are adjacent to

each other. Since an error power is output for the code domain error

power, consolidation of the power values is not appropriate. The

number of codes that are output therefore generally corresponds to

base spreading factor 16.

CHANNEL TABLE (TRACE1):

The following is output for each channel:

Channel type:

The channel type is coded by numbers as follows:

0 = PICH

1 = RRI

2 = DATA

3 = ACK

4 = DRC

5 = INACTIVE

Code class:

Code class of channel, values between 2 and 4.

Code number:

Code number of channel, values between 0 and 15.

Mapping:

0 = I branch

1 = Q branch

Absolute level:

In dBm

Relative level:

In dB, referred to the total or pilot power (see the

CDPower:PREFerence command)

Timing offset:

Referred to the pilot in seconds.

Phase offset:

Referred to the pilot in rad. If the evaluation of the timing

and phase offset is not active (see CDPower:TPMeas) or more

than 50 active channels are in the signal, the value 9 is

returned. For inactive channels, the value 9 is usually

returned.

The class specifies the spreading factor of the channel:

Class 4 corresponds to spreading factor 16 (symbol rate 76.8 ksps),

class 2 to the lowest permissible spreading factor 4 (symbol rate

307.2 ksps).

Eight values are thus transferred for all channels:

<Channel type>, <Code class>, <Code number>, <Mapping>,

<Absolute level>, <Relative level>, <Timing offset>, <Phase

offset>

All detected active channels are output first, followed by the

inactive or quasi-active channels. The channels are sorted in

ascending code number order (with identical code numbers: the I

branch first, followed by the Q branch). The unassigned codes are

displayed together with code class 4.

RESULT SUMMARY (TRACE2):

The results of RESULT SUMMARY are output in the following order:

<SLOT>, <PTOTal>, <PPICh>, <PRRI>,  <RHO>, <MACCuracy>,

<PCDerror>, <ACTive>, <FERRor>, <FERPpm>, <DRPich>,

<RHOVerall>, <TFRame>, <CERRor>, <IQOFfset>, <IQIMbalance>,

<SRATe>, <CHANnel>, <SFACtor>, <TOFFset>, <POFFset>,

<CDPRelative>, <CDPabsolute>, <EVMRms>, <EVMPeak>

The results have the following meanings and units:

Global results of selected half slot:

SLOT - Half-slot number

PTOTal - Total power in dBm

PPICh - Pilot power in dBm

PRRI - RRI power in dBm

RHO - RHO

MACCuracy - Composite EVM in %

PCDerror - Peak code domain error in dB

IQOFfset - IQ offset in %

IQIMbalance - IQ imbalance in %

Global results of all half slots:

FERRor - Frequency error in Hz

FERPpm - Frequency error in ppm

DRPich - Delta RRI/PICH in dB

CERRor - Chip rate error in ppm

TFRame - Trigger to frame

ACTive - Number of active channels

RHOverall - RHO over all half slots

Channel results:

SRATe - Symbol rate in ksps

CHANnel - Channel number

SFACtor - Spreading factor of channel

CDPRelative - Channel power relative in dB

(relative to total or PICH power, see command CDP:PREF)

EVMRms - Error vector magnitude RMS in %

TOFFset - Timing offset in s

POFFset - Phase offset in rad

CDPabsolute - Channel power absolute in dBm

EVMPeak - Error vector magnitude peak in %

POWER VS HALFSLOT, PEAK CODE DOMAIN ERR and COMPOSITE EVM (TRACE2):

The number of returned value pairs corresponds to the IQ capture

length. (See command CDPower:IQLength).

POWER VS HALFSLOT: <Half-slot number>,<Level value in dB>,

<Half-slot number>, <Level value in

dB>,.....;

PEAK CODE DOMAIN ERROR: <Half-slot number>, <Level value in

dB>, .....;

COMPOSITE EVM: <Half-slot number>, <Value in %>, .....;

SYMBOL EVM (TRACE2):

The number of values depends on the spreading factor:

Spreading factor 16 : 64 values

Spreading factor 8  : 128 values

Spreading factor 4  : 256 values

<Value in % symbol 0>, <Value in % symbol 1>,.....;

POWER VS SYMBOL (TRACE2):

The number of values depends on the spreading factor:

Spreading factor 16 : 64 values

Spreading factor 8  : 128 values

Spreading factor 4  : 256 values

<Value in dBm symbol 0>, <Value in dBm symbol 1>,.....;

SYMBOL CONST (TRACE2):

The number of value pairs depends on the spreading factor:

Spreading factor 16 : 64  values

Spreading factor 8  : 128 values

Spreading factor 4  : 256 values

Real and imaginary components are transferred as value pairs.

<re 0>,<im 0>,<re 1>,<im 1>,.....<re n>, <im n>

BIT STREAM (TRACE2):

The bit stream of a slot is output. A value is read out for each bit

(value range 0,1); each symbol consists of one bit for BPSK channels.

Spreading factor 16 : 64  values

Spreading factor 8  : 128 values

Spreading factor 4  : 256 values

If a channel is detected as being inactive, the invalid bits in the

bit stream are identified by "9".

Example of a bit stream trace: 0, 0, 1, 0 , 1, 1, 0 ....

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Source Trace) out of range.

BFFC0003  Parameter 3 (No of Values) is null pointer.

BFFC0004  Parameter 4 (Trace Data) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadTraceC2kCDP

ViStatus rsfsu_dataReadTraceC2kCDP (ViSession instrumentHandle,

ViInt32 sourceTrace,

ViPInt32 no_ofValues,

ViReal64 _VI_FAR traceData[]);

Purpose

This function reads trace data out of the instrument. Data can be queried

depending on the display mode.

(1) CTABLE:

In addition to the results of the channel table which are output using

the :TRACe[:DATA]? TRACE1 command, active timing and phase offset

measurement (refer to CDPower:TPMeas) also has the CTABle query command

which displays the maximum values of the TIMING and PHASE OFFSET together

with the associated channel.

Option FS-K82, FS-K83 - the following values are output:

<Max. time offset in s>, <Code number for max. time>, <Code

class for max. time>, <Max. phase offset in rad>, <Code number for max.

phase>, <Code class for max. phase>, <Reserved 1>, ..., <Reserved 6>

Option FS-K84 - the following values are output:

<Max. time offset in s>, <Channel Type>, <Code number for max. time>,

<Code class for max. time>, <Max. phase offset in rad>, <Channe type>,

<Code number for max. phase>, <Code class for max. phase>, <Reserved 1>,

..., <Reserved 4>

Option FS-K85 - the following values are output:

<Max. time offset in s>, <Code number for max. time>, <Code class for

max. time>, <Max. phase offset in rad>, <Code number for max. phase>,

<Code class for max. phase>, <Reserved 1>, ..., <Reserved 6>

Notes:

(1) This function is available for option FS-K82, FS-K83, FS-K84 or

FS-K85 only.

(2) The associated measurement window is selected with the function Set

Active Window (rsfsu_confSetActiveWindow).

Remote-control command(s):

TRACe[:DATA]? CTABle

FORMat REAL,32

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

sourceTrace

Variable Type       ViInt32

This parameter indicates which trace will be read out.

Valid Range:

1 - CTABLE (Default Value)

no_ofValues

Variable Type       ViInt32 (passed by reference)

Returns the number of trace data values.

traceData

Variable Type       ViReal64[]

Returns an array of trace data. Measured values are transferred

depending on the display mode (see function description).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Source Trace) out of range.

BFFC0003  Parameter 3 (No of Values) is null pointer.

BFFC0004  Parameter 4 (Trace Data) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadTraceCDP

ViStatus rsfsu_dataReadTraceCDP (ViSession instrumentHandle,

ViInt32 sourceTrace,

ViPInt32 no_ofValues,

ViReal64 _VI_FAR traceData[]);

Purpose

This function reads trace data out of the instrument. Only TRACE1 or

TRACE2 can be queried depending on the display mode.

Notes:

(1) This function is available for option FS-K76 or FS-K77 only.

(2) The associated measurement window is selected with the function Set

Active Window (rsfsu_confSetActiveWindow).

Remote-control command(s):

TRACe[:DATA] TRACE1 | TRACE2

FORMat ASCII

FORMat REAL,32

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

sourceTrace

Variable Type       ViInt32

This parameter indicates which trace will be read out.

Valid Range:

0 - Trace 1 (Default Value)

1 - Trace 2

no_ofValues

Variable Type       ViInt32 (passed by reference)

Returns the number of trace data values.

traceData

Variable Type       ViReal64[]

Returns an array of trace data. Only TRACE1 or TRACE2 can be queried

depending on the display mode. The following measured values are

transferred depending on the display mode.

For more dedails please refer also to the user manual.

---------- Option FS-K76 ----------

CODE DOMAIN POWER ABSOLUTE / CODE DOMAIN POWER RELATIVE (TRACE1):

The following is output for each channel:

Code class:

Code class of channel, values between 0 and 4

Code number:

Code number of channel, values between 1 and 16

Level:

- For CODE POWER ABS in dBm

- For CODE POWER REL in dB

Power detection:

0 - inactive channel

1 - active channel

Thus four values are transferred for all channels:

<Code class>, <Code number>, <Level>, <Power detection>

A maximum of 16 channels are read out, channels belonging together

are read out as one channel.

CODE DOMAIN ERROR POWER (TRACE1):

The following is output for each channel:

Code class:

Code class of channel, values between 0 and 4.

Code number:

Code number of channel, values between 1 and 16.

Error power:

in dB.

Power detection:

0 - inactive channel

1 - active channel

Thus four values are transferred for all channels:

<Code class>, <Code number>, <Level>, <Power detection>

An error power is output for the CDEP, so consolidation of the power

values is not meaningful. The number of output codes therefore

generally corresponds to the spreading factor 16.

CHANNEL TABLE (TRACE1):

The following is output for each channel:

Channel type:

The channel type is coded by numbers as follows:

0 = inactive

1 = midamble

2 = DPCH

3 = P-CCPCH

4 = S-CCPCH

5 = FPACH

6 = PDSCH

7 = PICH

Code class:

Code class of channel, values between 0 and 4.

Code number:

Code number of channel, values between 1 and 16.

Modulation type:

Modulation type of channel

0 = invalid (for midamble)

1 = QPSK

2 = 8PSK

Absolute level:

In dBm.

Relative level:

In dB.

Midamble shift:

Values between 1 and 16.

deltaMidD1:

Power offset between sum power of channels (belonging to

midamble(k), only data field 1) and midamble(k) Power.

deltaMidD2:

Power offset between sum power of channels (belonging to

midamble(k), only data field 2) and midamble(k) Power.

reserved 1:

Reserved for additions

reserved 2:

Reserved for additions

The class specifies the spreading factor of the channel Class 4 is

the highest spreading factor (16, data rate 17.6 kbps for QPSK, data

rate 26.4 kbps for 8PSK), class 0 is the lowest spreading factor (1,

data rate 281.6 kbps for QPSK, data rate 422.4 kbps for 8PSK).

Thus 11 values are transferred for all channels:

<Channel type>, <Code class>, <Code number>, <Modulation

type>, <Absolute level in dBm>, <Relative level in dB>,

<Midamble shift>, <deltaMidD1>, <deltaMidD2>, <Reserved 1>,

<Reserved 2>

In code sorting (CONF:CDP:CTAB:ORD CODE), first all midambles are

output, then the control channels, and finally the data channels in

ascending order of code number. In midamble sorting

(CONF:CDP:CTAB:ORD MID), first the midamble is output and then its

control and data channels.

RESULT SUMMARY (TRACE2):

This command polls the measured and calculated values of code domain

power analysis. The results are provided for the channel to which the

code selected by the CDPower:CODe command belongs.

Parameters:

Global results of selected slot:

SLOT - Slot number

PDATa - Power data fields in dBm

PD1 - Power data field 25.40 mm dBm

PD2 - Power data field 2 in dBm

PMIDamble - Power midamble in dBm

RHO - RHO

MACCuracy - Composite EVM in %

PCDerror - Peak code domain error in dB

FERRor - Frequency error in Hz

CERRor - Chip rate error in ppm

TFRame - Trigger to frame

IQIMbalance - IQ imbalance in %

IQOFfset - IQ offset in %

ACTive - Number of active channels

Channel results:

SRATe - Data rate in kbps

CHANnel - Channel number

SFACtor - Spreading factor of channel

CDPRelative - Channel power relative in dB

CDPabsolute - Channel power absolute in dBm

EVMRms - Error vector magnitude RMS in %

EVMPeak - Error vector magnitude Peak in %

The results of RESULT SUMMARY are output in the following order:

<SLOT>, <PDATa>, <PD1>, <PD2>, <PMIDamble>, <RHO>,

<MACCuracy>, <PCDerror>, <FERRor>, <CERRor>, <TFRame>,

<IQIMbalance>, <IQOFfset>, <ACTive>, <SRATe>, <CHANnel>,

<SFACtor>, <CDPRelative>, <CDPabsolute>, <EVMRms>,

<EVMPeak>, <Reserved 1>, <Reserved 2>, <Reserved 3>,

<Reserved 4>

The results have the following meanings and units:

Global results of selected slot:

SLOT - Slot number

PDATa - Power data fields in dBm

FERRor - Frequency error in Hz

PD1 - Power data field 25.40 mm dBm

CERRor - Chip rate error in ppm

PD2 - Power data field 2 in dBm

TFRame - Trigger to frame

PMIDamble - Power midamble in dBm

IQIMbalance - IQ imbalance in %

RHO - RHO

IQOFfset - IQ offset in %

MACCuracy - Composite EVM in %

ACTive - Number of active channels

PCDerror - Peak code domain error in dB

Channel results:

SRATe - Data rate in kbps

CHANnel - Channel number

SFACtor - Spreading factor of channel

CDPRelative - Channel power relative in dB

CDPabsolute - Channel power absolute in dBm

EVMRms - Error vector magnitude RMS in %

EVMPeak - Error vector magnitude in %

POWER VS SLOT ABS / REL (TRACE2):

The number of returned value tripels corresponds to the IQ capture

length. (See CDPower:IQLength command, range 2 to 63.)

POWER VS SLOT ABS: <Slot number>,<Level in

dBm>,<Validity>,.....;

POWER VS SLOT REL: <Slot number>,<Level in

dB>,<Validity>,.....;

Validity is coded as follows:

0 = inactive (channel not occupied)

1 = active (channel occupied)

2 = alias (code class of channel < 4, i.e. several channels

belong together)

PK CODE DOM ERR and COMPOSITE EVM (TRACE2):

The number of returned value pairs corresponds to the IQ capture

length. (See CDPower:IQLength command, range 2 to 63.)

PK CODE DOM ERR: <Slot number>, <Level in dB>, .....;

COMPOSITE EVM: <Slot number>, <Value in %>, .....;

SYMBOL EVM (TRACE2):

The number of values depends on the spreading factor:

Spreading factor 16: 44 values

Spreading factor 8: 88 values

Spreading factor 4: 176 values

Spreading factor 2: 352 values

Spreading factor 1: 704 values

<Value in % symbol 0>, <Value in % symbol 1>,.....;

POWER VS SYMBOL (TRACE2):

The number of values depends on the spreading factor:

Spreading factor 16: 44 values

Spreading factor 8: 88 values

Spreading factor 4: 176 values

Spreading factor 2: 352 values

Spreading factor 1: 704 values

<Value in dBm symbol 0>, <Value in dBm symbol 1>,.....;

SYMBOL CONST (TRACE2):

The number of value pairs depends on the spreading factor:

Spreading factor 16: 44 values

Spreading factor 8: 88 values

Spreading factor 4: 176 values

Spreading factor 2: 352 values

Spreading factor 1: 704 values

Real and imaginary parts are transferred as value pairs.

<re 0>,<im 0>,<re 1>,<im 1>,.....<re n>, <im n>

COMPOSITE CONST (TRACe2):

The number of value pairs corresponds to the chip number of 704 data

chips in a slot. Real and imaginary parts are transferred as value

pairs:

<re chip 0>, <im chip 0>, <re chip 1>, <im chip 1>,.....;

BIT STREAM (TRACE2):

The bit stream of a channel is output. A value is output for each bit

(range 0, 1), each symbol consists of two bits for QPSK channels and

three bits for 8PSK channels.

- The number of values depends on the spreading factor in QPSK:

Spreading factor 16: 88 values

Spreading factor 2: 704 values

Spreading factor 8: 176 values

Spreading factor 1: 1408 values

Spreading factor 4: 352 values

- In 8PSK:

Spreading factor 16: 132 values

Spreading factor 2: 1056 values

Spreading factor 8: 264 values

Spreading factor 1: 2112 values

Spreading factor 4: 528 values

---------- Option FS-K77 ----------

CODE DOMAIN POWER ABSOLUTE / CODE DOMAIN POWER RELATIVE (TRACE1):

The following is output for each channel:

Code class:

Code class of channel, values between 0 and 4

Code number:

Code number of channel, values between 1 and 16

Level:

- For CODE POWER ABS in dBm

- For CODE POWER REL in dB

Power detection:

0 - inactive channel

1 - active channel

Thus four values are transferred for all channels:

<Code class>, <Code number>, <Level>, <Power detection>

A maximum of 16 channels are read out, channels belonging together

are read out as one channel.

CODE DOMAIN ERROR POWER (TRACE1):

The following is output for each channel:

Code class:

Code class of channel, values between 0 and 4.

Code number:

Code number of channel, values between 1 and 16.

Error power:

in dB.

Power detection:

0 - inactive channel

1 - active channel

Thus four values are transferred for all channels:

<Code class>, <Code number>, <Level>, <Power detection>

An error power is output for the CDEP, so consolidation of the power

values is not meaningful. The number of output codes therefore

generally corresponds to the spreading factor 16.

CHANNEL TABLE (TRACE1):

The following is output for each channel:

Channel type:

The channel type is coded by numbers as follows:

0 = inactive

1 = midamble

2 = DPCH

3 = PRACH

4 = PUSCH

Code class:

Code class of channel, values between 0 and 4.

Code number:

Code number of channel, values between 1 and 16.

Modulation type:

Modulation type of channel

0 = invalid (for midamble)

1 = QPSK

2 = 8PSK

Absolute level:

In dBm.

Relative level:

In dB.

Midamble shift:

Values between 1 and 16.

deltaMidD1:

Power offset between sum power of channels (belonging to

midamble(k), only data field 1) and midamble(k) Power.

deltaMidD2:

Power offset between sum power of channels (belonging to

midamble(k), only data field 2) and midamble(k) Power.

reserved 1:

Reserved for additions

reserved 2:

Reserved for additions

The class specifies the spreading factor of the channel Class 4 is

the highest spreading factor (16, data rate 17.6 kbps for QPSK, data

rate 26.4 kbps for 8PSK), class 0 is the lowest spreading factor (1,

data rate 281.6 kbps for QPSK, data rate 422.4 kbps for 8PSK).

Thus 11 values are transferred for all channels:

<Channel type>, <Code class>, <Code number>, <Modulation

type>, <Absolute level in dBm>, <Relative level in dB>,

<Midamble shift>, <deltaMidD1>, <deltaMidD2>, <Reserved 1>,

<Reserved 2>

In code sorting (CONF:CDP:CTAB:ORD CODE), first all midambles are

output, then the control channels, and finally the data channels in

ascending order of code number. In midamble sorting

(CONF:CDP:CTAB:ORD MID), first the midamble is output and then its

control and data channels.

RESULT SUMMARY (TRACE2):

This command polls the measured and calculated values of code domain

power analysis. The results are provided for the channel to which the

code selected by the CDPower:CODe command belongs.

Parameters:

Global results of selected slot:

SLOT - Slot number

PDATa - Power data fields in dBm

PD1 - Power data field 25.40 mm dBm

PD2 - Power data field 2 in dBm

PMIDamble - Power midamble in dBm

RHO - RHO

MACCuracy - Composite EVM in %

PCDerror - Peak code domain error in dB

FERRor - Frequency error in Hz

CERRor - Chip rate error in ppm

TFRame - Trigger to frame

IQIMbalance - IQ imbalance in %

IQOFfset - IQ offset in %

ACTive - Number of active channels

Channel results:

SRATe - Data rate in kbps

CHANnel - Channel number

SFACtor - Spreading factor of channel

CDPRelative - Channel power relative in dB

CDPabsolute - Channel power absolute in dBm

EVMRms - Error vector magnitude RMS in %

EVMPeak - Error vector magnitude Peak in %

The results of RESULT SUMMARY are output in the following order:

<SLOT>, <PDATa>, <PD1>, <PD2>, <PMIDamble>, <RHO>,

<MACCuracy>, <PCDerror>, <FERRor>, <CERRor>, <TFRame>,

<IQIMbalance>, <IQOFfset>, <ACTive>, <SRATe>, <CHANnel>,

<SFACtor>, <CDPRelative>, <CDPabsolute>, <EVMRms>,

<EVMPeak>, <Reserved 1>, <Reserved 2>, <Reserved 3>,

<Reserved 4>

The results have the following meanings and units:

Global results of selected slot:

SLOT - Slot number

PDATa - Power data fields in dBm

FERRor - Frequency error in Hz

PD1 - Power data field 25.40 mm dBm

CERRor - Chip rate error in ppm

PD2 - Power data field 2 in dBm

TFRame - Trigger to frame

PMIDamble - Power midamble in dBm

IQIMbalance - IQ imbalance in %

RHO - RHO

IQOFfset - IQ offset in %

MACCuracy - Composite EVM in %

ACTive - Number of active channels

PCDerror - Peak code domain error in dB

Channel results:

SRATe - Data rate in kbps

CHANnel - Channel number

SFACtor - Spreading factor of channel

CDPRelative - Channel power relative in dB

CDPabsolute - Channel power absolute in dBm

EVMRms - Error vector magnitude RMS in %

EVMPeak - Error vector magnitude in %

POWER VS SLOT ABS / REL (TRACE2):

The number of returned value tripels corresponds to the IQ capture

length. (See CDPower:IQLength command, range 2 to 63.)

POWER VS SLOT ABS: <Slot number>,<Level in

dBm>,<Validity>,.....;

POWER VS SLOT REL: <Slot number>,<Level in

dB>,<Validity>,.....;

Validity is coded as follows:

0 = inactive (channel not occupied)

1 = active (channel occupied)

2 = alias (code class of channel < 4, i.e. several channels

belong together)

PK CODE DOM ERR and COMPOSITE EVM (TRACE2):

The number of returned value pairs corresponds to the IQ capture

length. (See CDPower:IQLength command, range 2 to 63.)

PK CODE DOM ERR: <Slot number>, <Level in dB>, .....;

COMPOSITE EVM: <Slot number>, <Value in %>, .....;

SYMBOL EVM (TRACE2):

The number of values depends on the spreading factor:

Spreading factor 16: 44 values

Spreading factor 8: 88 values

Spreading factor 4: 176 values

Spreading factor 2: 352 values

Spreading factor 1: 704 values

<Value in % symbol 0>, <Value in % symbol 1>,.....;

POWER VS SYMBOL (TRACE2):

The number of values depends on the spreading factor:

Spreading factor 16: 44 values

Spreading factor 8: 88 values

Spreading factor 4: 176 values

Spreading factor 2: 352 values

Spreading factor 1: 704 values

<Value in dBm symbol 0>, <Value in dBm symbol 1>,.....;

SYMBOL CONST (TRACE2):

The number of value pairs depends on the spreading factor:

Spreading factor 16: 44 values

Spreading factor 8: 88 values

Spreading factor 4: 176 values

Spreading factor 2: 352 values

Spreading factor 1: 704 values

Real and imaginary parts are transferred as value pairs.

<re 0>,<im 0>,<re 1>,<im 1>,.....<re n>, <im n>

COMPOSITE CONST (TRACe2):

The number of value pairs corresponds to the chip number of 704 data

chips in a slot. Real and imaginary parts are transferred as value

pairs:

<re chip 0>, <im chip 0>, <re chip 1>, <im chip 1>,.....;

BIT STREAM (TRACE2):

The bit stream of a channel is output. A value is output for each bit

(range 0, 1), each symbol consists of two bits for QPSK channels and

three bits for 8PSK channels.

- The number of values depends on the spreading factor in QPSK:

Spreading factor 16: 88 values

Spreading factor 2: 704 values

Spreading factor 8: 176 values

Spreading factor 1: 1408 values

Spreading factor 4: 352 values

- In 8PSK:

Spreading factor 16: 132 values

Spreading factor 2: 1056 values

Spreading factor 8: 264 values

Spreading factor 1: 2112 values

Spreading factor 4: 528 values

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Source Trace) out of range.

BFFC0003  Parameter 3 (No of Values) is null pointer.

BFFC0004  Parameter 4 (Trace Data) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadTraceIQ

ViStatus rsfsu_dataReadTraceIQ (ViSession instrumentHandle,

ViInt32 bufferSize,

ViPInt32 no_ofPoints,

ViReal64 _VI_FAR realParts_I[],

ViReal64 _VI_FAR imaginaryParts_Q[],

ViPReal64 meanSignalPower,

ViPReal64 peakSignalPower);

Purpose

This function will start a measurement and read I/Q trace data results.

The result values are scaled linear in unit Volt and correspond to the

voltage at the RF input of the instrument.

Notes:

(1) For large I/Q trace data is recommended to increase default timeout

value to the appropriate value which allows instrument handle all pending

tasks. To change default timeout value call following VISA function:

viSetAttribute (instrSession, VI_ATTR_TMO_VALUE, 60000);

This attribute specifies the minimum timeout value to use (in

milliseconds) when accessing the device associated with the given

session.

(2) If there are not I/Q data available in memory because the

corresponding measurement had not been started, the function will cause a

Query Error.

Example:

(1) Switch on I/Q data acquisition:

rsfsu_confTraceIQState (instrSession, VI_TRUE);

(2) Choose configuration:

rsfsu_confTraceIQ (instrSession, 32.0e6, 0, VI_TRUE,

1.0e6, 0, 1024);

(3) Start measurement and read results:

rsfsu_dataReadTraceIQ (instrSession, bufferSize, &points,

I, Q);

(4) Switch off I/Q data acquisition:

rsfsu_confTraceIQState (instrSession, VI_FALSE)

Remote-control command(s):

FORMat REAL,32

TRACe<1|2>:IQ:DATA:FORMat COMP

TRACe:IQ:DATA?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

bufferSize

Variable Type       ViInt32

Pass the number of doubles in the array you specify for the real

parts and imaginary parameters.

The buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter. If the data contains more

elemens than you specify with the Buffer Size parameter, the function

returns Buffer Size elements into the arrays and returns the buffer

size you must pass to get the all data.  For example, if the data

contains 16000 elements for real parts and 16000 elements for

imaginary part and the Buffer Size is 10000, the function places

10000 elements into the arrays and returns 16000.

Notes:

(1) If VI_ATTR_INTF_TYPE is equal to VI_INTF_ASRL, then use this

parameter to pass the expected number of points to be read from the

instrument. Data points will be read until instruments output buffer

is empty. If Buffer Size is lower than actual number of results

available in the instruments output buffer, then this function

returns an error code (BFFC0002  Parameter 2 (Buffer Size) out of

range.) and measurement results will not be valid. This is due to

limitation if the unit is remote-controlled via the RS-232-C

interface.

(2) This function requires that all response data are read out

completely before the instrument accepts further commands.

no_ofPoints

Variable Type       ViInt32 (passed by reference)

Returns the number of I/Q trace data points.

realParts_I

Variable Type       ViReal64[]

Returns an array of trace data (I-value x, real part) point. The

buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter.

imaginaryParts_Q

Variable Type       ViReal64[]

Returns an array of trace data (Q-value y, imaginary part) point. The

buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter.

meanSignalPower

Variable Type       ViReal64 (passed by reference)

Returns a mean signal power. This parameter will be ignored. You can

pass NULL pointer for this parameter.

peakSignalPower

Variable Type       ViReal64 (passed by reference)

Returns a peak signal power. This parameter will be ignored. You can

pass NULL pointer for this parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Buffer Size) out of range.

BFFC0003  Parameter 3 (No of Points) is null pointer.

BFFC0004  Parameter 4 (Real Parts (I)) is null pointer.

BFFC0005  Parameter 5 (Imaginary Parts (Q)) is null pointer.

BFFC0006  Parameter 6 (Mean Signal Power) is null pointer.

BFFC0007  Parameter 7 (Peak Signal Power) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadTraceIQFromMemory

ViStatus rsfsu_dataReadTraceIQFromMemory (ViSession instrumentHandle,

ViInt32 offsetSamples,

ViInt32 numberOfSamples,

ViPInt32 no_ofPoints,

ViReal64 _VI_FAR realParts_I[],

ViReal64 _VI_FAR imaginaryParts_Q[]);

Purpose

This function permits the readout of previously acquired (and frequency

response corrected) I/Q data from the memory, with indication of the

offset related to the start of measurement and with indication of the

number of measurement values. Therefore a previously acquired data set

can be read out in smaller portions. The maximum amount of available data

depends on the settings of command TRACe:IQ:SET (Configure IQ Data).

Note:

This function is placed here only for compatibility reason. See function

Read Memory I/Q Data (rsfsu_dataReadMemoryIQ) for more information.

Remote-control command(s):

FORMat REAL,32

TRACe:IQ:DATA:MEMory?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

offsetSamples

Variable Type       ViInt32

Determines the offset of the values to be read related to the start

of the acquired data.

Valid Range:

0 to number of measured samples - 1

Default Value:

0

numberOfSamples

Variable Type       ViInt32

Sets the number of samples, that should be read.

Valid Range:

1 to number of measured samples - offset samples

Default Value:

1

no_ofPoints

Variable Type       ViInt32 (passed by reference)

Returns the number of IQ data points.

realParts_I

Variable Type       ViReal64[]

Returns an array of trace data (I-value x, real part) point. The

buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter.

imaginaryParts_Q

Variable Type       ViReal64[]

Returns an array of trace data (Q-value y, imaginary part) point. The

buffer must contain at least as many elements as the value you

specify with the Buffer Size parameter.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Offset Samples) out of range.

BFFC0003  Parameter 3 (Number Of Samples) out of range.

BFFC0004  Parameter 4 (No of Points) is null pointer.

BFFC0005  Parameter 5 (Real Parts (I)) is null pointer.

BFFC0006  Parameter 6 (Imaginary Parts (Q)) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadTraceWCDMA

ViStatus rsfsu_dataReadTraceWCDMA (ViSession instrumentHandle,

ViInt32 sourceTrace,

ViPInt32 no_ofValues,

ViReal64 _VI_FAR traceData[]);

Purpose

This function reads trace data out of the instrument. Only TRACE1 or

TRACE2 can be queried depending on the display mode (CODE PWR ABSOLUTE /

RELATIVE , CHANNEL TABLE (TRACE1), RESULT SUMMARY (TRACE2), POWER VS SLOT

(TRACE2), SYMBOL EVM (TRACE2), PK CODE DOMAIN ERR and MODULATION ACCURACY

(TRACE2), SYMBOL CONST (TRACE2), BITSTREAM (TRACE2) and CODE DOMAIN ERROR

POWER (TRACE1 [Overview OFF] TRACE1 / TRACE2 [Overview ON])).

Notes:

(1) This function is available for option FS-K72, FS-K73 or FS-K74 only.

(2) The associated measurement window is selected with the function Set

Active Window (rsfsu_confSetActiveWindow).

Remote-control command(s):

TRACe<1|2>? TRACE1 | TRACE2 | TRACE3

FORMat ASCII

FORMat REAL,32

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

sourceTrace

Variable Type       ViInt32

This parameter indicates which trace will be read out.

Valid Range:

0 - Trace 1 (Default Value)

1 - Trace 2

no_ofValues

Variable Type       ViInt32 (passed by reference)

Returns the number of trace data values.

traceData

Variable Type       ViReal64[]

Returns an array of trace data. Only TRACE1 or TRACE2 can be queried

depending on the display mode. The following measured values are

transferred depending on the display mode:

For more dedails please refer also to the user manual.

---------- Option FS-K72, FS-K74 ----------

CODE PWR ABSOLUTE / RELATIVE , CHANNEL TABLE (TRACE1):

Each channel is defined by the class, the channel number, the

absolute level, the relative level and the timing offset. The

class denotes the spreading factor of the channel.

Class 9 corresponds to the highest spreading factor (512,

symbol rate 7.5 ksps), class 2 to the lowest admissible sp

reading factor (4, symbol rate 960 ksps).

Five values are transmitted for each channel:

<class>,<channel number>,<absolute level>,<relative

level>,<timing offset>, ...

For CODE PWR ABSOLUTE / RELATIVE, the channels are output in

ascending order sorted according to the code numbers, i.e. in

the same sequence as they are displayed on the screen. For

CHANNEL TABLE, the channels are sorted according to the code

classes, i.e. the unassigned channels are transmitted last.

The units are:

Absolute level - dBm

Relative level - dB referred to CPICH or total power

Timing offset  - chips

The example shows the results of a query for three channels

with the following configuration:

1st channel: spreading factor 512, channel number 7, timing

offset 0

2nd channel: spreading factor 4, channel number 1, timing

offset 256 chips

3rd channel: spreading factor 128, channel number 255, timing

offset 2560 hips

This yields the following result:

9,7,-40,-20,0,2,1,-40,-20,256,7,255,-40,-20,2560

The channels come in the same order as in the CDP diagram,

i.e. depending on their position in the code domain of

spreading factor 512.

CODE DOMAIN ERROR POWER (TRACE1):

Output: Five values are transmitted for each code class 9

channel. The channels are sorted according to the code

numbers.

Format: <code class>1, <code number>1, <CDEP>1, <channel

flag>1, <code class>2, <code number>2, <CDEP>2,

<channel flag>2 ,..., <code class>512, <code

number>512, <CDEP>512, <channel flag>512

Unit:     <[1]>, <[1]>, <[dB]>,<[1]>

Range:    <8> , <0...511>, <-Inf...Inf>, <0;1>

Quantity: 512

Explanation:

code class: [1] Highest code class of an uplink signal. It is

always set to 9 (CC9)

code number: [1] Code number of the evaluated CC9 channel

CDEP: [dB] Code domain error power value of the CC9 channel

channel flag: [1] Denotes if the CC9 channel belongs to an

assigned code channel or not

Range: 0b00 0d0 - CC9 is inactive

0b01 0d1 - CC9 channel belongs to an active

code channel

0b11 0d3 - CC9 channel belongs to an active

code channel with incorrect sent pilot

symbols

POWER VS SLOT (TRACE2):

16 pairs of slot (slot number of CPICH) and level values (for

16 slots) are always transferred.

<slot number>, <level value in dB>,<slot number>,<level value

in dB>,...

RESULT SUMMARY (TRACE2):

The results of the RESULT SUMMARY are output in the following

order:

<composite EVM>,<peak CDE>,<carr freq Error>,<chip rate

error>,<total power>,<trg to frame>,<EVM peak channel>,<EVM

mean channel>, <class>,<channel number>,<power abs.

channel>,<power rel. channel>,<timing offset>,<I/Q

offset>,<I/Q imbalance>

The units are:

EVM peak channel/mean channel, composite EVM, I/Q

offset/imbalance: %

Peak CDE, total power and power abs. channel: dB

Power rel. Channel: dB referred to CPICH or total power

Carr freq error : Hz

Chip Rate Error: ppm

Timing offset: chips

Trg to Frame: us

COMPOSITE EVM PEAK and CODE DOMAIN ERR (TRACE2):

15 pairs of slot (slot number of CPICH) and values are always

transferred.

COMPOSITE EVM: <slot number>, <value in %>, ...

PEAK CODE DOMAIN ERR: <slot number>, <level value in dB>, ...

COMPOSITE CONSTELLATION:

The real and the imaginary part of the chip constellation at

the selected slot are transferred. The values are normalised

to the square root of the average power at the selected slot

Output: List of I/Q values of all chips per slot

Format: Re1,Im1,Re2,Im2,...,Re2560,Im2560

Unit: [1]

Quantity: 2560

POWER VERSUS SYMBOL:

The power of each symbol at the selected slot are transfer

red. The values give the difference to the reference power in

dB. The number of the symbols depends on the spreading factor

of the selected channel.

Output: List of symbol power related to the reference power

Format: Val1,Val2,...,Val(NOF)

Unit: [dB]

Quantity: NOFsymbols = 10*2^(8 - Code Class)

SYMBOL CONSTALLATION(TRACE2):

The real and the imaginary part are transferred as a pair:

<re 0>,<im 0>,<re 1>,<im 1>,...<re n>,<im n>

The number of level values depends on the spreading factor:

Spreading factor 512 5 values

Spreading factor 256 10 values

Spreading factor 128 20 values

Spreading factor 64  40 values

Spreading factor 32  80 values

Spreading factor 16  160 values

Spreading factor 8   320 values

Spreading factor 4   640 value

SYMBOL EVM (TRACE2):

The number of level values depends on the spreading factor:

Spreading factor 512 5 values

Spreading factor 256 10 values

Spreading factor 128 20 values

Spreading factor 64  40 values

Spreading factor 32  80 values

Spreading factor 16  160 values

Spreading factor 8   320 values

Spreading factor 4   640 value

BITSTREAM (TRACE2):

The bit stream of one slot is transferred. One value is

transferred per bit (range 0,1,). Each symbol contains of two

consecutive bits. The number of symbols is not constant and

may vary for each sweep. Specific symbols in the bit stream

may be invalid depending on the channel type and the bit rate

symbols without power). The assigned invalid bits are marked

by "9". The number of Bits of an 16QAM modulated channel is

twice of that of an QPSK modulated channel.

Example for a bit stream trace: 0,1,0,9,0,1,...

---------- Option FS-K73 ----------

CODE PWR ABSOLUTE / RELATIVE, CHANNEL TABLE (TRACE1):

Each channel is defined by the class, the channel number, the

absolute level, the relative level and the timing offset. The

class denotes the spreading factor of the channel.

Class 8 corresponds to the highest spreading factor (256,

symbol rate 15 ksps), class 2 to the lowest admissible

spreading factor (4, symbol rate 960 ksps).

Five values are transmitted for each channel:

<class>,<cannel number>,<absolute level>,<relative level>,<I/Q

component>,...

For CODE PWR ABSOLUTE / RELATIVE, the channels are output in

ascending order sorted according to the code numbers, i.e in

the same sequence as they are displayed on the screen. For

CHANNEL TABLE, the channels are sorted according to the code

classes, i.e. the unassigned channels are transmitted last.

The units are:

Absolute level - dBm,

Relative level - dB referred to the total power of the signal

The example shows the results of a query for three channels

with the following configuration:

1st channel: spreading factor 256, ch number 0, component Q

2nd channel: spreading factor 4, ch number 1, component I

3rd channel: spreading factor 4, ch number 1, component Q

This yields the following result:

8,0,-20.0,0.0,0,2,-20.0,0.0,1,2,1,-20.0,0.0,0

The channels come in the same order as in the CDP diagram,

i.e. depending on their position in the code domain of

spreading factor 256.

CODE DOMAIN ERROR POWER (TRACE1 [Overview OFF] TRACE1 / TRACE2

[Overview ON]):

Output: Five values are transmitted for each code class 8

channel. The channels are sorted according to the code

numbers.

Format: <code class>1, <code number>1, <CDEP>1, <channel

flag>1, <code class>2, <code number>2, <CDEP>2,

<channel flag>2, ,..., <code class>256, <code

number>256, <CDEP>256, <channel flag>256

Unit: <[1]>,<[1]>,<[dB]>,<[1]>

Range: <8>,<0...256>,<-Inf...Inf>,<0;1>

Quantity: 256

Explanation:

code class: [1] Highest code class of an uplink signal. It

is always set to 8 (CC8).

code number: [1] Code number of the evaluated CC8 channel.

CDEP: [dB] Code domain error power value of the CC8 channel.

channel flag: [1] Indicates whether the CC8 channel belongs

to an assigned code channel or not.

Range: 0b00 0d0 - CC8 is inactive

0b01 0d1 - CC8 channel belongs to an active code

channel

RESULT SUMMARY (TRACE2):

The results of the RESULT SUMMARY are output in the following

order:

<composite EVM>,<peak CDE>,<carr freq Error>,<chip rate

error>,<total power>,<trg to frame>,<EVM peak channel>,<EVM

mean channel>, <class>,<channel number>,<power abs.

channel>,<power rel. channel>,<I/Q component>,<pilotm

length>,<IQ offset>,<IQ imbalance>

The units are:

EVM peak channel/mean channel, composite EVM, IQ

offset/imbalance: %

Peak CDE, total power and power abs. channel dB.

Power rel. Channel: dB referred to the total power of the

signal.

Carr freq error: Hz

Chip Rate Error: ppm

Trg to Frame: us

Pilot length: bits

I/Q component: absolute

POWER VS SLOT (TRACE2):

15 pairs of slot (slotnumber of CPICH) and level values (for

15 slots) are always transferred.

<slot number>, <level value in dB>,<slot number>,<level value

in dB>,...

SYMBOL EVM (TRACE2):

The number of level values depends on the spreading factor:

Spreading factor 256 10 values

Spreading factor 128 20 values

Spreading factor 64  40 values

Spreading factor 32  80 values

Spreading factor 16  160 values

Spreading factor 8   320 values

Spreading factor 4   640 values

PEAK CODE DOMAIN ERR and COMPOSITE EVM (TRACE2):

15 pairs of slot (slotnumber of CPICH) and values are always

transferred.

PEAK CODE DOMAIN ERR: <slot number>, <level value in dB>,...

COMPOSITE EVM: <slot number>, <lvalue in %>,...

SYMBOL CONST(TRACE2):

The real and the imaginary part are transferred as a pair:

<re 0>,<im 0>,<re 1>,<im 1>,...<re n>,<im n>

For the channels have exclusively I or Q components in FS-K73,

the <re> or <im> values are 0, depending on the selected

component.

The number of level values depends on the spreading factor:

Spreading factor 256 10 values

Spreading factor 128 20 values

Spreading factor 64  40 values

Spreading factor 32  80 values

Spreading factor 16  160 values

Spreading factor 8   320 values

Spreading factor 4   640 values

BITSTREAM (TRACE2):

The bitstream of one slot is transferred. One value is

transferred per bit (range 0,1,). The number of symbols is not

constant and may vary for each sweep. Specific symbols in the

bitstream may be invalid depending on the channel type and the

bit rate (symbols without power). The assigned invalid bits

are marked by "9".

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Source Trace) out of range.

BFFC0003  Parameter 3 (No of Values) is null pointer.

BFFC0004  Parameter 4 (Trace Data) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataReadTraceWCDP

ViStatus rsfsu_dataReadTraceWCDP (ViSession instrumentHandle,

ViInt32 sourceTrace,

ViPInt32 no_ofValues,

ViReal64 _VI_FAR traceData[]);

Purpose

This function reads trace data out of the instrument. Data can be queried

depending on the display mode.

(1) --- PWCDP ---

PWCDP can be set only for base station tests (FS-K72/74), if CODE PWR

ABSOLUTE / RELATIVE, CHANNEL TABLE is selected for Trace 1. The pilot

length is transmitted in addition to the same five values as transmitted

for TRACE1. The pilot length is specified in symbols.

Six values are transmitted for each assigned channel:

<class>,<channel number>,<absolute level>,<relative

level>,<timing offset>,<pilot length>,...

The pilot length is specified in bit.

(2) --- CTABLE ---

CTABLE can be set only if CODE PWR ABSOLUTE / RELATIVE, CHANNEL TABLE is

selected for Trace 1. The same data as for TRACE1 are output. In

addition, the pilot length is output as the 6th value for FS-K72/74 (0 as

the 6th value for FS-K73) and active / inactive (1/0) as the 7th value

for both options.

FS-K72/74: Seven values are transmitted for each assigned

channel: <class>,<cannel number>,<absolute

level>,<relative level>,<timing offset>,<pilot

length>,<active|inactive>...

FS-K73: Seven values are transmitted for each channel, the

sixth value (reserved for pilot length) being

constantly 0:

<class>,<cannel number>,<absolute level>,<relative

level>,<I/Q component>,0,<state>...

The pilot length is specified in symbols.

(3) --- CWCDP ---

CWCDp can be set if CODE PWR ABSOLUTE / RELATIVE, CHANNEL TABLE

is selected for Trace 1. The pilot length, channel state, channel type,

modulation type and a reserved value are transmitted in addition to the

values as transmitted for TRACE1.

For each channel 10 values are transmitted:

<code class>,<cannel number>,<absolute level>,<relative

level>,<timing offset>,<pilot length>,<active flag>,<channel

type>,<modulation type>,<reserved>...

For CODE PWR ABSOLUTE / RELATIVE, the channels are output in ascending

order sorted according to the code numbers, i.e. in the same sequence as

they are displayed on the screen. For CHANNEL TABLE, the channels are

sorted according to the code classes, i.e. the unassigned channels are

transmitted last.

Parameter number, Parameter range, unit, explanation:

1) <code class>, {2...9}, [1], Code class of the channel

2) <cannel number>, {0...511}, [1], Code number of the channel

3) <absolute level>, {-Inf...Inf}, [dBm], Absolute level of the

code channel at the selected channel slot. (The channel slot

can be marked by the SELECTED CPICH slot.)

4) <relative level>, {-Inf...Inf}, [dB], Relative level of the

code channel at the selected channel slot referred to CPICH

or total power. (The channel slot can be marked by the

SELECTED CPICH slot.)

5) <timing offset>, {0...38400}, [chips], Timing offset of the

code channel to the frame start. The value is measured in

chips. The step width is 256 chips in case of code class 2 to

8, and 512 chips in case of code class 9.

6) <pilot length>, {0,2,4,8,16}, [symbols], Pilot length of the

code channel. According to the 3GPP standardization, the

pilot length range depends on the code class.

7) <active flag>, {0,1}, [1], Flag to characterize if a channel

is active or not:

0 - channel is not active

1 - channel is active

8) <channel type>, {0...13}, [1], Channel type characterisation:

0 - DPCH Dedicated Physical Channel of a standard Frame

1 - PICH Paging Indication Channel

2 - CPICH Common Pilot Channel

3 - PSCH Primary Synchronisation Channel

4 - SSCH Secondary Synchronisation Channel

5 - PCCPCH Primary Common Control Physical Channel

6 - SCCPCH Secondary Common Control Physical Channel

7 - HS_SCCH HSDPA: High Speed Shared Control Channel

8 - HS_PDSCH HSDPA: High Speed Physical Downlink Shared

Channel

9 - CHAN Channel without any pilot symbols (QPSK modulated)

10 - CPRSD Dedicated Physical Channel in compressed mode

11 - CPR-TPC Dedicated Physical Channel in compressed mode

TPC symbols are sent in the first slot of the gap.

12 - CPR-SF/2 Dedicated Physical Channel in compressed mode

using half spreading factor (SF/2).

13 - CPR-SF/2-TPC Dedicated Physical Channel in compressed

mode using half spreading factor (SF/2).

TPC symbols are sent in the first slot of the gap.

9) <modulation type>, {2,4,15}, [1], Modulation type of the code

channel at the selected channel slot. (The channel slot can

be marked by adjusting SELECT CPICH slot.)

2 - QPSK modulation type QPSK

4 - 16QAM modulation type 16QAM

15 - NONE there is no power in the selected channel slot.

[slot is switched OFF.] (According to 3GPP the power of

a HSDPA channel can be switched every 2ms, which are

3 slots.)

10) <reserved>, {0}, [1], reserved for new functionality in the

future.

(4) --- ABITstream ---

ABITstream can be set only if CALC2:FEED "XTIM:CDP:BSTReam" is selected

(in the lower bit stream window). This command returns the bit streams of

all 15 slots one after the other. The number of Bits of an 16QAM

modulated channel is twice of that of an QPSK modulated channel.

Notes:

(1) This function is available for option FS-K72, FS-K73 or FS-K74 only.

(2) The associated measurement window is selected with the function Set

Active Window (rsfsu_confSetActiveWindow).

Remote-control command(s):

TRACe[:DATA]? PWCD | CTABle | CWCDp | ABITstream

FORMat REAL,32

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

sourceTrace

Variable Type       ViInt32

This parameter indicates which trace will be read out.

Valid Range:

0 - PWCDP (Default Value)

1 - CTABLE

2 - CWCDP

3 - ABITstream

no_ofValues

Variable Type       ViInt32 (passed by reference)

Returns the number of trace data values.

traceData

Variable Type       ViReal64[]

Returns an array of trace data. Measured values are transferred

depending on the display mode (see function description).

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Source Trace) out of range.

BFFC0003  Parameter 3 (No of Values) is null pointer.

BFFC0004  Parameter 4 (Trace Data) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataSEMeasurementResults

ViStatus rsfsu_dataSEMeasurementResults (ViSession instrumentHandle,

ViInt32 noOfValues,

ViReal64 _VI_FAR frequency[],

ViReal64 _VI_FAR level[],

ViReal64 _VI_FAR deltaLimit[],

ViPInt32 returnedValues);

Purpose

This function returns reads the peak list in the spurious measurement. As

results a list of frequency, level and delta to limit line values is

returned. A delta limit of +200dB indicates no limit check is active.

Remote-control command(s):

FORMat REAL,32

TRACe<1|2>[:DATA]? SPURIous

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

noOfValues

Variable Type       ViInt32

This control sets the number of values returned in arrays.

The array size has to be at least equal to this number.

Valid Values: > 0

Default Value: 10

frequency

Variable Type       ViReal64[]

Returns an array of frequency list in Hz.

level

Variable Type       ViReal64[]

Returns an array of level list in dBm.

deltaLimit

Variable Type       ViReal64[]

Returns an array of delta to limit line values in dB. A delta limit

of +200dB indicates no limit check is active.

returnedValues

Variable Type       ViInt32 (passed by reference)

Returns number of value sets (frequency, level and delta to limit

line) read (available to return) from the instrument. Number of

values returned in arrays is limited with the parameter No of Values.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (No Of Values) out of range.

BFFC0003  Parameter 3 (Frequency) is null pointer.

BFFC0004  Parameter 4 (Level) is null pointer.

BFFC0005  Parameter 5 (Delta Limit) is null pointer.

BFFC0006  Parameter 6 (Returned Values) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dataWriteTrace

ViStatus rsfsu_dataWriteTrace (ViSession instrumentHandle,

ViInt32 destinationTrace,

ViInt32 no_ofPoints,

ViReal64 _VI_FAR traceData[]);

Purpose

This function writes trace data to the instrument.

Note:

The associated measurement window is selected with the function Set

Active Window (rsfsu_confSetActiveWindow).

Remote-control command(s):

TRACe

FORMat ASCII

FORMat REAL,32

SYSTem:COMMunicate:GPIB:SELF:RTERminator EOI

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

destinationTrace

Variable Type       ViInt32

This control sets the destination trace data to be writen to the

instrument.

Valid Range:

0 - Trace 1 (Default Value)

1 - Trace 2

2 - Trace 3

no_ofPoints

Variable Type       ViInt32

Sets the number of trace points to be writen to the instrument.

Valid Range: 0 to 250000

Default Value: none

traceData

Variable Type       ViReal64[]

Trace data array.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Destination Trace) out of range.

BFFC0003  Parameter 3 (No. of Points) out of range.

BFFC0004  Parameter 4 (Trace Data) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_dispUpdate

ViStatus rsfsu_dispUpdate (ViSession instrumentHandle,

ViBoolean displayInRemote);

Purpose

This function switches on or off the update of all display elements

during remote control.

Note:

The best performance is obtained when the display output is switched off

during remote control.

Remote-control command(s):

SYSTtem:DISPlay:UPDate ON | OFF

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

displayInRemote

Variable Type       ViBoolean

This control enables or disables instrument display when in remote

control mode.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Display In Remote) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_error_message

ViStatus rsfsu_error_message (ViSession instrumentHandle,

ViStatus statusCode,

ViChar _VI_FAR message[]);

Purpose

This function takes the Status Code returned by the instrument driver

functions, interprets it and returns it as a user readable string.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

statusCode

Variable Type       ViStatus

This control accepts the Status Code returned from the instrument

driver functions.

Default Value:

0 - VI_SUCCESS

message

Variable Type       ViChar[]

This control returns the interpreted Status Code as a user readable

message string.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Status Code) out of range.

BFFC0003  Parameter 3 (Message) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_error_query

ViStatus rsfsu_error_query (ViSession instrumentHandle,

ViPInt32 errorCode,

ViChar _VI_FAR errorMessage[]);

Purpose

This function reads an error code from the instrument's error queue.

Remote-control command(s):

SYSTem:ERRor?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

errorCode

Variable Type       ViInt32 (passed by reference)

This control returns the error code read from the instrument's error

queue.

errorMessage

Variable Type       ViChar[]

This control returns the error message string read from the

instrument's error message queue.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Error Code) is null pointer.

BFFC0003  Parameter 3 (Error Message) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_errorCheckState

ViStatus rsfsu_errorCheckState (ViSession instrumentHandle,

ViBoolean state);

Purpose

This function switches ON/OFF state checking of the instrument (reading

of the Standard Event Register and checking it for error). All driver

functions are using state checking. Switch this function to ON when debug

your application. For better bus throughput and instrument performance

switch it OFF. As default is state checking switched to ON!

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

state

Variable Type       ViBoolean

This control selects ON or OFF instrument state checking.

Valid Range:

VI_FALSE (0) - Off

VI_TRUE  (1) - On (Default Value)

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_getADemodResultValues

ViStatus rsfsu_getADemodResultValues (ViSession instrumentHandle,

ViInt32 window,

ViInt32 demodulation,

ViInt32 resultType,

ViInt32 arraySize,

ViReal64 _VI_FAR resultValues[],

ViPInt32 returnedValues);

Purpose

This function returns analog demodulation result values.

Notes:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

*CLS

SENSe1:ADEMod:AM:RESult? WRITe | AVERage | MAXHold | MINHold

SENSe1:ADEMod:FM:RESult? WRITe | AVERage | MAXHold | MINHold

SENSe1:ADEMod:SPECtrum:RESult? WRITe | AVERage | MAXHold | MINHold

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

demodulation

Variable Type       ViInt32

This control determines the function.

Valid Range:

0 - FM (Default Value)

1 - AM

2 - RF Spectrum

resultType

Variable Type       ViInt32

Selects the result type 1, to be measured in parallel by selected

demodulation.

Valid Range:

0 - Current Meas. Value

1 - Average

2 - Maximum

3 - Minimum

Note:

(1) Current Meas. Value: The current measured values are recorded.

(2) Average: The measurement results are averaged over the given

number of measurements.

(3) Maximum: The maximum result over the given number of measurements

is obtained.

(4) Minimum: The minimum result over the given number of measurements

is obtained.

(5) View: The measurement result is frozen on the display, i.e. it is

not recalculated in further measurements.

(6) The result type indicated must be one of those configured by

Analog Demodulation Type (confVSAMADemodType). Otherwise a query

error will be generated.

arraySize

Variable Type       ViInt32

This parameter determines the size of array result values. The Result

Values array must contain at least this number of items.

resultValues

Variable Type       ViReal64[]

Returns the result type.

returnedValues

Variable Type       ViInt32 (passed by reference)

Returns the number of returned values.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Demodulation) out of range.

BFFC0004  Parameter 4 (Result Type) out of range.

BFFC0005  Parameter 5 (Array Size) out of range.

BFFC0006  Parameter 6 (Result Values) is null pointer.

BFFC0007  Parameter 7 (Returned Values) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_getPeaksCount

ViStatus rsfsu_getPeaksCount (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViPInt32 count);

Purpose

This function returns the number of peaks found during the search. If no

peak search had been executed before a value of 0 is returned.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:FPEaks:COUNt?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 2

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 4

Default Value:

1

count

Variable Type       ViInt32 (passed by reference)

Returns the number of peaks.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Count) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_getPeaksValues

ViStatus rsfsu_getPeaksValues (ViSession instrumentHandle,

ViInt32 window, ViInt32 markerNumber,

ViInt32 axis, ViReal64 _VI_FAR values[],

ViInt32 numOfValues,

ViPInt32 returned_value);

Purpose

This function queries the list of values of the found peaks. The number

of values available can be queried with Get Peaks Count

(rsfsu_getPeaksCount).

Note:

(1) The values are sorted according sort mode. The mode can be set with

function Search Peaks (rsfsu_searchPeaks).

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:FPEaks:X?

CALCulate<1|2>:MARKer<1..4>:FUNCtion:FPEaks:Y?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 2

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 - 4

Default Value:

1

axis

Variable Type       ViInt32

This control determines the readed values.

Valid Range:

0 - X values (Default Value)

1 - Y values

values

Variable Type       ViReal64[]

Returns the array of peak values.

numOfValues

Variable Type       ViInt32

This comand sets the size of array of values.

Valid Range:

0 to free memory

returned_value

Variable Type       ViInt32 (passed by reference)

Returns the number of actualy returned values.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Axis) out of range.

BFFC0005  Parameter 5 (Values) is null pointer.

BFFC0006  Parameter 6 (Num. Of Values) out of range.

BFFC0007  Parameter 7 (Returned value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_getPowerOfSignalPulses

ViStatus rsfsu_getPowerOfSignalPulses (ViSession instrumentHandle,

ViInt32 window,

ViInt32 markerNumber,

ViReal64 timeOffset,

ViReal64 measurementTime,

ViReal64 period,

ViInt32 numOf_pulses,

ViReal64 _VI_FAR values[]);

Purpose

This function measures the power of a sequence of signal pulses with

identical distance in timing. This sort of signal is eg typical for the

slots of a GSM signal. The number of pulses to measure can be selected as

well as measurement time and period of the pulses. In order to define the

position of the first pulse within the measured curve a corresponding

offset can be selected.

The evaluation is always performed on the measurement data of a

previously recorded trace. The data measured during the selected

measurement time will be summarized to one result value per pulse with

respect to the selected detector. The sequence of result values will be

returned as a comma separated list of values.

Remote-control command(s):

CALCulate<1|2>:MARKer1:FUNCtion:MSUMmary?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 2

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 to 1

Default Value:

1

Note:

(1) This control has no effect.

timeOffset

Variable Type       ViReal64

Sets the time offset of first pulse.

Default Value:

0 s

Note:

(1) This value is not range checked

measurementTime

Variable Type       ViReal64

Sets the time of measurement.

Default Value:

50.0E-6 s

Note:

(1) This value is not range checked

period

Variable Type       ViReal64

Sets the period.

Valid Range:

s

Default Value:

500.0E-6s

numOf_pulses

Variable Type       ViInt32

Sets the number of pulses, that should be measured.

Valid Range:

Default Value:

6

Note:

(1) This value should be small or equal to size of array of values.

values

Variable Type       ViReal64[]

This control returns the array of result values.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Time Offset) out of range.

BFFC0005  Parameter 5 (Measurement Time) out of range.

BFFC0006  Parameter 6 (Period) out of range.

BFFC0007  Parameter 7 (Num Of pulses) out of range.

BFFC0008  Parameter 8 (Values) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_getStatusRegister

ViStatus rsfsu_getStatusRegister (ViSession instrumentHandle,

ViInt32 statusRegistersQuery,

ViPInt32 registerValue);

Purpose

This function contains the commands for the querying status reporting

system.

Note:

For detailed description of Status Reporting System see also operating

manual.

Remote-control command(s):

STATus:PRESetSTATus:OPERation:EVENt?

STATus:OPERation:CONDition?

STATus:QUEStionable:EVENt?

STATus:QUEStionable:CONDition?

STATus:QUEStionable:POWer:EVENt?

STATus:QUEStionable:POWer:CONDition?

STATus:QUEStionable:LIMit<1|2>:EVENt?

STATus:QUEStionable:LIMit<1|2>:CONDition?

STATus:QUEStionable:SYNC:EVENt?

STATus:QUEStionable:SYNC:CONDition?

STATus:QUEStionable:ACPLimit:EVENt?

STATus:QUEStionable:ACPLimit:CONDition?

STATus:QUEStionable:FREQuency:EVENt?

STATus:QUEStionable:FREQuency:CONDition?

STATus:QUEStionable:TRANsducer:EVENt?

STATus:QUEStionable:TRANsducer:CONDition?

STATus:QUEStionable?

STATus:QUEStionable:LMARgin<1|2>:EVENt?

STATus:QUEStionable:LMARgin<1|2>:CONDition?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

statusRegistersQuery

Variable Type       ViInt32

This control selects the registers to be queried.

Valid Range:

0 - Operation Event (Default Value)

1 - Operation Condition

2 - Main Event

3 - Main Condition

4 - Power Event

5 - Power Condition

6 - Limit Event Screen A

7 - Limit Condition Screen A

8 - LMargin Event Screen A

9 - LMargin Condition Screen A

10 - Sync Event

11 - Sync Condition

12 - ACP Limit Event

13 - ACP Limit Condition

14 - Frequency Event

15 - Frequency Condition

16 - RESERVED

17 - RESERVED

18 - Queue Next

19 - Limit Event Screen B

20 - Limit Condition Screen B

21 - LMargin Event Screen B

22 - LMargin Condition Screen B

Notes:

(1) Operation Event:

This command queries the contents of the EVENt section of the

STATus:OPERation register.

(2) Operation Condition:

This command queries the CONDition section of the STATus:OPERation

register.

(3) Main Event:

This command queries the contents of the EVENt section of the

STATus:QUEStionable register.

(4) Main Condition:

This command queries the CONDition section of the STATus:QUEStionable

register.

(5) Power Event:

This command queries the contents of the EVENt section of the

STATus:QUEStionable:POWer register.

(6) Power Condition:

This command queries the content of the CONDition section of the

STATus:QUEStionable:POWer register.

(7) Limit Event Screen A(B):

This command queries the contents of the EVENt section of the

STATus:QUEStionable:LIMit register.

(8) Limit Condition Screen A(B):

This command queries the contents of the CONDition section of the

STATus:QUEStionable:LIMit register.

(9) LMargin Event Screen A(B):

This command queries the contents of the EVENt section of the

STATus:QUEStionable:LMARgin register.

(10) LMargin Condition Screen A(B):

This command queries the contents of the CONDition section of the

STATus:QUEStionable:LMARgin register.

(11) ACP Limit Event:

This command queries the contents of the EVENt section of the

STATus:QUEStionable:ACPLimit register.

(12) ACP Limit Condition:

This command queries the contents of the CONDition section of the

STATus:QUEStionable: ACPLimit register.

(13) Frequency Event:

This command queries the contents of the EVENt section of the

STATus:QUEStionableFREQuency register.

(14) Frequency Condition:

This command queries the contents of the CONDition section of the

STATus:QUEStionable:FREQuency register.

(15) Queue Next:

This command queries the earliest entry to the error queue, thus

deleting it.

(16) Sync Event:

This command queries the contents of the EVENt section of the

STATus:QUEStionable:SYNC register.

(17) Sync Condition:

This command queries the contents of the CONDition section of the

STATus:QUEStionable:SYNC register.

registerValue

Variable Type       ViInt32 (passed by reference)

Returns content of selected register.

Note:

For detailed description of Status Reporting System see also

operating manual.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Status Registers Query) out of range.

BFFC0003  Parameter 3 (Register Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_getSweepCount

ViStatus rsfsu_getSweepCount (ViSession instrumentHandle,

ViInt32 window,

ViPInt32 currentSweepCount);

Purpose

This function returns the current number of started sweeps. A sweep count

value should be set and the device should be in single sweep mode.

Remote-control command(s):

[SENSe<1|2>:]SWEep:COUNt:CURRent?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

currentSweepCount

Variable Type       ViInt32 (passed by reference)

This control queries the number of started sweeps.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Current Sweep Count) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_getVSAMParam

ViStatus rsfsu_getVSAMParam (ViSession instrumentHandle, ViInt32 window,

ViInt32 parameterToRead, ViPReal64 value);

Purpose

This function returns actual setting of analog demodulation parameter.

Notes:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

SENSe1:ADEMod:RLENgth?

SENSe1:ADEMod:SRATe?

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window to set up FSK ref deviation.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) Only Window A can be used now.

parameterToRead

Variable Type       ViInt32

This control determines the parameter.

Valid Range:

0 - Record Length (Default Value)

1 - Sampling Rate

value

Variable Type       ViReal64 (passed by reference)

Returns parameter value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Parameter To Read) out of range.

BFFC0004  Parameter 4 (Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_getVSFMOffset

ViStatus rsfsu_getVSFMOffset (ViSession instrumentHandle,

ViInt32 window, ViInt32 resultType,

ViPReal64 value);

Purpose

This function calculates the FM offset of the current measured data set.

Notes:

(1) This function is available only in conjunction with option FS-K7.

Remote-control command(s):

SENSe1:ADEMod:FM:OFFSet? IMMediate

SENSe1:ADEMod:FM:OFFSet? AVERage

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

resultType

Variable Type       ViInt32

This control determines the result type.

Valid Range:

0 - Immediate (Default Value)

1 - Average

Note:

(1) Immediate: The current measurement results are used for

calculating the FM offset.

(2) Average: The measurement results averaged over the given number

of measurements are used for calculation of the FM offset.

(3) If no average measurement has been active during the last

measurement sequence, only the type Immediate will return a valid

result (data to calculate the FM offset are taken from the last

measured data set). Type Average will cause a query error in this

case.

value

Variable Type       ViReal64 (passed by reference)

Returns measured value.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Result Type) out of range.

BFFC0004  Parameter 4 (Value) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_init

ViStatus rsfsu_init (ViRsrc resourceName, ViBoolean IDQuery,

ViBoolean resetDevice,

ViPSession instrumentHandle);

Purpose

This function performs the following initialization actions:

- Opens a session to the Default Resource Manager resource and a session

to the specified device using the interface and address specified in the

Resource_Name control.

- Performs an identification query on the Instrument.

- Resets the instrument to a known state.

- Sends initialization commands to the instrument that set any necessary

programmatic variables such as Headers Off, Short Command form, and Data

Transfer Binary to the state necessary for the operation of the

instrument driver.

- Returns an Instrument Handle which is used to differentiate between

different sessions of this instrument driver.

- Each time this function is invoked a Unique Session is opened.  It is

possible to have more than one session open for the same resource.

Parameter List

resourceName

Variable Type       ViRsrc

This control specifies the interface and address of the device that

is to be initialized (Instrument Descriptor). The exact grammar to be

used in this control is shown in the note below.

Default Value:  "GPIB::20"

Notes:

(1) Based on the Instrument Descriptor, this operation establishes a

communication session with a device.  The grammar for the Instrument

Descriptor is shown below.  Optional parameters are shown in square

brackets ([]).

Interface   Grammar

------------------------------------------------------

Serial      ASRL<port>[::INSTR]

GPIB        GPIB[board]::primary address[::secondary address]

[::INSTR]

RSIB        RSIB::remote_host::INSTR

The ASRL keyword is used for serial interface.

The GPIB keyword is used for GPIB interface.

The RSIB keyword is used for RSIB interface.

Examples:

(1) Serial - "ASRL1"

(2) GPIB   - "GPIB::20"

(3) RSIB   - "RSIB::192.168.1.33::INSTR"

The default value for optional parameters are shown below.

Optional Parameter          Default Value

-----------------------------------------

board                       0

secondary address           none - 31

IDQuery

Variable Type       ViBoolean

This control specifies if an ID Query is sent to the instrument

during the initialization procedure.

Valid Range:

VI_FALSE (0) - Skip Query

VI_TRUE  (1) - Do Query (Default Value)

Notes:

(1) Under normal circumstances the ID Query ensures that the

instrument initialized is the type supported by this driver. However

circumstances may arise where it is undesirable to send an ID Query

to the instrument.  In those cases; set this control to "Skip Query"

and this function will initialize the selected interface, without

doing an ID Query.

resetDevice

Variable Type       ViBoolean

This control specifies if the instrument is to be reset to its

power-on settings during the initialization procedure.

Valid Range:

VI_FALSE (0) - Don't Reset

VI_TRUE  (1) - Reset Device (Default Value)

Notes:

(1) If you do not want the instrument reset. Set this control to

"Don't Reset" while initializing the instrument.

instrumentHandle

Variable Type       ViSession (passed by reference)

This control returns an Instrument Handle that is used in all

subsequent function calls to differentiate between different sessions

of this instrument driver.

Notes:

(1) Each time this function is invoked a Unique Session is opened.

It is possible to have more than one session open for the same

resource.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0001  Parameter 1 (Resource Name) out of range.

BFFC0002  Parameter 2 (ID Query) out of range.

BFFC0003  Parameter 3 (Reset Device) out of range.

BFFC0004  Parameter 4 (Instrument Handle) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_OPTList

ViStatus rsfsu_OPTList (ViSession instrumentHandle,

ViChar _VI_FAR optionsList[],

ViChar _VI_FAR instrumentVersion[]);

Purpose

This function returns the instrument version number and the list of

options which the instrument is equipped with.

Remote-control command(s):

*OPT?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

optionsList

Variable Type       ViChar[]

Returns the list of options present in the instrument. For more

detail see the User's manual.

instrumentVersion

Variable Type       ViChar[]

Returns instrument version.

FSU 3  - 0.0 Hz to 3.6 GHz

FSU 8  - 0.0 Hz to 8.0 GHz

FSU 26 - 0.0 Hz to 26.5 GHz

FSU 46 - 0.0 Hz to 46.0 GHz

FSU 50 - 0.0 Hz to 50.0 GHz

UNKNOWN - Unknown instrument model

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Options List) is null pointer.

BFFC0003  Parameter 3 (Instrument Version) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_readInstrData

ViStatus rsfsu_readInstrData (ViSession instrumentHandle,

ViInt32 numberBytesToRead,

ViChar _VI_FAR readBuffer[],

ViPInt32 numBytesRead);

Purpose

This function reads data from the instrument's output buffer and returns

it to the specified variable in memory.

Notes:

(1) Because the instrument may return both numeric and text data in

response to queries this function returns the data in string format.

(2) If valid data is not available at the instrument's output buffer when

this function is called the instrument will hang up and the function will

not return until it times out.  If the time-out is disabled this function

will hang indefinitely and it may be necessary to reboot the computer to

break out.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

numberBytesToRead

Variable Type       ViInt32

The number of bytes specified by this control should be greater than

or equal to the number of bytes which are to be read from the

instrument. If the actual number of bytes to be read is greater than

the number this control specifies then multiple reads will be

required to empty the instrument's output buffer.

If the instrument's output buffer is not emptied the instrument may

return invalid data when future reads are performed.

Default Value:  50 (See NOTE 2)

Notes:

(1) If NO DATA is available at the instrument's output buffer when

this function is called the instrument will hang up and the function

will not return until it times out.  If the time-out is disabled,

this function will hang indefinitely and it may be necessary to

reboot the computer.

(2) If the number of bytes expected is greater than 50 the value of

this control may be increased. If the actual number of bytes read is

larger than the declared size of the read buffer a run-time error

will be generated.

readBuffer

Variable Type       ViChar[]

The incoming data from the instrument is placed into this variable.

Notes:

(1) Because the instrument may return both numeric and text data in

response to queries this function returns the data in string format.

(2) This function does not overwrite any old data left in the string

variable from the last time the function was called. String data in

LabWindows/CVI is terminated with an ASCII null(0x0) and string

manipulation functions will only recognize data before the ASCII

null.

(3) The declared size of the string variable must be greater than the

actual number of bytes read from the instrument if it is not a

run-time error will be generated.

numBytesRead

Variable Type       ViInt32 (passed by reference)

This variable contains the actual number of bytes read from the

instrument. This is the value which is returned by the read function.

Notes:

(1) If the actual number of bytes read is less than the number of

bytes specified in the Number Bytes To Read control then the output

buffer has probably been emptied. If the read function fails and the

number of bytes read is 0, the most probable cause for the failure is

there was no data available at the instrument's output buffer.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Number Bytes To Read) out of range.

BFFC0003  Parameter 3 (Read Buffer) is null pointer.

BFFC0004  Parameter 4 (NumBytes Read) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_readMultiFrameData

ViStatus rsfsu_readMultiFrameData (ViSession instrumentHandle,

ViInt32 firstFrame,

ViInt32 lastFrame,

ViInt32 no_ofSearchData,

ViString searchPattern,

ViPInt32 searchPosition,

ViReal64 _VI_FAR dataCompositeEVM[],

ViPInt32 no_ofDataCompositeEVM,

ViReal64 _VI_FAR dataPCDE[],

ViPInt32 no_ofDataPCDE,

ViReal64 _VI_FAR dataPowerVsSlot[],

ViPInt32 no_ofDataPowerVsSlot,

ViReal64 _VI_FAR dataPowerVsSymbol[],

ViPInt32 no_ofDataPowerVsSymbol,

ViReal64 _VI_FAR dataSymbolEVM[],

ViPInt32 no_ofDataSymbolEVM,

ViChar _VI_FAR dataBitStream[],

ViPInt32 no_ofDataBitStream,

ViReal64 _VI_FAR dataCarrierFrequencyError[],

ViPInt32 no_ofDataCarrierFreqError,

ViReal64 _VI_FAR dataModulationType[],

ViPInt32 no_ofModulationType,

ViChar _VI_FAR dataABitStream[],

ViPInt32 no_ofDataABitStream,

ViReal64 _VI_FAR reserved2[],

ViPInt32 reserved3,

ViReal64 _VI_FAR reserved4[],

ViPInt32 reserved5,

ViReal64 _VI_FAR reserved6[],

ViPInt32 reserved7,

ViReal64 _VI_FAR reserved8[],

ViPInt32 reserved9);

Purpose

Purpose of this function is to reads various data such as

Composite EVM

Peak Code Domain Error (PCDE)

Power versus slot

Power versus symbol

Symbol EVM

Bit Stream

Carrier Frequency Error

Modulation Type

of a single channel of a 3GPP-FDD signal. All data which are assigned to

a channel and / or slot is read out for all slots for all frames within a

range of frames to be defined by the user.

Optionally, the occurrence of a certain pattern within the bit stream can

be checked. The position of the first occurrence of the search pattern is

reported to the calling function.

The read out for all data can be suppressed by passing a NULL pointer as

data array and is such faster, if not all information is needed.

All data are returned in the format set elsewhere in the program, e.g.

power is measured relatively or absolutely, depending on instrument

setting.

Note:

(1) This function is available only in conjunction with option FS-K72,

FS-K73 or FS-K74.

(2) This function requires FW 2.3x / 3.3x or above.

(3) Collecting of all data might take longer time.

Remote-control command(s):

CALCulate:MARKer:FUNCtion:WCDPower:RESult MTYPe | FERRor

CALCulate2:FEED 'XTIM:CDP:ERR:SUMMary'

CALCulate1:FEED 'XTIMe:CDPower:MACCuracy'

CALCulate1:FEED 'XTIMe:CDPower:ERRor:PCDomain'

CALCulate1:FEED 'XTIMe:CDPower:PVSLot'

CALCulate1:FEED 'XTIMe:CDPower:PVSYmbol'

CALCulate1:FEED 'XTIMe:CDPower:SYMBol:EVM'

CALCulate1:FEED 'XTIMe:CDPower:BSTReam'

CALCulate2:FEED 'XTIM:CDP:ERR:SUMMary'

SENSe1:CDPower:FRAMe:VALue

SENSe1:CDPower:SLOT

TRACe<1|2>? TRACE1 | TRACE2 | TRACE3

TRACe[:DATA]? ABITstream

*WAI

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

firstFrame

Variable Type       ViInt32

The first frame to read the data from.

Valid Range: 0 to <set capture length>

Default Value: 0

Note:

First Frame must be equal or less the number of captured frames (see

command set capture length :CDPower:IQLength (new for FW 2.3x/3.3x)).

Maximum number of captured frames is:

--- FSP ---

Downlink (K72/K74):

EXT TRIGGER - none

FREE RUN - none

Uplink (K73):

EXT TRIGGER - 1 slot

FREE RUN - 1 slot

--- FSP (with B70) ---

Downlink (K72/K74):

EXT TRIGGER - 3 frames

FREE RUN - 2 frames

Uplink (K73):

EXT TRIGGER - 3 frames

FREE RUN - 2 frames

--- FSU ---

Downlink (K72/K74):

EXT TRIGGER - 3 frames

FREE RUN - 2 frames

Uplink (K73):

EXT TRIGGER - 3 frames

FREE RUN - 2 frames

--- FSQ ---

Downlink (K72/K74):

EXT TRIGGER - 100 frames

FREE RUN - 100 frames

Uplink (K73):

EXT TRIGGER - 100 frames

FREE RUN - 100 frames

lastFrame

Variable Type       ViInt32

The last frame to read the data from.

Valid Range: <First Frame> to 100 (or <set capture length>)

Default Value: 1

Note:

Last Frame must be equal or less the number of captured frames (see

command set capture length :CDPower:IQLength (new for FW 2.3x/3.3x)).

Maximum number of captured frames is:

--- FSP ---

Downlink (K72/K74):

EXT TRIGGER - none

FREE RUN - none

Uplink (K73):

EXT TRIGGER - 1 slot

FREE RUN - 1 slot

--- FSP (with B70) ---

Downlink (K72/K74):

EXT TRIGGER - 3 frames

FREE RUN - 2 frames

Uplink (K73):

EXT TRIGGER - 3 frames

FREE RUN - 2 frames

--- FSU ---

Downlink (K72/K74):

EXT TRIGGER - 3 frames

FREE RUN - 2 frames

Uplink (K73):

EXT TRIGGER - 3 frames

FREE RUN - 2 frames

--- FSQ ---

Downlink (K72/K74):

EXT TRIGGER - 100 frames

FREE RUN - 100 frames

Uplink (K73):

EXT TRIGGER - 100 frames

FREE RUN - 100 frames

no_ofSearchData

Variable Type       ViInt32

Size of the Search Pattern.

Valid Range: 0 to maxint

Default Value: 0

Notes:

(1) If < 1, no search is performed.

(2) Data Bit Stream and No of Data Bit Stream parameters must be

configured to measure stream data.

searchPattern

Variable Type       ViString

The position of the first occurrence of the defined search pattern in

the bit stream data will be then searched.

Valid Range: none

Default Value: VI_NULL

Example for a bit stream trace: "01090100101..."

Example for search pattern: "901"

Notes:

(1) If VI_NULL is passed, no search is performed.

(2) Data Bit Stream and No of Data Bit Stream parameters must be

configured to measure stream data.

searchPosition

Variable Type       ViInt32 (passed by reference)

Returns the position of the first occurrence of the search pattern in

the bit stream data. The A-Bit stream is used, if both bit streams

are read in.

If VI_NULL is passed, no search will be performed.

If Search Position = -1, the Search Pattern is not contained in the

data stream.

Data Bit Stream and No of Data Bit Stream parameters must be

configured to measure stream data.

Examples:

Search pattern: "001"

Data Stream: "001..."

Search Position = 0

Search pattern: "001"

Data Stream: "00001"

Search Position = 4

Search pattern: "111"

Data Stream: "0000...0"

Search Position = -1

Value 9 stands for "invalid data":

Search pattern: "1991"

Data Stream: "01991..."

Search Position = 1

dataCompositeEVM

Variable Type       ViReal64[]

Array for Composite EVM data. 15 values per frame are read in.

If VI_NULL is passed, no Composite EVM data are read in.

The size of array must be at least 15 * <number of frames>

no_ofDataCompositeEVM

Variable Type       ViInt32 (passed by reference)

Returns the number of Compsite EVM data read in.

If VI_NULL is passed or parameter value is -1, no data are read in.

dataPCDE

Variable Type       ViReal64[]

Array for Peak Code Domain Error (PCDE) data. 15 values per frame are

read in.

If VI_NULL is passed, no PCDE data are read in.

The size of array must be at least 15 * <number of frames>

no_ofDataPCDE

Variable Type       ViInt32 (passed by reference)

Returns the number of Peak Code Domain Error (PCDE) data read in.

If VI_NULL is passed or parameter value is -1, no data are read in.

dataPowerVsSlot

Variable Type       ViReal64[]

Array for Power Vs Slot data. 15 values per frame are read in.

If VI_NULL is passed, no Power Vs Slot data are read in.

The size of array must be at least 15 * <number of frames>

Note:

For options FS-K72/K74 instrument returns 16 values per frame; the

last value of each frame is equal to the first value of the next

frame. This function returns 15 values per frame; the first value of

each frame is suppressed.

no_ofDataPowerVsSlot

Variable Type       ViInt32 (passed by reference)

Returns the number of power versus symbol data read in.

If VI_NULL is passed or parameter value is -1, no data are read in.

dataPowerVsSymbol

Variable Type       ViReal64[]

Array for power versus symbol data. Up to 640 * 15 values per frame

are read in, depending on the type of channel.

If VI_NULL is passed, no power versus symbol data are read in.

The size of array must be at least up to 640 * 15 * <number of

frames>.

no_ofDataPowerVsSymbol

Variable Type       ViInt32 (passed by reference)

Returns the number of power versus symbol data read in.

If VI_NULL is passed or parameter value is -1, no data are read in.

dataSymbolEVM

Variable Type       ViReal64[]

Array for symbol EVM data. Up to 640 * 15 values per frame are read

in, depending on the type of channel.

If VI_NULL is passed, no PCDE data are read in.

The size of array must be at least up to 640 * 15 * <number of

frames>.

no_ofDataSymbolEVM

Variable Type       ViInt32 (passed by reference)

Returns the number of power versus symbol data read in.

If VI_NULL is passed or parameter value is -1, no data are read in.

dataBitStream

Variable Type       ViChar[]

Array for bit stream data. Up to 640 * 15 * 4 values per frame are

read in, depending on the type of channel.

If VI_NULL is passed, no bit stream data are read in.

The size of array must be at least up to 640 * 15 * 4 * <number of

frames>. Each bit is stored in one array element.

One value is transferred per bit (range 0,1,9). Each symbol contains

of two consecutive bits. The number of symbols is not constant and

may vary for each sweep. Specific symbols in the bit stream may be

invalid depending on the channel type and the bit rate (symbols

without power). The assigned invalid bits are marked by "9".

Note:

This parameter should only be used up to FW-Version 2.3.x (NT) /

3.3.x (XP). Use dataABitStream instead.

no_ofDataBitStream

Variable Type       ViInt32 (passed by reference)

Returns the number of bit stream data.

If VI_NULL is passed or parameter value is -1, no data are read in.

Note:

This parameter should only be used up to FW-Version 2.3.x (NT) /

3.3.x (XP). Use dataABitStream instead.

dataCarrierFrequencyError

Variable Type       ViReal64[]

Array for carrier frequency error data. 15 values per frame are read

in.

If VI_NULL is passed, no bit carrier frequency error data are read

in.

The size of array must be at least 15 * <number of frames>.

no_ofDataCarrierFreqError

Variable Type       ViInt32 (passed by reference)

Returns the number of carrier frequency data.

If VI_NULL is passed or parameter value is -1, no data are read in.

dataModulationType

Variable Type       ViReal64[]

Array for modulation type data. 15 values per frame are read in.

If VI_NULL is passed, no modulation type data are read in.

The size of array must be at least 15 * <number of frames>.

The output data have the following meaning:

2  - QPSK modulation type QPSK

4  - 16QAM modulation type 16QAM

15 - NONE there is no power in the selected channel slot.

[slot is switched OFF.] (According to 3GPP the power of a

HSDPA channel can be switched every 2ms, which are 3

slots.)

no_ofModulationType

Variable Type       ViInt32 (passed by reference)

Returns the number of modulation type data.

If VI_NULL is passed or parameter value is -1, no data are read in.

dataABitStream

Variable Type       ViChar[]

Array for bit stream data. Up to 640 * 15 * 4 values per frame are

read in, depending on the type of channel.

If VI_NULL is passed, no bit stream data are read in.

The size of array must be at least up to 640 * 15 * 4 * <number of

frames>. Each bit is stored in one array element.

Note:

This parameter is supported only beginning with FW-Version 2.4.x (NT)

/ 3.4.x (XP) reserved, up to FW 2 / 3.3.x.

no_ofDataABitStream

Variable Type       ViInt32 (passed by reference)

Returns the number of bit stream data.

If VI_NULL is passed, or if it points to -1, no data are read in.

Note:

This parameter is supported only beginning with FW-Version 2.4.x (NT)

/ 3.4.x (XP) reserved, up to FW 2 / 3.3.x.

reserved2

Variable Type       ViReal64[]

Reserved for future use.

reserved3

Variable Type       ViInt32 (passed by reference)

Reserved for future use.

reserved4

Variable Type       ViReal64[]

Reserved for future use.

reserved5

Variable Type       ViInt32 (passed by reference)

Reserved for future use.

reserved6

Variable Type       ViReal64[]

Reserved for future use.

reserved7

Variable Type       ViInt32 (passed by reference)

Reserved for future use.

reserved8

Variable Type       ViReal64[]

Reserved for future use.

reserved9

Variable Type       ViInt32 (passed by reference)

Reserved for future use.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (First Frame) out of range.

BFFC0003  Parameter 3 (Last Frame) out of range.

BFFC0004  Parameter 4 (No of Search Data) out of range.

BFFC0005  Parameter 5 (Search Pattern) out of range.

BFFC0006  Parameter 6 (Search Position) is null pointer.

BFFC0007  Parameter 7 (Data Composite EVM) is null pointer.

BFFC0008  Parameter 8 (No of Data Composite EVM) is null pointer.

BFFC0009  Parameter 9 (Data PCDE) is null pointer.

BFFC000A  Parameter 10 (No of Data PCDE) is null pointer.

BFFC000B  Parameter 11 (Data Power Vs Slot) is null pointer.

BFFC000C  Parameter 12 (No of Data Power Vs Slot) is null pointer.

BFFC000D  Parameter 13 (Data Power Vs Symbol) is null pointer.

BFFC000E  Parameter 14 (No of Data Power Vs Symbo) is null pointer.

BFFC000F  Parameter 15 (Data Symbol EVM) is null pointer.

BFFC0010  Parameter 16 (No of Data Symbol EVM) is null pointer.

BFFC0011  Parameter 17 (Data Bit Stream) is null pointer.

BFFC0012  Parameter 18 (No of Data Bit Stream) is null pointer.

BFFC0013  Parameter 19 (Data Carrier Frequency Er) is null pointer.

BFFC0014  Parameter 20 (No of Data Carrier Freq E) is null pointer.

BFFC0015  Parameter 21 (Data Modulation Type) is null pointer.

BFFC0016  Parameter 22 (No of Modulation Type) is null pointer.

BFFC0017  Parameter 23 (Data ABit Stream) is null pointer.

BFFC0018  Parameter 24 (No of Data ABit Stream) is null pointer.

BFFC0019  Parameter 25 (Reserved 2) is null pointer.

BFFC001A  Parameter 26 (Reserved 3) is null pointer.

BFFC001B  Parameter 27 (Reserved 4) is null pointer.

BFFC001C  Parameter 28 (Reserved 5) is null pointer.

BFFC001D  Parameter 29 (Reserved 6) is null pointer.

BFFC001E  Parameter 30 (Reserved 7) is null pointer.

BFFC001F  Parameter 31 (Reserved 8) is null pointer.

BFFC0020  Parameter 32 (Reserved 9) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_readToFile

ViStatus rsfsu_readToFile (ViSession instrumentHandle, ViString source,

ViString destination);

Purpose

This function is used to read data from the instrument and write it to a

user specified file on the host computer.

Remote-control command(s):

MMEMory:DATA? <file_name>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

source

Variable Type       ViString

This control selects file for the data transfer from instrument to

control computer.

Valid Range: not tested

Default Value: "D:\\user\\data\\test.hcp"

destination

Variable Type       ViString

This control defines destination file to which the data transfer from

instrument to control computer applies.

Valid Range: not tested

Default Value: "\\test.hcp"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Source) out of range.

BFFC0003  Parameter 3 (Destination) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_reset

ViStatus rsfsu_reset (ViSession instrumentHandle);

Purpose

This function resets the instrument to a known state and sends

initialization commands to the instrument that set any necessary

programmatic variables such as Headers Off, Short Command form, and Data

Transfer Binary to the state necessary for the operation of the

instrument driver.

Remote-control command(s):

*RST

*CLS

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_revision_query

ViStatus rsfsu_revision_query (ViSession instrumentHandle,

ViChar _VI_FAR instrumentDriverRevision[],

ViChar _VI_FAR firmwareRevision[]);

Purpose

This function returns the revision numbers of the instrument driver and

instrument firmware, and tells the user with which  instrument firmware

this revision of the driver is compatible.

This instrument driver's Revision Number is "Rev 1.7, 04/2005, CVI 7.1"

and is compatible with Instrument FSU 3 Firmware Revision "3.60".

Remote-control command(s):

*IDN?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

instrumentDriverRevision

Variable Type       ViChar[]

This control returns the Instrument Driver Software Revision.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

firmwareRevision

Variable Type       ViChar[]

This control returns the Instrument Firmware Revision.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Instrument Driver Revisio) is null pointer.

BFFC0003  Parameter 3 (Firmware Revision) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_searchPeaks

ViStatus rsfsu_searchPeaks (ViSession instrumentHandle, ViInt32 window,

ViInt32 markerNumber, ViInt32 numberOfPeaks,

ViInt32 sorting);

Purpose

This function is useful for min/max peak searching. It sets the sort mode

of searched peaks and searches the entered number of peaks.

Remote-control command(s):

CALCulate<1|2>:MARKer<1..4>:FUNCtion:FPEaks:SORT X | Y

CALCulate<1|2>:MARKer<1..4>:FUNCtion:FPEaks

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range: 1 to 2

Default Value: 1

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

markerNumber

Variable Type       ViInt32

Selects the marker.

Valid Range:

1 to 4

Default Value:

1

numberOfPeaks

Variable Type       ViInt32

Sets the number of peaks, should be searched.

Valid Range:

1 to 50

Default Value:

1

sorting

Variable Type       ViInt32

Sets the sort mode of searched peaks.

Valid Range:

0 - Ascending (Default Value)

1 - Descending

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Marker Number) out of range.

BFFC0004  Parameter 4 (Number Of Peaks) out of range.

BFFC0005  Parameter 5 (Sorting) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_self_test

ViStatus rsfsu_self_test (ViSession instrumentHandle,

ViPInt32 selfTestResult,

ViChar _VI_FAR selfTestMessage[]);

Purpose

This function runs the instrument's self test routine and returns the

test result(s).

Remote-control command(s):

*TST?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

selfTestResult

Variable Type       ViInt32 (passed by reference)

This control contains the value returned from the instrument self

test.  Zero means success.  For any other code, see the device's

operator's manual.

selfTestMessage

Variable Type       ViChar[]

This control contains the string returned from the self test. See the

device's operation manual for an explanation of the string's

contents.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Self Test Result) is null pointer.

BFFC0003  Parameter 3 (Self-Test Message) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_self_test_result

ViStatus rsfsu_self_test_result (ViSession instrumentHandle,

ViInt32 bufferSize,

ViChar _VI_FAR resultTable[]);

Purpose

This function reads the results of the selftest out of the instrument.

The lines of the result table are output as string data separated by

commas.

Remote-control command(s):

*CLS

DIAGnostic:SERVice:STESt:RESult?

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

bufferSize

Variable Type       ViInt32

This control specifies the size of the buffer.

Note:

This value is not range checked.

resultTable

Variable Type       ViChar[]

This control contains the results of the selftest out of the

instrument. The lines of the restlt table are output as string data

separated by commas.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Buffer Size) out of range.

BFFC0003  Parameter 3 (Result Table) is null pointer.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_SEStopMeasurement

ViStatus rsfsu_SEStopMeasurement (ViSession instrumentHandle);

Purpose

This function stops the measurement. The measurement data can be analyzed

then.

Remote-control command(s):

ABORt

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_setADEMBandwidth

ViStatus rsfsu_setADEMBandwidth (ViSession instrumentHandle,

ViInt32 window, ViInt32 bandwidth);

Purpose

This function sets the bandwidth for FM demodulation. Depending on the

selected demodulation bandwidth, the instrument selects the required

sampling rate.

Note:

(1) This VI is available only with option FS-K7 (FM Demodulator).

Remote-control command(s):

SENSe<1|2>:BWIDth:DEModulation

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

bandwidth

Variable Type       ViInt32

This control sets the bandwidth for analog demodulation. Depending on

the selected demodulation bandwidth, the instrument selects the

required sampling rate. The available values of the demodulation

bandwidths are determined by the sampling rates.

Valid Range:

0 - 10 MHz

1 - 8 MHz

2 - 5 MHz   (Default Value)

3 - 3 MHz

4 - 1.6 MHz

5 - 800 kHz

6 - 400 kHz

7 - 200 kHz

8 - 100 kHz

9 - 50 kHz

10 - 25 kHz

11 - 12.5 kHz

12 - 6.4 kHz

13 - 3.2 kHz

14 - 1.6 kHz

15 - 800 Hz

16 - 400 Hz

17 - 200 Hz

18 - 100 Hz

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Bandwidth) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_setRFInYIGFilterState

ViStatus rsfsu_setRFInYIGFilterState (ViSession instrumentHandle,

ViInt32 window, ViInt32 state);

Purpose

This function sets the state of the input YIG filter.

Remote-control command(s):

INPut<1|2>:FILTer:YIG OFF | ON

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

state

Variable Type       ViInt32

Sets the state of YIG filter.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_setRFInYIGFilterTempCorr

ViStatus rsfsu_setRFInYIGFilterTempCorr (ViSession instrumentHandle,

ViBoolean state);

Purpose

This function turns on or off the automatic correction of YIG filter's

frequency drift.

Remote-control command(s):

CORRection:YIG:TEMP:AUTO OFF | ON

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

state

Variable Type       ViBoolean

This control turns on or off the automatic correction of YIG filter's

frequency drift.

Valid Range:

VI_FALSE (0) - Off (Default Value)

VI_TRUE  (1) - On

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (State) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_setStatusRegister

ViStatus rsfsu_setStatusRegister (ViSession instrumentHandle,

ViInt32 registerOperation,

ViInt32 questionableRegister,

ViInt32 enable, ViInt32 PTransition,

ViInt32 NTransition);

Purpose

This function sets the Enable, NTransition, and PTransition bits     of

the Main, Power, Limit, LMargin, ACP Limit, Frequency, status

questionable and operating registers.

Note:

(1) For detailed description of Status Reporting System see also

operating manual.

Remote-control command(s):

STATus:PRESet

STATus:QUEStionable:POWer:ENABle

STATus:QUEStionable:LIMit<1|2>:ENABle

STATus:QUEStionable:LMARgin<1|2>:ENABle

STATus:QUEStionable:SYNC:ENABle

STATus:QUEStionable:ACPLimit:ENABle

STATus:QUEStionable:FREQuency:ENABle

STATus:QUEStionable:POWer:PTRansition

STATus:QUEStionable:LIMit<1|2>:PTRansition

STATus:QUEStionable:LMARgin<1|2>:PTRansition

STATus:QUEStionable:SYNC:PTRansition

STATus:QUEStionable:ACPLimit:PTRansition

STATus:QUEStionable:FREQuency:PTRansition

STATus:QUEStionable:POWer:NTRansition

STATus:QUEStionable:LIMit<1|2>:NTRansition

STATus:QUEStionable:LMARgin<1|2>:NTRansition

STATus:QUEStionable:SYNC:NTRansition

STATus:QUEStionable:ACPLimit:NTRansition

STATus:QUEStionable:FREQuency:NTRansition

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

registerOperation

Variable Type       ViInt32

Specifies the operation to perform.

Valid Range:

0 - Preset Registers (Default Value)

1 - Set Questionable Register

2 - Set Operation Register

Notes:

(1) Preset Registers:

This command resets the edge detectors and ENABle parts of all

registers to a defined value. All PTRansition parts are set to FFFFh,

i.e., all transitions from 0 to 1 are detected. All NTRansition parts

are set to 0, i.e., a transition from 1 to 0 in a CONDition bit is

not detected. The ENABle part of the STATus:OPERation and

STATus:QUEStionable registers are set to 0, i.e., all events in these

registers are not passed on.

questionableRegister

Variable Type       ViInt32

This control selects the registers that will be configured.

Valid Range:

0 - Main(Default Value)

1 - Power

2 - Limit Screen A

3 - LMargin Screen A

4 - Sync

5 - ACP Limit

6 - Frequency

7 - Limit Screen B

8 - LMargin Screen B

enable

Variable Type       ViInt32

This control sets the enable bits of the status register.

Valid range: 0 to 65535

Default Value: 0

Note:

(1) The ENABle register selectively enables the individual events of

the associated EVENt section for the sum bit in the status byte.

PTransition

Variable Type       ViInt32

This control sets the PTransition bits (edge detectors) of the status

register.

Valid range: 0 to 65535

Default Value: 0

NTransition

Variable Type       ViInt32

This control sets the NTransition bits (edge detectors) of the status

register.

Valid range: 0 to 65535

Default Value: 0

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Register Operation) out of range.

BFFC0003  Parameter 3 (Questionable Register) out of range.

BFFC0004  Parameter 4 (Enable) out of range.

BFFC0005  Parameter 5 (PTransition) out of range.

BFFC0006  Parameter 6 (NTransition) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_setVideoBWType

ViStatus rsfsu_setVideoBWType (ViSession instrumentHandle,

ViInt32 window, ViInt32 type);

Purpose

This function selects the position of the video filter within the signal

path, as soon as the resolution bandwidth is <= 100 kHz.

Remote-control command(s):

SENSe<1|2>:BANDwidth:VIDeo:TYPE LINear | LOGarithmic

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

window

Variable Type       ViInt32

Selects the display window that will be configured.

Valid Range:

1 - Window A (Default Value)

2 - Window B

Notes:

(1) If the instrument is in full screen mode and the screen selected

is not the active screen then the instrument will generate an error.

type

Variable Type       ViInt32

Sets the position of the video filter within the signal path.

Valid Ranges:

0 - Logarithmic

1 - Linear (Default Value)

Notes:

(1) With selection Linear the video filter is placed in front of the

logarithmic amplifier (default). With selection Logarithmic the video

filter is placed behind the logarithmic amplifier.

The major difference between the two positions affects the settling

characteristics of falling signal slopes: With selection Linear the

falling slope is measured as being less "steep" than with selection

Logarithmic.

This behaviour is a result from the transformation of linear power to

logarithmic power values: reducing a linear power value by 50%

results in a logarithmic reduction of only 3 dB.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Window) out of range.

BFFC0003  Parameter 3 (Type) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_writeFromFile

ViStatus rsfsu_writeFromFile (ViSession instrumentHandle,

ViString source, ViString destination);

Purpose

This function is used to read data from the host computer and write it to

a user specified file in the instrument.

Remote-control command(s):

MMEMory:DATA <file_name>

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

source

Variable Type       ViString

This control selects file for the data transfer from control computer

to instrument.

Valid Range: not tested

Default Value: "\\test.hcp"

destination

Variable Type       ViString

This control defines destination file to which the data transfer from

control computer to instrument applies.

Valid Range: not tested

Default Value: "D:\\user\\data\\test.hcp"

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Source) out of range.

BFFC0003  Parameter 3 (Destination) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors

________________________________________________________________________________

rsfsu_writeInstrData

ViStatus rsfsu_writeInstrData (ViSession instrumentHandle,

ViString writeBuffer);

Purpose

This function writes commands and queries to the instrument to modify

parameters and query device settings.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

writeBuffer

Variable Type       ViString

The user can use this control to send common commands and queries to

the instrument. This control can also be used to write any valid

command to the instrument.

Default Value:  "*RST"

Notes:

(1) The command or query to be sent to the instrument may be a

literal enclosed in double quotes i.e. "*RST" or may be contained in

a variable of type string.

Return Value

Returns the status code of this operation. The status code

either indicates success or describes an error or warning

condition. You examine the status code from each call to an

instrument driver function to determine if an error occurred. To

obtain a text description of the status code, call the

rsfsu_error_message function.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

This driver defines the following status codes:

Status    Description

-------------------------------------------------

BFFC0002  Parameter 2 (Write Buffer) out of range.

BFFC09F0  Instrument status error

BFFC09F1  Option not installed

BFFC09F2  Conflict with setup parameters

BFFC09F3  Data are not available

BFFC09F4  Unknown option

BFFC09F5  Calibration failed

BFFC09F6  Function does not support FSEx parameter value

BFFC09F7  Overload. Instrument does not return I/Q values

This instrument driver also returns errors and warnings defined

by other sources. The following table defines the ranges of

additional status codes that this driver can return. The table

lists the different include files that contain the defined

constants for the particular status codes:

Numeric Range (in Hex)   Status Code Types

-------------------------------------------------

3FFF0000 to 3FFFFFFF     VISA     Warnings

3FFC0000 to 3FFCFFFF     VXIPnP   Driver Warnings

BFFF0000 to BFFFFFFF     VISA     Errors

BFFC0000 to BFFCFFFF     VXIPnP   Driver Errors